## AnalysisDataTest
RunID | InstrumentID | RunDate | NumberOfCyclesR1 | NumberofCyclesR2 | FlowCellInDate? | PR2BottleInDate? | ReagentKitInDate? | %ReadsPassingFilter | %ReadsPassingFilterOutsideThreshold? | NumberofReadsPassingFilter | ReadCountWithinThreshold? | MeanClusterDensityPassingFilter | MeanClusterDensityWithinThreshold | COVClusterDensity | COVClusterDensityExceedsThreshold? | %MedianClusterDensityPassingFilter | GapBetweenClusterDensitiesPassingFilterAndNotExceedsThreshold? | ProportionOverQ30 | IlluminaQ30Threshold | ProportionOverQ30ExceedsThreshold? | ProportionOverQ30Read1 | ProportionOverQ30ExceedsThresholdRead1 | ProportionOverQ30Read2 | ProportionOverQ30ExceedsThresholdRead2 | K-SPValue | K-SPValueSignificant? | K-SUnpackPvalue | M-WUnpackPvalue | NumberOfCyclesExceedingThresholdUnderQ30Read1 | SlopeOfLineOfBestFitRead1 | SlopeOfLineOfBestFitRead1ExceedsThreshold? | NumberOfCyclesExceedingThresholdUnderQ30Read2 | SlopeOfLineOfBestFitRead2 | SlopeOfLineOfBestFitRead2ExceedsThreshold? | ProportionOfUndeterminedReads? | COVIndexes | COVIndexesExceedsThreshold? | OutlyingIndexes | OutlyingIndexes(Low) | OutlyingIndexes(High) | StateOfNTC
150731_M00766_0121_000000000-AFMYC | M00766 | 2015-07-31 | 75 | 75 | in date | in date | in date | 89.5212 | no | 27660854 | low | 1152.6219 | low | 0.036004 | no | 89.6942 | no | 0.94787 | 0.85 | no | 0.95439 | no | 0.94056 | no | 0.99584 | no | 0.98595 | nan | 0 | -0.00062082 | yes | 0 | -0.00041404 | no | OK | 0.82507 | yes | ['15M07986'] | [] | ['15M07986'] | NTC OK
130510_M00766_0014_000000000-A3PFJ | M00766 | 2013-05-10 | 156 | 151 | in date | in date | in date | 88.5142 | no | 18378480 | low | 995.2515 | low | 0.013484 | no | 88.6611 | no | 0.95249 | 0.75 | no | 0.96867 | no | 0.93502 | no | 0.95413 | no | 3.3216e-08 | nan | 0 | -0.0002029 | no | 0 | -0.00047138 | no | OK | 0.51666 | yes | no outlying indexes |  |  | No NTC on run
151125_M00766_0158_000000000-AJJC1 | M00766 | 2015-11-25 | 151 | 151 | in date | in date | in date | 78.1687 | yes | 3563033 | low | 150.7739 | very low | 0.031274 | no | 78.1863 | yes | 0.9558 | 0.8 | no | 0.96712 | no | 0.94724 | no | 0.95413 | no | 0.17297 | 0.010728 | 0 | -0.00022353 | no | 0 | -0.00043897 | no | OK | 0.29368 | yes | ['15M12680' '15M12866'] | ['15M12680'] | ['15M12866'] | NTC OK
150911_M00766_0131_000000000-AGLDT | M00766 | 2015-09-11 | 75 | 75 | in date | in date | in date | 91.8669 | no | 29894066 | low | 1233.1842 | OK | 0.017257 | no | 92.0012 | no | 0.95954 | 0.85 | no | 0.97065 | no | 0.94872 | no | 0.84095 | no | 0.0036196 | 0.00029275 | 0 | -0.0004291 | no | 0 | -0.00031438 | no | OK | 0.30964 | yes | no outlying indexes |  |  | NTC OK
150429_M02641_0043_000000000-AD8KJ | M02641 | 2015-04-29 | 75 | 75 | in date | in date | in date | 77.8177 | yes | 32906358 | low | 1415.0412 | high | 0.018771 | no | 78.178 | yes | 0.8983 | 0.85 | no | 0.92026 | no | 0.87594 | no | 0.84095 | no | 0.00033968 | nan | 0 | -0.00089215 | yes | 0 | -0.00067513 | yes | OK | 0.50312 | yes | no outlying indexes |  |  | NTC OK
160218_M00766_0010_000000000-AMF48 | M00766 | 2016-02-18 | 75 | 75 | in date | in date | in date | 90.3938 | no | 26783167 | low | 1115.6764 | low | 0.019331 | no | 90.2226 | no | 0.95211 | 0.85 | no | 0.96432 | no | 0.93856 | no | 0.84095 | no | 0.0028637 | 0.00027102 | 0 | -0.00054184 | yes | 0 | -0.00048552 | no | OK | 0.32057 | yes | ['15M15396'] | [] | ['15M15396'] | NTC OK
150706_M00766_0118_000000000-AF512 | M00766 | 2015-07-06 | 200 | 200 | in date | in date | in date | 94.4229 | no | 16447526 | low | 844.1499 | very low | 0.016368 | no | 94.5523 | no | 0.86041 | 0.75 | no | 0.89416 | no | 0.82397 | no | 0.84095 | no | 3.1972e-28 | nan | 1 | -0.001537 | yes | 1 | -0.0023974 | yes | OK | 0.17724 | yes | no outlying indexes |  |  | NTC OK
150325_M00766_0084_000000000-AD8KN | M00766 | 2015-03-25 | 75 | 75 | in date | in date | in date | 93.0443 | no | 26726137 | low | 1096.1198 | very low | 0.025085 | no | 93.5238 | no | 0.96479 | 0.85 | no | 0.97449 | no | 0.95451 | no | 0.84095 | no | 0.047651 | 0.00076142 | 0 | -0.00034315 | no | 0 | -0.00025609 | no | OK | 0.617 | yes | no outlying indexes |  |  | NTC OK
151022_M02641_0042_000000000-AJ5B5 | M02641 | 2015-10-22 | 200 | 200 | in date | in date | in date | 89.6462 | no | 28861426 | low | 1217.8503 | OK | 0.010571 | no | 89.8396 | no | 0.67585 | 0.7 | yes | 0.7507 | no | 0.59122 | yes | 0.67794 | no | 4.6203e-88 | 7.0766e-116 | 12 | -0.0033975 | yes | 75 | -0.0048318 | yes | OK | 0.26348 | yes | ['15M09575' '15M10221'] | [] | ['15M09575' '15M10221'] | NTC OK
160210_M02641_0077_000000000-AME81 | M02641 | 2016-02-10 | 75 | 75 | in date | in date | in date | 93.9536 | no | 19770059 | low | 818.1394 | very low | 0.024181 | no | 94.3148 | no | 0.97229 | 0.85 | no | 0.97917 | no | 0.96582 | no | 0.95413 | no | 0.58184 | 6.9811e-05 | 0 | -0.00021377 | no | 0 | -0.00022485 | no | OK | 0.54079 | yes | no outlying indexes |  |  | NTC OK
151210_M00766_0168_000000000-AJDDH | M00766 | 2015-12-10 | 75 | 75 | in date | in date | in date | 91.376 | no | 26145512 | low | 1086.9972 | very low | 0.021546 | no | 90.8931 | no | 0.956 | 0.85 | no | 0.96684 | no | 0.94501 | no | 0.84095 | no | 0.10224 | 0.014367 | 0 | -0.00051081 | yes | 0 | -0.0004594 | no | OK | 0.4035 | yes | ['15M13946'] | ['15M13946'] | [] | NTC OK
160128_M00766_0002_000000000-ALAJ1 | M00766 | 2016-01-28 | 75 | 75 | in date | in date | in date | 88.2213 | no | 28975202 | low | 1238.6174 | OK | 0.0080531 | no | 88.3898 | no | 0.95207 | 0.85 | no | 0.96129 | no | 0.94181 | no | 0.84095 | no | 0.078659 | 0.012869 | 0 | -0.00042109 | no | 0 | -0.00041041 | no | OK | 0.33634 | yes | no outlying indexes |  |  | NTC OK
150313_M02641_0037_000000000-ACC8R | M02641 | 2015-03-13 | 151 | 151 | in date | in date | in date | 94.5096 | no | 17647098 | low | 906.6463 | low | 0.017228 | no | 93.7282 | no | 0.94213 | 0.8 | no | 0.9595 | no | 0.9242 | no | 0.67794 | no | 2.7478e-10 | nan | 0 | -0.00051188 | yes | 0 | -0.00060154 | yes | OK | 0.11578 | yes | no outlying indexes |  |  | NTC OK
160209_M02641_0076_000000000-AMF44 | M02641 | 2016-02-09 | 75 | 75 | in date | in date | in date | 89.416 | no | 23928513 | low | 1023.4734 | very low | 0.02047 | no | 89.4341 | no | 0.95438 | 0.85 | no | 0.96439 | no | 0.94402 | no | 0.84095 | no | 0.15811 | 4.2223e-80 | 0 | -0.00036997 | no | 0 | -0.00030272 | no | OK | 0.28461 | yes | ['15M15069'] | [] | ['15M15069'] | NTC OK
160105_M00766_0174_000000000-AL4LB | M00766 | 2016-01-05 | 151 | 151 | in date | in date | in date | 91.2681 | no | 14218489 | low | 749.2227 | very low | 0.036601 | no | 91.4658 | no | 0.93032 | 0.8 | no | 0.95273 | no | 0.9068 | no | 0.95413 | no | 1.8718e-11 | nan | 0 | -0.00053469 | yes | 0 | -0.00071583 | yes | OK | 0.63317 | yes | ['15M12287'] | [] | ['15M12287'] | NTC OK
140917_M02641_0018_000000000-AA3H7 | M02641 | 2014-09-17 | 75 | 75 | in date | in date | in date | 91.5859 | no | 24528146 | low | 1028.5468 | very low | 0.035153 | no | 91.5237 | no | 0.96229 | 0.85 | no | 0.96999 | no | 0.95452 | no | 0.84095 | no | 0.88947 | nan | 0 | -0.00031376 | no | 0 | -0.00016688 | no | OK | 0.27666 | yes | ['14M00640'] | ['14M00640'] | [] | NTC OK
130909_M00766_0047_000000000-A5PEH | M00766 | 2013-09-09 | 151 | 151 | in date | in date | in date | 91.8202 | no | 17818894 | low | 943.064 | low | 0.013531 | no | 91.9043 | no | 0.91659 | 0.8 | no | 0.93844 | no | 0.89323 | no | 0.95413 | no | 2.3926e-11 | 3.6095e-18 | 0 | -0.00083727 | yes | 0 | -0.00095948 | yes | OK | 0.36674 | yes | ['Control'] | ['Control'] | [] | No NTC on run
160205_M02641_0073_000000000-ALY9Y | M02641 | 2016-02-05 | 75 | 75 | in date | in date | in date | 81.6149 | no | 30119640 | low | 1331.3992 | OK | 0.025391 | no | 80.9054 | yes | 0.92844 | 0.85 | no | 0.94565 | no | 0.90982 | no | 0.84095 | no | 0.023743 | nan | 0 | -0.0004287 | no | 0 | -0.00027127 | no | OK | 0.37123 | yes | no outlying indexes |  |  | NTC OK
150127_M00766_0072_000000000-ABRK7 | M00766 | 2015-01-27 | 151 | 151 | in date | in date | out of date | 94.2571 | no | 17854832 | low | 926.0782 | low | 0.029511 | no | 94.3435 | no | 0.93257 | 0.8 | no | 0.95668 | no | 0.90691 | no | 0.67794 | no | 1.1008e-18 | 1.0313e-37 | 0 | -0.00054076 | yes | 0 | -0.00067649 | yes | OK | 0.23558 | yes | ['14M12186'] | [] | ['14M12186'] | NTC OK
151204_M02641_0056_000000000-AJDC6 | M02641 | 2015-12-04 | 200 | 200 | in date | in date | in date | 93.7017 | no | 25575432 | low | 1042.2805 | very low | 0.036678 | no | 94.2736 | no | 0.85421 | 0.7 | no | 0.88062 | no | 0.82404 | no | 0.50766 | no | 2.3201e-29 | nan | 13 | -0.0018496 | yes | 19 | -0.0025865 | yes | OK | 0.1513 | yes | ['15M04443'] | [] | ['15M04443'] | NTC OK
160212_M00766_0008_000000000-AL607 | M00766 | 2016-02-12 | 151 | 151 | in date | in date | in date | 87.5244 | no | 20878061 | low | 1123.9732 | OK | 0.026805 | no | 87.1194 | no | 0.85079 | 0.8 | no | 0.8926 | no | 0.80667 | no | 0.84095 | no | 5.1098e-39 | nan | 0 | -0.002166 | yes | 10 | -0.0030426 | yes | OK | 0.42813 | yes | ['15M12553' '15M12790'] | ['15M12553' '15M12790'] | [] | NTC OK
160218_M02641_0080_000000000-AMF3Y | M02641 | 2016-02-18 | 75 | 75 | in date | in date | in date | 92.7164 | no | 22919923 | low | 955.1644 | very low | 0.010077 | no | 92.7435 | no | 0.95689 | 0.85 | no | 0.97361 | no | 0.93955 | no | 0.84095 | no | 0.00019591 | nan | 0 | -0.00031985 | no | 0 | -0.00022652 | no | OK | 0.48739 | yes | no outlying indexes |  |  | NTC OK
151104_M00766_0151_000000000-AK8EW | M00766 | 2015-11-04 | 151 | 151 | in date | in date | in date | 82.696 | no | 16143415 | low | 883.0439 | low | 0.069792 | yes | 83.2962 | yes | 0.91934 | 0.8 | no | 0.94177 | no | 0.89677 | no | 0.95413 | no | 1.0753e-07 | 5.9859e-42 | 0 | -0.00035735 | no | 0 | -0.00067331 | yes | OK | 0.40846 | yes | no outlying indexes |  |  | NTC OK
150423_M02641_0041_000000000-ACN0P | M02641 | 2015-04-23 | 75 | 75 | in date | in date | in date | 91.9691 | no | 30009055 | low | 1228.4129 | OK | 0.027957 | no | 92.2905 | no | 0.9501 | 0.85 | no | 0.96725 | no | 0.93306 | no | 0.84095 | no | 2.6066e-06 | 3.4467e-12 | 0 | -0.00047255 | no | 0 | -0.00049099 | no | OK | 0.53918 | yes | ['15M03290'] | [] | ['15M03290'] | NTC OK
151119_M02641_0052_000000000-AJJ6B | M02641 | 2015-11-19 | 75 | 75 | in date | in date | out of date | 94.5733 | no | 24895626 | low | 1016.266 | very low | 0.029336 | no | 94.9966 | no | 0.96439 | 0.85 | no | 0.97782 | no | 0.95208 | no | 0.84095 | no | 0.00044318 | 1.1267e-06 | 0 | -0.00028501 | no | 0 | -0.00037122 | no | OK | 0.33896 | yes | no outlying indexes |  |  | NTC OK
151208_M02641_0057_000000000-AJHRF | M02641 | 2015-12-08 | 151 | 151 | in date | in date | in date | 97.1211 | no | 4496775 | low | 228.8877 | very low | 0.02523 | no | 97.1074 | no | 0.94501 | 0.8 | no | 0.95817 | no | 0.93055 | no | 0.95413 | no | 0.083233 | 0.003379 | 0 | -0.0007036 | yes | 0 | -0.00058399 | yes | OK | 0.58595 | yes | ['15M12287'] | [] | ['15M12287'] | NTC OK
141024_M02641_0024_000000000-A8P84 | M02641 | 2014-10-24 | 151 | 151 | in date | in date | in date | 94.9748 | no | 15281418 | low | 794.7029 | very low | 0.031108 | no | 95.0899 | no | 0.94073 | 0.8 | no | 0.95744 | no | 0.92359 | no | 0.84095 | no | 1.9766e-08 | 4.7689e-23 | 0 | -0.00055662 | yes | 0 | -0.00062808 | yes | OK | 0.077865 | yes | no outlying indexes |  |  | NTC OK
130417_M00766_0007_000000000-A3R1F | M00766 | 2013-04-17 | 151 | 151 | in date | in date | in date | 93.6874 | no | 14918743 | low | 773.0359 | very low | 0.024644 | no | 93.6673 | no | 0.97122 | 0.8 | no | 0.98017 | no | 0.96226 | no | 0.84095 | no | 0.0022684 | 4.2948e-06 | 0 | -0.00013884 | no | 0 | -0.00025711 | no | No index information for 130417_M00766_0007_000000000-A3R1F
140804_M02641_0011_000000000-AAD6A | M02641 | 2014-08-04 | 251 | 251 | in date | in date | in date | 92.8278 | no | 18236023 | low | 947.4631 | low | 0.014242 | no | 92.5473 | no | 0.75551 | 0.75 | no | 0.79909 | no | 0.70633 | yes | 0.84095 | no | 4.579e-31 | nan | 19 | -0.0025558 | yes | 41 | -0.0033381 | yes | OK | 0.15644 | yes | no outlying indexes |  |  | No NTC on run
140620_M02641_0004_000000000-A8R7V | M02641 | 2014-06-20 | 151 | 151 | in date | in date | in date | 95.2324 | no | 15895393 | low | 815.6811 | very low | 0.010793 | no | 95.1913 | no | 0.92041 | 0.8 | no | 0.96394 | no | 0.87468 | no | 0.35842 | no | 5.5248e-61 | nan | 0 | -0.0004185 | no | 0 | -0.00070599 | yes | High | 0.41518 | yes | ['14M02611'] | ['14M02611'] | [] | NTC OK
150910_M02641_0023_000000000-AFLHH | M02641 | 2015-09-10 | 151 | 151 | in date | in date | in date | 89.5554 | no | 20794433 | low | 1100.488 | OK | 0.02384 | no | 88.8672 | no | 0.85359 | 0.8 | no | 0.88668 | no | 0.81835 | no | 0.95413 | no | 2.3039e-18 | nan | 0 | -0.0021309 | yes | 9 | -0.0024176 | yes | OK | 0.92371 | yes | ['15M04253'] | [] | ['15M04253'] | NTC OK
150325_M02641_0038_000000000-AD8KY | M02641 | 2015-03-25 | 75 | 75 | in date | in date | in date | 92.8928 | no | 28292154 | low | 1150.8931 | low | 0.016396 | no | 92.9934 | no | 0.96863 | 0.85 | no | 0.97441 | no | 0.9626 | no | 0.95413 | no | 0.97963 | nan | 0 | -0.00030202 | no | 0 | -0.00017631 | no | OK | 0.27087 | yes | no outlying indexes |  |  | NTC OK
150820_M02641_0017_000000000-AGJJA | M02641 | 2015-08-20 | 75 | 75 | in date | in date | in date | 93.4116 | no | 26592471 | low | 1089.8361 | very low | 0.015153 | no | 93.7457 | no | 0.96481 | 0.85 | no | 0.97351 | no | 0.95717 | no | 0.84095 | no | 0.072701 | 0.016306 | 0 | -0.00030311 | no | 0 | -0.00031655 | no | OK | 0.26272 | yes | no outlying indexes |  |  | NTC OK
151021_M00766_0144_000000000-AG028 | M00766 | 2015-10-21 | 75 | 75 | in date | in date | in date | 85.0537 | no | 24774347 | low | 1085.7652 | very low | 0.013905 | no | 84.6897 | yes | 0.94129 | 0.85 | no | 0.95646 | no | 0.92591 | no | 0.84095 | no | 0.0030374 | 4.0367e-11 | 0 | -0.00055587 | yes | 0 | -0.00055374 | yes | OK | 0.46617 | yes | ['15M10881'] | ['15M10881'] | [] | NTC OK
160118_M02641_0066_000000000-AJETF | M02641 | 2016-01-18 | 75 | 75 | in date | in date | in date | 98.1898 | no | 9321517 | low | 373.3045 | very low | 0.028008 | no | 98.3733 | no | 0.97492 | 0.85 | no | 0.97945 | no | 0.97382 | no | 0.99998 | no | 1 | 0.5283 | 0 | -0.00022747 | no | 0 | -0.000157 | no | OK | 0.35791 | yes | ['15M13946'] | ['15M13946'] | [] | NTC OK
151111_M02641_0049_000000000-AJD47 | M02641 | 2015-11-11 | 200 | 200 | in date | in date | in date | 92.8535 | no | 30208881 | low | 1238.2172 | OK | 0.021332 | no | 93.316 | no | 0.83991 | 0.7 | no | 0.87018 | no | 0.80592 | no | 0.50766 | no | 8.1049e-40 | 6.164e-109 | 14 | -0.0019638 | yes | 19 | -0.0027647 | yes | OK | 0.28743 | yes | ['15M10790'] | [] | ['15M10790'] | NTC OK
150430_M02641_0044_000000000-AD8VC | M02641 | 2015-04-30 | 75 | 75 | in date | in date | in date | 86.9368 | no | 27790413 | low | 1196.1245 | low | 0.022911 | no | 87.4623 | no | 0.92905 | 0.85 | no | 0.95387 | no | 0.90368 | no | 0.84095 | no | 1.2394e-10 | 1.7232e-11 | 0 | -0.00060018 | yes | 0 | -0.0005903 | yes | OK | 0.4048 | yes | no outlying indexes |  |  | NTC OK
150114_M00766_0069_000000000-AA68B | M00766 | 2015-01-14 | 151 | 151 | in date | in date | in date | 77.8262 | yes | 20704375 | low | 1148.0758 | OK | 0.024397 | no | 77.3315 | yes | 0.86381 | 0.8 | no | 0.91949 | no | 0.81333 | no | 0.95413 | no | 2.6748e-38 | nan | 0 | -0.00060681 | yes | 4 | -0.00057587 | yes | OK | 0.18923 | yes | ['14M09058'] | ['14M09058'] | [] | NTC OK
150306_M02641_0034_000000000-ACRUF | M02641 | 2015-03-06 | 75 | 75 | in date | in date | in date | 91.3955 | no | 24975924 | low | 1049.1583 | very low | 0.0067808 | no | 91.297 | no | 0.95124 | 0.85 | no | 0.9701 | no | 0.93153 | no | 0.95413 | no | 3.4327e-07 | nan | 0 | -0.00040766 | no | 0 | -0.00082529 | yes | OK | 0.17399 | yes | no outlying indexes |  |  | NTC OK
151002_M02641_0033_000000000-AFLDA | M02641 | 2015-10-02 | 151 | 151 | in date | in date | in date | 88.0983 | no | 23364962 | low | 1229.9955 | high | 0.042476 | no | 87.9251 | no | 0.90553 | 0.8 | no | 0.937 | no | 0.87138 | no | 0.84095 | no | 1.4773e-23 | 1.9239e-34 | 0 | -0.00058099 | yes | 0 | -0.0010877 | yes | OK | 0.54428 | yes | ['8M4079'] | [] | ['8M4079'] | NTC OK
150810_M00766_0123_000000000-AGTUY | M00766 | 2015-08-10 | 151 | 151 | in date | in date | in date | 89.2945 | no | 22283310 | low | 1188.9462 | OK | 0.016194 | no | 88.9362 | no | 0.92177 | 0.8 | no | 0.95497 | no | 0.88475 | no | 0.84095 | no | 2.7093e-32 | nan | 0 | -0.00035243 | no | 0 | -0.00083315 | yes | OK | 0.36536 | yes | ['15M07650'] | [] | ['15M07650'] | NTC OK
150205_M02641_0029_000000000-ACCE2 | M02641 | 2015-02-05 | 151 | 151 | in date | in date | in date | 87.5711 | no | 18965301 | low | 1006.645 | OK | 0.019074 | no | 87.6822 | no | 0.8298 | 0.8 | no | 0.8663 | no | 0.79202 | yes | 0.99584 | no | 1.5184e-12 | 8.2211e-115 | 0 | -0.0022136 | yes | 2 | -0.0018972 | yes | OK | 0.25963 | yes | ['13M12234' '15M00575'] | ['13M12234' '15M00575'] | [] | NTC OK
150729_M02641_0008_000000000-AGET9 | M02641 | 2015-07-29 | 200 | 200 | in date | in date | in date | 91.862 | no | 30439243 | low | 1258.6104 | OK | 0.012085 | no | 92.3982 | no | 0.84428 | 0.7 | no | 0.87341 | no | 0.81136 | no | 0.67794 | no | 1.7873e-34 | 2.8556e-270 | 1 | -0.0016885 | yes | 2 | -0.0024684 | yes | OK | 0.27015 | yes | ['15M05450'] | [] | ['15M05450'] | NTC OK
150501_M00766_0093_000000000-AF9MH | M00766 | 2015-05-01 | 75 | 75 | in date | in date | in date | 87.2253 | no | 34467195 | low | 1450.6461 | high | 0.022012 | no | 87.0197 | no | 0.94139 | 0.85 | no | 0.95639 | no | 0.92569 | no | 0.84095 | no | 0.00057945 | 2.9759e-05 | 0 | -0.00056762 | yes | 0 | -0.00052432 | yes | OK | 0.38312 | yes | ['14M11012'] | [] | ['14M11012'] | NTC OK
151211_M00766_0169_000000000-AK5DW | M00766 | 2015-12-11 | 300 | 300 | in date | in date | in date | 93.0016 | no | 21872195 | low | 891.8405 | very low | 0.023515 | no | 93.6566 | no | 0.54826 | 0.7 | yes | 0.59108 | yes | 0.48586 | yes | 0.24079 | no | 3.3919e-41 | 2.4542e-59 | 169 | -0.0035153 | yes | 181 | -0.004229 | yes | OK | 0.21949 | yes | ['14M10927b'] | [] | ['14M10927b'] | NTC OK
140603_M00766_0037_000000000-A7WU9 | M00766 | 2014-06-03 | 251 | 251 | in date | in date | in date | 80.6392 | no | 20279789 | low | 1121.2903 | OK | 0.020185 | no | 81.2363 | yes | 0.67262 | 0.75 | yes | 0.87096 | no | 0.46478 | yes | 0.24079 | no | 0 | 0 | 0 | -0.0011435 | yes | 73 | -0.0033779 | yes | OK | 0.33214 | yes | no outlying indexes |  |  | No NTC on run
150528_M00766_0105_000000000-AF9MJ | M00766 | 2015-05-28 | 75 | 75 | in date | in date | in date | 91.4093 | no | 30434248 | low | 1259.146 | OK | 0.033027 | no | 91.4154 | no | 0.9591 | 0.85 | no | 0.96908 | no | 0.94896 | no | 0.84095 | no | 0.011641 | 5.539e-05 | 0 | -0.00043918 | no | 0 | -0.00032915 | no | OK | 0.43106 | yes | ['14M12059'] | [] | ['14M12059'] | NTC OK
150130_M00766_0073_000000000-ACBVB | M00766 | 2015-01-30 | 251 | 251 | in date | in date | in date | 81.2781 | no | 23929178 | low | 1312.9896 | very high | 0.017972 | no | 81.1332 | yes | 0.66236 | 0.75 | yes | 0.70216 | yes | 0.61962 | yes | 0.84095 | no | 6.7473e-31 | nan | 47 | -0.0033611 | yes | 64 | -0.0034493 | yes | OK | 0.3138 | yes | no outlying indexes |  |  | No NTC on run
150610_M02641_0059_000000000-AFN4H | M02641 | 2015-06-10 | 75 | 75 | in date | in date | in date | 82.3028 | no | 25297112 | low | 1022.9657 | very low | 0.024105 | no | 83.0805 | yes | 0.91102 | 0.85 | no | 0.93192 | no | 0.89027 | no | 0.84095 | no | 0.00015742 | 7.1708e-18 | 0 | -0.00077999 | yes | 0 | -0.00063882 | yes | OK | 0.52264 | yes | ['15M03072'] | [] | ['15M03072'] | NTC OK
160119_M02641_0067_000000000-ALN4P | M02641 | 2016-01-19 | 151 | 151 | in date | in date | in date | 87.2125 | no | 18949772 | low | 995.7574 | low | 0.091472 | yes | 86.1573 | no | 0.9035 | 0.8 | no | 0.92594 | no | 0.87789 | no | 0.95413 | no | 2.2467e-08 | nan | 0 | -0.00078266 | yes | 0 | -0.0012195 | yes | OK | 0.31982 | yes | ['15M14816'] | [] | ['15M14816'] | NTC OK
150930_M02641_0031_000000000-AGJGT | M02641 | 2015-09-30 | 75 | 75 | in date | in date | in date | 93.3234 | no | 29303547 | low | 1206.8802 | OK | 0.049805 | no | 93.5156 | no | 0.9573 | 0.85 | no | 0.9715 | no | 0.94372 | no | 0.84095 | no | 7.6841e-05 | nan | 0 | -0.00045184 | no | 0 | -0.00030011 | no | OK | 0.51261 | yes | no outlying indexes |  |  | NTC OK
150518_M00766_0100_000000000-AATJ7 | M00766 | 2015-05-18 | 151 | 151 | in date | in date | in date | 89.6123 | no | 7467630 | low | 351.2212 | very low | 0.03759 | no | 88.7216 | no | 0.94106 | 0.8 | no | 0.95572 | no | 0.92587 | no | 0.84095 | no | 0.55116 | 0.39443 | 0 | -0.00037129 | no | 0 | -0.00084743 | yes | OK | 2.6175 | yes | ['Faststart' '360Taq4mM' '360Taq2mM'] | [] | ['Faststart' '360Taq4mM' '360Taq2mM'] | NTC OK
150401_M00766_0086_000000000-ACBU5 | M00766 | 2015-04-01 | 151 | 151 | in date | in date | in date | 90.4494 | no | 18801773 | low | 986.6817 | low | 0.029182 | no | 90.7591 | no | 0.8401 | 0.8 | no | 0.87195 | no | 0.80659 | no | 0.99584 | no | 6.8105e-11 | 1.2646e-33 | 0 | -0.0022957 | yes | 6 | -0.0022224 | yes | High | 2.2561 | yes | ['15M00594' '15M00616'] | [] | ['15M00594' '15M00616'] | NTC OK
130613_M00766_0018_000000000-A4FN3 | M00766 | 2013-06-13 | 251 | 251 | in date | in date | in date | 62.2349 | yes | 14759376 | low | 877.4421 | low | 0.0052548 | no | 61.6486 | yes | 0.68299 | 0.75 | yes | 0.7656 | no | 0.58473 | yes | 0.67794 | no | 5.4172e-57 | 1.5425e-86 | 28 | -0.0020629 | yes | 55 | -0.0033012 | yes | OK | 0.15 | yes | no outlying indexes |  |  | No NTC on run
150605_M02641_0058_000000000-AFF15 | M02641 | 2015-06-05 | 75 | 75 | in date | in date | in date | 86.9846 | no | 20446753 | low | 822.314 | very low | 0.026851 | no | 87.5712 | no | 0.93738 | 0.85 | no | 0.94849 | no | 0.92669 | no | 0.95413 | no | 0.71981 | nan | 0 | -0.00062227 | yes | 0 | -0.00041516 | no | OK | 0.38966 | yes | ['14M08356'] | ['14M08356'] | [] | NTC OK
160204_M00766_0004_000000000-AL4H0 | M00766 | 2016-02-04 | 151 | 151 | in date | in date | in date | 95.4825 | no | 14166224 | low | 726.9693 | very low | 0.021398 | no | 95.3691 | no | 0.95549 | 0.8 | no | 0.97197 | no | 0.93852 | no | 0.84095 | no | 3.1187e-14 | 2.8941e-31 | 0 | -0.00029231 | no | 0 | -0.00059973 | yes | OK | 0.56692 | yes | ['15M13181'] | [] | ['15M13181'] | NTC OK
150507_M02641_0046_000000000-AF7N5 | M02641 | 2015-05-07 | 75 | 75 | in date | in date | in date | 92.9775 | no | 27595623 | low | 1124.3595 | low | 0.040298 | no | 93.1424 | no | 0.95227 | 0.85 | no | 0.96981 | no | 0.93556 | no | 0.84095 | no | 9.3853e-07 | 5.8544e-20 | 0 | -0.00053503 | yes | 0 | -0.0001815 | no | OK | 0.39453 | yes | no outlying indexes |  |  | NTC OK
150116_M00766_0070_000000000-ACC3G | M00766 | 2015-01-16 | 151 | 151 | in date | in date | in date | 86.297 | no | 14147272 | low | 752.8015 | very low | 0.040986 | no | 86.2322 | no | 0.93362 | 0.8 | no | 0.95795 | no | 0.91769 | no | 0.95413 | no | 0.027224 | 4.5409e-07 | 0 | -0.00014326 | no | 3 | 0.00022032 | no | OK | 0.37649 | yes | ['14M11762-2' '3M0956-2' '3M0956' '14M11762'] | ['14M11762-2' '3M0956-2' '3M0956' '14M11762'] | [] | NTC OK
150220_M00766_0077_000000000-AA2U8 | M00766 | 2015-02-20 | 26 | 26 | in date | in date | in date | 11.2989 | yes | 3422864 | low | 175.3978 | very low | 0.069387 | yes | 0 | yes | 0.63211 | 0.9 | yes | 0.88444 | yes | 0.47035 | yes | 0.95413 | no | 6.5681e-12 | 7.5094e-13 | 0 | 0 | no | 26 | 0 | no | High | 0.6497 | yes | no outlying indexes |  |  | No NTC on run
130708_M00766_0023_000000000-A20R8 | M00766 | 2013-07-08 | 151 | 151 | out of date | in date | out of date | 90.9082 | no | 8144995 | None | 872.0913 | None | 0.015455 | no | 90.7757 | no | 0.86589 | 0.85 | no | 0.91799 | no | 0.81047 | yes | 0.67794 | no | 9.6929e-19 | 6.9078e-114 | 0 | -0.0013571 | yes | 0 | -0.0030958 | yes | OK | 0.10661 | yes | no outlying indexes |  |  | No NTC on run
160107_M02641_0063_000000000-AJDC9 | M02641 | 2016-01-07 | 75 | 75 | in date | in date | in date | 94.1624 | no | 24963706 | low | 1022.5057 | very low | 0.017799 | no | 94.5508 | no | 0.9694 | 0.85 | no | 0.97637 | no | 0.96261 | no | 0.84095 | no | 0.22798 | nan | 0 | -0.00029644 | no | 0 | -0.00032503 | no | OK | 0.51228 | yes | no outlying indexes |  |  | NTC OK
150807_M02641_0012_000000000-AFMYT | M02641 | 2015-08-07 | 75 | 75 | in date | in date | in date | 96.8172 | no | 16466109 | low | 661.2988 | very low | 0.043831 | no | 97.2034 | no | 0.97599 | 0.85 | no | 0.98258 | no | 0.97147 | no | 0.95413 | no | 0.81203 | 0.23825 | 0 | -0.00028401 | no | 0 | -0.0001526 | no | OK | 0.35887 | yes | ['15M08559'] | [] | ['15M08559'] | NTC OK
151116_M00766_0154_000000000-AJF4C | M00766 | 2015-11-16 | 151 | 151 | in date | in date | in date | 87.9376 | no | 13835502 | low | 734.2783 | very low | 0.054577 | yes | 88.5491 | no | 0.9313 | 0.8 | no | 0.95109 | no | 0.90966 | no | 0.95413 | no | 5.1553e-08 | nan | 0 | -0.00033626 | no | 0 | -0.00061836 | yes | OK | 0.25181 | yes | ['15M12301'] | ['15M12301'] | [] | NTC OK
150820_M00766_0126_000000000-AGHUU | M00766 | 2015-08-20 | 75 | 75 | in date | in date | in date | 92.7672 | no | 28026234 | low | 1148.9642 | low | 0.020853 | no | 93.2904 | no | 0.96152 | 0.85 | no | 0.97405 | no | 0.94902 | no | 0.84095 | no | 0.00031602 | nan | 0 | -0.00033273 | no | 0 | -0.0004084 | no | OK | 0.70143 | yes | no outlying indexes |  |  | NTC OK
150507_M00766_0095_000000000-AF7N3 | M00766 | 2015-05-07 | 75 | 75 | in date | in date | in date | 92.6658 | no | 27410104 | low | 1119.7122 | low | 0.036603 | no | 93.5327 | no | 0.96159 | 0.85 | no | 0.9741 | no | 0.94953 | no | 0.84095 | no | 0.00099419 | 8.626e-08 | 0 | -0.00036389 | no | 0 | -0.00042845 | no | OK | 0.29241 | yes | ['14M10600'] | [] | ['14M10600'] | NTC OK
160208_M02641_0075_000000000-AMF40 | M02641 | 2016-02-08 | 75 | 75 | in date | in date | in date | 79.3999 | yes | 30335883 | low | 1362.3448 | OK | 0.022595 | no | 80.0194 | yes | 0.9085 | 0.85 | no | 0.93381 | no | 0.88187 | no | 0.84095 | no | 9.8314e-08 | 2.2541e-295 | 0 | -0.00078696 | yes | 0 | -0.00037771 | no | OK | 0.359 | yes | ['16M00691'] | [] | ['16M00691'] | NTC OK
160115_M00766_0178_000000000-ALKJV | M00766 | 2016-01-15 | 151 | 151 | in date | in date | in date | 83.9557 | no | 15235349 | low | 831.5431 | very low | 0.066527 | yes | 83.4762 | yes | 0.91226 | 0.8 | no | 0.93609 | no | 0.8844 | no | 0.95413 | no | 5.102e-09 | nan | 0 | -0.00057903 | yes | 0 | -0.0011478 | yes | OK | 0.31466 | yes | ['15M13506'] | ['15M13506'] | [] | NTC OK
160216_M02641_0079_000000000-AL75M | M02641 | 2016-02-16 | 151 | 151 | in date | in date | in date | 92.5527 | no | 17259426 | low | 905.0069 | low | 0.019188 | no | 93.4136 | no | 0.9397 | 0.8 | no | 0.96218 | no | 0.91691 | no | 0.67794 | no | 4.7356e-16 | nan | 0 | -0.00030351 | no | 0 | -0.00073347 | yes | OK | 0.44782 | yes | ['15M13455' '15M13442'] | ['15M13455' '15M13442'] | [] | NTC OK
150811_M02641_0014_000000000-AGK0F | M02641 | 2015-08-11 | 151 | 151 | in date | in date | in date | 92.4253 | no | 22253985 | low | 1160.0949 | OK | 0.016889 | no | 92.3445 | no | 0.90106 | 0.8 | no | 0.9283 | no | 0.87132 | no | 0.95413 | no | 7.8772e-25 | nan | 0 | -0.0009934 | yes | 0 | -0.0015021 | yes | OK | 0.17315 | yes | no outlying indexes |  |  | NTC OK
151218_M00766_0171_000000000-AK5FL | M00766 | 2015-12-18 | 226 | 226 | in date | in date | in date | 92.0182 | no | 20535429 | low | 846.6865 | very low | 0.020936 | no | 91.9926 | no | 0.57401 | 0.7 | yes | 0.58212 | yes | 0.54171 | yes | 0.67794 | no | 3.7819e-14 | 8.2184e-29 | 24 | -0.005182 | yes | 97 | -0.0055294 | yes | OK | 0.25591 | yes | no outlying indexes |  |  | NTC OK
150911_M02641_0024_AF50J | M02641 | 2015-09-11 | 151 | 151 | out of date | in date | in date | 91.3886 | no | 20193785 | low | 1043.6462 | OK | 0.026601 | no | 91.1381 | no | 0.85976 | 0.8 | no | 0.89102 | no | 0.82694 | no | 0.84095 | no | 4.1997e-18 | nan | 0 | -0.0020335 | yes | 5 | -0.0021409 | yes | OK | 0.64106 | yes | ['15M04270'] | [] | ['15M04270'] | NTC OK
150227_M02641_0032_000000000-ACNF3 | M02641 | 2015-02-27 | 75 | 75 | in date | in date | in date | 90.2941 | no | 25317785 | low | 1066.3007 | very low | 0.076375 | yes | 90.5194 | no | 0.95732 | 0.85 | no | 0.9695 | no | 0.94388 | no | 0.84095 | no | 0.0023737 | nan | 0 | -0.00027283 | no | 0 | -0.00038603 | no | OK | 0.80447 | yes | ['14M08864'] | [] | ['14M08864'] | NTC OK
150925_M02641_0028_000000000-AGK08 | M02641 | 2015-09-25 | 75 | 75 | in date | in date | in date | 95.0618 | no | 24906667 | low | 1010.387 | very low | 0.01775 | no | 95.4096 | no | 0.96734 | 0.85 | no | 0.97861 | no | 0.95784 | no | 0.84095 | no | 0.0070482 | 1.5449e-09 | 0 | -0.00035963 | no | 0 | -0.00022766 | no | OK | 0.33049 | yes | no outlying indexes |  |  | NTC OK
151130_M00766_0163_000000000-AJGC9 | M00766 | 2015-11-30 | 151 | 151 | in date | in date | in date | 79.0357 | yes | 21816702 | low | 1234.7264 | high | 0.043221 | no | 79.6269 | yes | 0.91879 | 0.8 | no | 0.94269 | no | 0.8946 | no | 0.95413 | no | 4.5777e-17 | 3.5682e-28 | 0 | -0.0002497 | no | 0 | -0.00046641 | no | OK | 0.37872 | yes | ['15M12680' '15M12866'] | ['15M12680'] | ['15M12866'] | NTC OK
150615_M00766_0113_000000000-AFBH3 | M00766 | 2015-06-15 | 251 | 251 | in date | in date | in date | 89.384 | no | 20944521 | low | 1111.4379 | OK | 0.037506 | no | 89.0846 | no | 0.71363 | 0.75 | yes | 0.81632 | no | 0.60535 | yes | 0.50766 | no | 3.3983e-154 | 5.608e-166 | 24 | -0.0023213 | yes | 53 | -0.004186 | yes | OK | 0.11445 | yes | ['15M01148'] | [] | ['15M01148'] | NTC OK
130902_M00766_0045_000000000-A5BVD | M00766 | 2013-09-02 | 156 | 151 | in date | in date | in date | 92.4439 | no | 17107054 | low | 889.9218 | low | 0.020943 | no | 92.3467 | no | 0.95012 | 0.75 | no | 0.96778 | no | 0.93094 | no | 0.84095 | no | 1.4736e-06 | nan | 0 | -0.00028935 | no | 1 | -0.00072347 | yes | OK | 0.59937 | yes | no outlying indexes |  |  | No NTC on run
151009_M00766_0140_000000000-AGK0B | M00766 | 2015-10-09 | 75 | 75 | in date | in date | in date | 87.3055 | no | 24687917 | low | 1048.2614 | very low | 0.024799 | no | 87.0808 | no | 0.95594 | 0.85 | no | 0.96314 | no | 0.94902 | no | 0.99584 | no | 0.95838 | 0.16034 | 0 | -0.00045147 | no | 0 | -0.00036923 | no | OK | 0.4225 | yes | ['15M10481'] | ['15M10481'] | [] | NTC OK
160127_M02641_0069_000000000-AMDH4 | M02641 | 2016-01-27 | 75 | 75 | in date | in date | in date | 76.69 | yes | 29747487 | low | 1328.6387 | OK | 0.096919 | yes | 78.2455 | yes | 0.91416 | 0.85 | no | 0.93371 | no | 0.89378 | no | 0.84095 | no | 0.00062682 | nan | 0 | -0.00074637 | yes | 0 | -0.00037367 | no | OK | 0.52107 | yes | no outlying indexes |  |  | NTC OK
151012_M02641_0037_000000000-AFNA5 | M02641 | 2015-10-12 | 151 | 151 | in date | in date | in date | 93.0707 | no | 19894718 | low | 1031.818 | OK | 0.019111 | no | 93.4892 | no | 0.88194 | 0.8 | no | 0.90915 | no | 0.85233 | no | 0.95413 | no | 1.2493e-20 | 1.0755e-37 | 0 | -0.0019675 | yes | 8 | -0.0026129 | yes | OK | 0.45894 | yes | ['15M09548'] | [] | ['15M09548'] | NTC OK
150227_M00766_0079_000000000-ACPR9 | M00766 | 2015-02-27 | 75 | 75 | in date | in date | in date | 91.9723 | no | 29954337 | low | 1225.6634 | OK | 0.019386 | no | 91.9816 | no | 0.96464 | 0.85 | no | 0.97372 | no | 0.95482 | no | 0.84095 | no | 0.028842 | 2.0571e-07 | 0 | -0.00030049 | no | 0 | -0.00039375 | no | OK | 0.1812 | yes | no outlying indexes |  |  | NTC OK
150730_M00766_0120_000000000-AFN4E | M00766 | 2015-07-30 | 75 | 75 | in date | in date | in date | 94.0611 | no | 23945923 | low | 972.2454 | very low | 0.052603 | yes | 94.2795 | no | 0.9647 | 0.85 | no | 0.97668 | no | 0.95322 | no | 0.84095 | no | 0.0076657 | nan | 0 | -0.00038193 | no | 0 | -0.0003459 | no | OK | 0.34994 | yes | no outlying indexes |  |  | NTC OK
151001_M02641_0032_000000000-AGLE0 | M02641 | 2015-10-01 | 75 | 75 | in date | in date | in date | 87.3795 | no | 35725439 | low | 1504.6951 | very high | 0.017939 | no | 86.8081 | no | 0.9359 | 0.85 | no | 0.95659 | no | 0.91522 | no | 0.84095 | no | 5.6512e-08 | nan | 0 | -0.00058153 | yes | 0 | -0.00045814 | no | OK | 0.38331 | yes | no outlying indexes |  |  | NTC OK
151026_M00766_0148_000000000-AFLVV | M00766 | 2015-10-26 | 151 | 151 | in date | in date | in date | 88.3189 | no | 18895886 | low | 1013.3256 | OK | 0.032209 | no | 89.0093 | no | 0.93493 | 0.8 | no | 0.95661 | no | 0.91268 | no | 0.84095 | no | 9.4968e-15 | 5.2193e-112 | 0 | -0.00031813 | no | 0 | -0.00073043 | yes | OK | 0.28355 | yes | ['15M10912'] | [] | ['15M10912'] | NTC OK
150623_M00766_0116_000000000-AFLEW | M00766 | 2015-06-23 | 151 | 151 | in date | in date | in date | 87.7026 | no | 20733081 | low | 1104.0961 | OK | 0.034923 | no | 86.8671 | no | 0.84466 | 0.8 | no | 0.89266 | no | 0.79387 | yes | 0.84095 | no | 1.473e-43 | 7.7036e-97 | 0 | -0.0019268 | yes | 9 | -0.0023623 | yes | OK | 0.42771 | yes | no outlying indexes |  |  | NTC OK
151019_M02641_0040_000000000-AJDVH | M02641 | 2015-10-19 | 151 | 151 | in date | in date | in date | 90.6324 | no | 24425659 | OK | 1287.4209 | very high | 0.019011 | no | 90.7816 | no | 0.92696 | 0.8 | no | 0.94837 | no | 0.90505 | no | 0.95413 | no | 6.5136e-18 | nan | 0 | -0.00048911 | no | 0 | -0.00086419 | yes | OK | 0.45517 | yes | ['15M10923'] | ['15M10923'] | [] | NTC OK
150612_M02641_0061_000000000-AFMVT | M02641 | 2015-06-12 | 75 | 75 | in date | in date | in date | 79.341 | yes | 25070490 | low | 1016.9582 | very low | 0.03704 | no | 80.1022 | yes | 0.90347 | 0.85 | no | 0.92645 | no | 0.88019 | no | 0.84095 | no | 3.0151e-05 | 3.0045e-11 | 0 | -0.00078584 | yes | 0 | -0.00069582 | yes | OK | 0.44048 | yes | ['15M04137'] | [] | ['15M04137'] | NTC OK
150410_M00766_0089_000000000-ACBU4 | M00766 | 2015-04-10 | 151 | 151 | in date | in date | out of date | 94.4563 | no | 16861157 | low | 867.0412 | low | 0.021851 | no | 94.4152 | no | 0.93636 | 0.8 | no | 0.9672 | no | 0.90877 | no | 0.50766 | no | 1.738e-17 | nan | 0 | -0.00043159 | no | 0 | -0.001257 | yes | OK | 0.24869 | yes | ['14M04128'] | ['14M04128'] | [] | NTC OK
160127_M00766_0001_000000000-ALR2L | M00766 | 2016-01-27 | 151 | 151 | in date | in date | in date | 94.4219 | no | 16728171 | low | 869.4499 | low | 0.024493 | no | 94.8131 | no | 0.94892 | 0.8 | no | 0.95934 | no | 0.93843 | no | 0.95413 | no | 5.3787e-06 | nan | 0 | -0.00045839 | no | 0 | -0.00066696 | yes | OK | 0.97347 | yes | no outlying indexes |  |  | NTC OK
150622_M00766_0115_000000000-AFN2H | M00766 | 2015-06-22 | 75 | 75 | in date | in date | in date | 93.2875 | no | 26234040 | low | 1075.6113 | very low | 0.020578 | no | 93.4073 | no | 0.96472 | 0.85 | no | 0.97473 | no | 0.95557 | no | 0.84095 | no | 0.048507 | 0.0078432 | 0 | -0.00038391 | no | 0 | -0.0002776 | no | OK | 0.38333 | yes | ['15M05252'] | [] | ['15M05252'] | NTC OK
140205_M00766_0021_000000000-A7RVR | M00766 | 2014-02-05 | 251 | 251 | in date | in date | in date | 96.2303 | no | 10640092 | low | 549.8594 | very low | 0.015526 | no | 96.2427 | no | 0.77554 | 0.75 | no | 0.82459 | no | 0.7232 | yes | 0.84095 | no | 3.5805e-28 | 1.8115e-29 | 1 | -0.0023111 | yes | 65 | -0.0033538 | yes | High | 0.11219 | yes | no outlying indexes |  |  | No NTC on run
150902_M00766_0128_000000000-AFLDG | M00766 | 2015-09-02 | 151 | 151 | in date | in date | in date | 89.3025 | no | 21077655 | low | 1129.9527 | OK | 0.015246 | no | 89.1889 | no | 0.83273 | 0.8 | no | 0.88185 | no | 0.78 | yes | 0.67794 | no | 7.5008e-47 | 3.9097e-271 | 2 | -0.0022787 | yes | 13 | -0.0028073 | yes | OK | 0.46847 | yes | no outlying indexes |  |  | NTC OK
150407_M00766_0088_000000000-ACCAT | M00766 | 2015-04-07 | 151 | 151 | in date | in date | in date | 86.0053 | no | 20081955 | low | 1075.4261 | OK | 0.031951 | no | 86.1466 | no | 0.8006 | 0.8 | no | 0.8519 | no | 0.74788 | yes | 0.95413 | no | 6.3802e-32 | nan | 0 | -0.002369 | yes | 12 | -0.001854 | yes | OK | 0.62039 | yes | no outlying indexes |  |  | NTC OK
150619_M00766_0114_000000000-AFMYN | M00766 | 2015-06-19 | 75 | 75 | in date | in date | in date | 92.1565 | no | 28125244 | low | 1158.4526 | low | 0.034411 | no | 93.2083 | no | 0.95863 | 0.85 | no | 0.97242 | no | 0.94406 | no | 0.84095 | no | 0.00010309 | nan | 0 | -0.00041056 | no | 0 | -0.00058758 | yes | OK | 0.33916 | yes | no outlying indexes |  |  | NTC OK
151102_M02641_0045_000000000-AFN3H | M02641 | 2015-11-02 | 151 | 151 | in date | in date | in date | 94.4283 | no | 15063309 | low | 779.0627 | very low | 0.048767 | no | 94.8292 | no | 0.95449 | 0.8 | no | 0.96887 | no | 0.93956 | no | 0.84095 | no | 2.4447e-05 | 2.8037e-38 | 0 | -0.00027595 | no | 0 | -0.00058311 | yes | OK | 0.18495 | yes | ['15M10676' '15M10427'] | ['15M10427'] | ['15M10676'] | NTC OK
150526_M02641_0053_000000000-AFHDV | M02641 | 2015-05-26 | 200 | 200 | in date | in date | in date | 48.4208 | yes | 7559009 | low | 310.8307 | very low | 0.22965 | yes | 50.5562 | yes | 0.64122 | 0.7 | yes | 0.51285 | yes | 0.75917 | no | 0.50766 | no | 1.0758e-55 | 2.1338e-299 | 146 | -0.0039621 | yes | 1 | -0.0018861 | yes | OK | 0.33972 | yes | ['14M12202'] | [] | ['14M12202'] | NTC OK
150612_M00766_0112_000000000-AFMW5 | M00766 | 2015-06-12 | 75 | 75 | in date | in date | in date | 85.2766 | no | 34132691 | low | 1448.1255 | high | 0.019585 | no | 85.0541 | no | 0.93156 | 0.85 | no | 0.95081 | no | 0.91105 | no | 0.84095 | no | 3.193e-06 | 6.9549e-35 | 0 | -0.0006103 | yes | 0 | -0.00064344 | yes | OK | 0.42363 | yes | no outlying indexes |  |  | NTC OK
150904_M02641_0022_000000000-AH991 | M02641 | 2015-09-04 | 200 | 200 | in date | in date | in date | 90.7156 | no | 32586889 | low | 1348.7337 | OK | 0.0084128 | no | 90.8839 | no | 0.80842 | 0.7 | no | 0.83792 | no | 0.774 | no | 0.50766 | no | 6.5299e-46 | 4.3193e-60 | 33 | -0.002385 | yes | 45 | -0.0029542 | yes | OK | 0.26214 | yes | ['15M05232' '15M02845'] | [] | ['15M05232' '15M02845'] | NTC OK
151106_M00766_0152_000000000-AJ5W5 | M00766 | 2015-11-06 | 151 | 151 | in date | in date | in date | 89.6104 | no | 16168604 | low | 863.0096 | low | 0.024136 | no | 90.7829 | no | 0.9317 | 0.8 | no | 0.95041 | no | 0.91232 | no | 0.95413 | no | 3.0502e-07 | 1.2764e-25 | 0 | -0.00049715 | no | 0 | -0.00054569 | yes | OK | 0.26734 | yes | ['15M11860' '15M11870'] | ['15M11860'] | ['15M11870'] | NTC OK
150112_M00766_0068_000000000-A8PD8 | M00766 | 2015-01-12 | 151 | 151 | in date | in date | in date | 94.817 | no | 15773684 | low | 797.5388 | very low | 0.028582 | no | 94.9895 | no | 0.92451 | 0.8 | no | 0.93876 | no | 0.90863 | no | 0.99584 | no | 0.00017877 | 1.2387e-09 | 0 | -0.00090845 | yes | 0 | -0.00084942 | yes | OK | 0.20577 | yes | no outlying indexes |  |  | NTC OK
140905_M02641_0017_000000000-AA3FN | M02641 | 2014-09-05 | 75 | 75 | in date | in date | in date | 90.0019 | no | 23949422 | low | 1015.424 | very low | 0.055448 | yes | 89.6438 | no | 0.95625 | 0.85 | no | 0.96176 | no | 0.95019 | no | 0.99584 | no | 0.99831 | nan | 0 | -0.00042771 | no | 0 | -0.00021245 | no | OK | 0.47404 | yes | ['14M07917'] | [] | ['14M07917'] | NTC OK
151009_M02641_0036_000000000-AFN30 | M02641 | 2015-10-09 | 151 | 151 | in date | in date | in date | 91.4518 | no | 21458630 | low | 1123.3539 | OK | 0.024973 | no | 91.2766 | no | 0.86713 | 0.8 | no | 0.89737 | no | 0.83419 | no | 0.95413 | no | 2.2346e-20 | 0 | 0 | -0.0020801 | yes | 10 | -0.0026428 | yes | OK | 0.89911 | yes | ['15M09439' '15M09554'] | [] | ['15M09439' '15M09554'] | NTC OK
140715_M02641_0008_000000000-A8P81 | M02641 | 2014-07-15 | 151 | 151 | in date | in date | in date | 98.3026 | no | 5446264 | low | 277.542 | very low | 0.020978 | no | 98.4264 | no | 0.94495 | 0.8 | no | 0.95412 | no | 0.93542 | no | 0.99584 | no | 0.593 | 0.077423 | 0 | -0.00077434 | yes | 0 | -0.00075567 | yes | OK | 0.23316 | yes | ['14M05256'] | [] | ['14M05256'] | NTC OK
160108_M02641_0064_000000000-ALNH2 | M02641 | 2016-01-08 | 200 | 200 | in date | in date | in date | 91.9702 | no | 21242610 | low | 1109.3897 | OK | 0.022049 | no | 92.3443 | no | 0.8579 | 0.75 | no | 0.90283 | no | 0.81012 | no | 0.67794 | no | 5.5542e-59 | nan | 0 | -0.0013823 | yes | 1 | -0.0025063 | yes | OK | 0.31714 | yes | ['15M11705'] | [] | ['15M11705'] | NTC OK
141125_M00766_0062_000000000-A7BKW | M00766 | 2014-11-25 | 151 | 151 | out of date | in date | in date | 94.983 | no | 14809543 | low | 762.9367 | very low | 0.018589 | no | 95.0964 | no | 0.89379 | 0.8 | no | 0.96312 | no | 0.82977 | no | 0.24079 | no | 2.4117e-86 | nan | 0 | -0.00041911 | no | 4 | -0.00042946 | no | OK | 0.33497 | yes | no outlying indexes |  |  | NTC OK
141222_M00766_0067_000000000-ACCB3 | M00766 | 2014-12-22 | 151 | 151 | in date | in date | in date | 67.7653 | yes | 14253970 | low | 760.392 | very low | 0.029815 | no | 68.1874 | yes | 0.82139 | 0.8 | no | 0.85456 | no | 0.78426 | yes | 0.95413 | no | 3.8283e-06 | nan | 0 | -0.0021203 | yes | 11 | -0.0013537 | yes | OK | 0.27283 | yes | no outlying indexes |  |  | NTC OK
160204_M02641_0072_000000000-AK6CE | M02641 | 2016-02-04 | 75 | 75 | in date | in date | in date | 91.6672 | no | 25918178 | low | 1060.0101 | very low | 0.015091 | no | 91.5669 | no | 0.95246 | 0.85 | no | 0.96475 | no | 0.93912 | no | 0.84095 | no | 0.00023857 | 2.281e-41 | 0 | -0.00053931 | yes | 0 | -0.00055384 | yes | OK | 0.32617 | yes | no outlying indexes |  |  | NTC OK
150515_M00766_0098_000000000-AF9ND | M00766 | 2015-05-15 | 75 | 75 | in date | in date | in date | 96.8787 | no | 15942131 | low | 638.6353 | very low | 0.030068 | no | 96.9762 | no | 0.97995 | 0.85 | no | 0.98705 | no | 0.97446 | no | 0.84095 | no | 0.57572 | nan | 0 | -0.00024399 | no | 0 | -0.00016041 | no | OK | 0.33654 | yes | no outlying indexes |  |  | NTC OK
160104_M02641_0062_000000000-AL603 | M02641 | 2016-01-04 | 101 | 101 | in date | in date | in date | 93.4734 | no | 19858211 | low | 1030.1337 | OK | 0.01816 | no | 93.4478 | no | 0.95548 | 0.8 | no | 0.97027 | no | 0.94346 | no | 0.84095 | no | 1.6436e-05 | 1.1366e-31 | 0 | -0.00025903 | no | 0 | -0.00019241 | no | OK | 0.31626 | yes | ['15M11705'] | [] | ['15M11705'] | NTC OK
160208_M00766_0006_000000000-AMF4G | M00766 | 2016-02-08 | 75 | 75 | in date | in date | in date | 85.7346 | no | 32902642 | low | 1404.3574 | high | 0.015626 | no | 85.8175 | no | 0.93494 | 0.85 | no | 0.95198 | no | 0.91641 | no | 0.84095 | no | 1.9762e-05 | 5.3645e-08 | 0 | -0.00062874 | yes | 0 | -0.00051979 | yes | OK | 0.40774 | yes | ['16M00692'] | ['16M00692'] | [] | NTC OK
150914_M00766_0132_000000000-AF41F | M00766 | 2015-09-14 | 151 | 151 | in date | in date | in date | 86.8892 | no | 3991745 | low | 184.4777 | very low | 0.012611 | no | 87.677 | no | 0.95124 | 0.8 | no | 0.96221 | no | 0.94056 | no | 0.95413 | no | 0.17881 | 0.0092674 | 0 | -0.00035999 | no | 0 | -0.00086178 | yes | OK | 0.34857 | yes | ['15M07808'] | ['15M07808'] | [] | NTC OK
141113_M02641_0025_000000000-A8RTY | M02641 | 2014-11-13 | 151 | 151 | in date | in date | in date | 97.3253 | no | 9834883 | low | 502.4441 | very low | 0.049911 | no | 97.4887 | no | 0.91483 | 0.8 | no | 0.92802 | no | 0.90038 | no | 0.95413 | no | 0.00015291 | 0 | 0 | -0.00132 | yes | 0 | -0.0014718 | yes | OK | 0.73682 | yes | no outlying indexes |  |  | NTC OK
150521_M00766_0102_000000000-AF9MV | M00766 | 2015-05-21 | 75 | 75 | in date | in date | in date | 89.6458 | no | 30563261 | low | 1273.6592 | OK | 0.043729 | no | 90.9553 | no | 0.95232 | 0.85 | no | 0.96492 | no | 0.93927 | no | 0.84095 | no | 0.0097477 | nan | 0 | -0.00050706 | yes | 0 | -0.00046026 | no | OK | 0.3581 | yes | ['14M09383'] | [] | ['14M09383'] | NTC OK
150731_M02641_0009_000000000-AFN3M | M02641 | 2015-07-31 | 75 | 75 | in date | in date | in date | 90.2753 | no | 32967275 | low | 1364.8961 | OK | 0.020875 | no | 90.6855 | no | 0.9485 | 0.85 | no | 0.96322 | no | 0.93353 | no | 0.84095 | no | 0.00029106 | 1.2672e-06 | 0 | -0.00052951 | yes | 0 | -0.00044329 | no | OK | 0.29214 | yes | no outlying indexes |  |  | NTC OK
150819_M02641_0016_000000000-AGJHU | M02641 | 2015-08-19 | 75 | 75 | in date | in date | in date | 94.1214 | no | 26775268 | low | 1083.954 | very low | 0.032934 | no | 94.4666 | no | 0.96462 | 0.85 | no | 0.97519 | no | 0.95489 | no | 0.84095 | no | 0.013871 | 2.5966e-09 | 0 | -0.00041278 | no | 0 | -0.00025665 | no | OK | 0.35264 | yes | no outlying indexes |  |  | NTC OK
150529_M00766_0106_000000000-AEVP5 | M00766 | 2015-05-29 | 151 | 151 | in date | in date | in date | 88.3589 | no | 15502868 | low | 806.7578 | very low | 0.032653 | no | 89.3692 | no | 0.91331 | 0.8 | no | 0.92662 | no | 0.89824 | no | 0.99584 | no | 0.34729 | 5.8531e-07 | 0 | -0.00082214 | yes | 0 | -0.0010282 | yes | OK | 0.32149 | yes | ['13M70706' '13M70790'] | ['13M70706' '13M70790'] | [] | NTC OK
140625_M00766_0041_000000000-A8P78 | M00766 | 2014-06-25 | 251 | 251 | in date | in date | in date | 93.6991 | no | 12306226 | low | 638.3155 | very low | 0.015947 | no | 93.5516 | no | 0.80794 | 0.75 | no | 0.83721 | no | 0.77401 | no | 0.95413 | no | 1.446e-13 | nan | 15 | -0.0023266 | yes | 34 | -0.0029216 | yes | OK | 0.12175 | yes | ['12M01046'] | ['12M01046'] | [] | NTC OK
150501_M02641_0045_000000000-AD6UA | M02641 | 2015-05-01 | 75 | 75 | in date | in date | in date | 80.8519 | no | 25913017 | low | 1131.0451 | low | 0.015126 | no | 81.3041 | yes | 0.93089 | 0.85 | no | 0.92992 | no | 0.93269 | no | 0.99998 | no | 0.013057 | nan | 0 | -0.00054803 | yes | 0 | -0.00041428 | no | OK | 0.4481 | yes | no outlying indexes |  |  | NTC OK
130508_M00766_0013_000000000-A3RB5 | M00766 | 2013-05-08 | 151 | 151 | in date | in date | in date | 89.695 | no | 17060727 | low | 897.222 | low | 0.013058 | no | 89.6111 | no | 0.92341 | 0.8 | no | 0.94982 | no | 0.89785 | no | 0.67794 | no | 1.7122e-13 | 9.1751e-65 | 0 | -0.00060464 | yes | 0 | -0.00096571 | yes | High | 0.15582 | yes | no outlying indexes |  |  | No NTC on run
150928_M02641_0029_000000000-AGHV5 | M02641 | 2015-09-28 | 75 | 75 | in date | in date | in date | 93.773 | no | 28292454 | low | 1144.9595 | low | 0.021423 | no | 93.5295 | no | 0.96266 | 0.85 | no | 0.97493 | no | 0.95191 | no | 0.84095 | no | 0.012643 | nan | 0 | -0.00041356 | no | 0 | -0.00026611 | no | OK | 0.2835 | yes | no outlying indexes |  |  | NTC OK
150821_M02641_0018_000000000-AGJNU | M02641 | 2015-08-21 | 151 | 151 | in date | in date | in date | 96.3232 | no | 11702340 | low | 597.5554 | very low | 0.030041 | no | 96.5218 | no | 0.96101 | 0.8 | no | 0.97032 | no | 0.95086 | no | 0.95413 | no | 0.0070316 | nan | 0 | -0.00034988 | no | 0 | -0.000607 | yes | OK | 0.23539 | yes | no outlying indexes |  |  | NTC OK
150306_M00766_0080_000000000-ACNEP | M00766 | 2015-03-06 | 75 | 75 | in date | in date | in date | 94.106 | no | 26017554 | low | 1059.9167 | very low | 0.02142 | no | 94.169 | no | 0.97008 | 0.85 | no | 0.97849 | no | 0.96173 | no | 0.84095 | no | 0.078515 | 0.00029715 | 0 | -0.00028625 | no | 0 | -0.00023866 | no | OK | 0.20405 | yes | ['13M11429'] | ['13M11429'] | [] | NTC OK
150302_M02641_0033_000000000-ACC9K | M02641 | 2015-03-02 | 151 | 151 | in date | in date | in date | 86.4847 | no | 20383236 | low | 1091.5584 | OK | 0.0091317 | no | 86.4229 | no | 0.8794 | 0.8 | no | 0.92333 | no | 0.83429 | no | 0.95413 | no | 4.0711e-39 | 8.175e-86 | 0 | -0.00085422 | yes | 0 | -0.00080079 | yes | OK | 0.50664 | yes | no outlying indexes |  |  | NTC OK
| 63 |  |  |  | 3 | 0 | 4 |  | 11 |  | 120 |  | 33 |  | 10 |  | 19 |  |  | 9

## AnalysisDataTestSorted
RunID | InstrumentID | RunDate | NumberOfCyclesR1 | NumberofCyclesR2 | FlowCellInDate? | PR2BottleInDate? | ReagentKitInDate? | %ReadsPassingFilter | %ReadsPassingFilterOutsideThreshold? | NumberofReadsPassingFilter | ReadCountWithinThreshold? | MeanClusterDensityPassingFilter | MeanClusterDensityWithinThreshold | COVClusterDensity | COVClusterDensityExceedsThreshold? | %MedianClusterDensityPassingFilter | GapBetweenClusterDensitiesPassingFilterAndNotExceedsThreshold? | ProportionOverQ30 | IlluminaQ30Threshold | ProportionOverQ30ExceedsThreshold? | ProportionOverQ30Read1 | ProportionOverQ30ExceedsThresholdRead1 | ProportionOverQ30Read2 | ProportionOverQ30ExceedsThresholdRead2 | K-SPValue | K-SPValueSignificant? | K-SUnpackPvalue | M-WUnpackPvalue | NumberOfCyclesExceedingThresholdUnderQ30Read1 | SlopeOfLineOfBestFitRead1 | SlopeOfLineOfBestFitRead1ExceedsThreshold? | NumberOfCyclesExceedingThresholdUnderQ30Read2 | SlopeOfLineOfBestFitRead2 | SlopeOfLineOfBestFitRead2ExceedsThreshold? | ProportionOfUndeterminedReads? | COVIndexes | COVIndexesExceedsThreshold? | OutlyingIndexes | OutlyingIndexes(Low) | OutlyingIndexes(High) | StateOfNTC
130417_M00766_0007_000000000-A3R1F | M00766 | 2013-04-17 | 151 | 151 | in date | in date | in date | 93.6874 | no | 14918743 | low | 773.0359 | very low | 0.024644 | no | 93.6673 | no | 0.97122 | 0.8 | no | 0.98017 | no | 0.96226 | no | 0.84095 | no | 0.0022684 | 4.2948e-06 | 0 | -0.00013884 | no | 0 | -0.00025711 | no | No index information for 130417_M00766_0007_000000000-A3R1F
130508_M00766_0013_000000000-A3RB5 | M00766 | 2013-05-08 | 151 | 151 | in date | in date | in date | 89.695 | no | 17060727 | low | 897.222 | low | 0.013058 | no | 89.6111 | no | 0.92341 | 0.8 | no | 0.94982 | no | 0.89785 | no | 0.67794 | no | 1.7122e-13 | 9.1751e-65 | 0 | -0.00060464 | yes | 0 | -0.00096571 | yes | High | 0.15582 | yes | no outlying indexes |  |  | No NTC on run
130510_M00766_0014_000000000-A3PFJ | M00766 | 2013-05-10 | 156 | 151 | in date | in date | in date | 88.5142 | no | 18378480 | low | 995.2515 | low | 0.013484 | no | 88.6611 | no | 0.95249 | 0.75 | no | 0.96867 | no | 0.93502 | no | 0.95413 | no | 3.3216e-08 | nan | 0 | -0.0002029 | no | 0 | -0.00047138 | no | OK | 0.51666 | yes | no outlying indexes |  |  | No NTC on run
130613_M00766_0018_000000000-A4FN3 | M00766 | 2013-06-13 | 251 | 251 | in date | in date | in date | 62.2349 | yes | 14759376 | low | 877.4421 | low | 0.0052548 | no | 61.6486 | yes | 0.68299 | 0.75 | yes | 0.7656 | no | 0.58473 | yes | 0.67794 | no | 5.4172e-57 | 1.5425e-86 | 28 | -0.0020629 | yes | 55 | -0.0033012 | yes | OK | 0.15 | yes | no outlying indexes |  |  | No NTC on run
130708_M00766_0023_000000000-A20R8 | M00766 | 2013-07-08 | 151 | 151 | out of date | in date | out of date | 90.9082 | no | 8144995 | None | 872.0913 | None | 0.015455 | no | 90.7757 | no | 0.86589 | 0.85 | no | 0.91799 | no | 0.81047 | yes | 0.67794 | no | 9.6929e-19 | 6.9078e-114 | 0 | -0.0013571 | yes | 0 | -0.0030958 | yes | OK | 0.10661 | yes | no outlying indexes |  |  | No NTC on run
130902_M00766_0045_000000000-A5BVD | M00766 | 2013-09-02 | 156 | 151 | in date | in date | in date | 92.4439 | no | 17107054 | low | 889.9218 | low | 0.020943 | no | 92.3467 | no | 0.95012 | 0.75 | no | 0.96778 | no | 0.93094 | no | 0.84095 | no | 1.4736e-06 | nan | 0 | -0.00028935 | no | 1 | -0.00072347 | yes | OK | 0.59937 | yes | no outlying indexes |  |  | No NTC on run
130909_M00766_0047_000000000-A5PEH | M00766 | 2013-09-09 | 151 | 151 | in date | in date | in date | 91.8202 | no | 17818894 | low | 943.064 | low | 0.013531 | no | 91.9043 | no | 0.91659 | 0.8 | no | 0.93844 | no | 0.89323 | no | 0.95413 | no | 2.3926e-11 | 3.6095e-18 | 0 | -0.00083727 | yes | 0 | -0.00095948 | yes | OK | 0.36674 | yes | ['Control'] | ['Control'] | [] | No NTC on run
140205_M00766_0021_000000000-A7RVR | M00766 | 2014-02-05 | 251 | 251 | in date | in date | in date | 96.2303 | no | 10640092 | low | 549.8594 | very low | 0.015526 | no | 96.2427 | no | 0.77554 | 0.75 | no | 0.82459 | no | 0.7232 | yes | 0.84095 | no | 3.5805e-28 | 1.8115e-29 | 1 | -0.0023111 | yes | 65 | -0.0033538 | yes | High | 0.11219 | yes | no outlying indexes |  |  | No NTC on run
140603_M00766_0037_000000000-A7WU9 | M00766 | 2014-06-03 | 251 | 251 | in date | in date | in date | 80.6392 | no | 20279789 | low | 1121.2903 | OK | 0.020185 | no | 81.2363 | yes | 0.67262 | 0.75 | yes | 0.87096 | no | 0.46478 | yes | 0.24079 | no | 0 | 0 | 0 | -0.0011435 | yes | 73 | -0.0033779 | yes | OK | 0.33214 | yes | no outlying indexes |  |  | No NTC on run
140620_M02641_0004_000000000-A8R7V | M02641 | 2014-06-20 | 151 | 151 | in date | in date | in date | 95.2324 | no | 15895393 | low | 815.6811 | very low | 0.010793 | no | 95.1913 | no | 0.92041 | 0.8 | no | 0.96394 | no | 0.87468 | no | 0.35842 | no | 5.5248e-61 | nan | 0 | -0.0004185 | no | 0 | -0.00070599 | yes | High | 0.41518 | yes | ['14M02611'] | ['14M02611'] | [] | NTC OK
140625_M00766_0041_000000000-A8P78 | M00766 | 2014-06-25 | 251 | 251 | in date | in date | in date | 93.6991 | no | 12306226 | low | 638.3155 | very low | 0.015947 | no | 93.5516 | no | 0.80794 | 0.75 | no | 0.83721 | no | 0.77401 | no | 0.95413 | no | 1.446e-13 | nan | 15 | -0.0023266 | yes | 34 | -0.0029216 | yes | OK | 0.12175 | yes | ['12M01046'] | ['12M01046'] | [] | NTC OK
140715_M02641_0008_000000000-A8P81 | M02641 | 2014-07-15 | 151 | 151 | in date | in date | in date | 98.3026 | no | 5446264 | low | 277.542 | very low | 0.020978 | no | 98.4264 | no | 0.94495 | 0.8 | no | 0.95412 | no | 0.93542 | no | 0.99584 | no | 0.593 | 0.077423 | 0 | -0.00077434 | yes | 0 | -0.00075567 | yes | OK | 0.23316 | yes | ['14M05256'] | [] | ['14M05256'] | NTC OK
140804_M02641_0011_000000000-AAD6A | M02641 | 2014-08-04 | 251 | 251 | in date | in date | in date | 92.8278 | no | 18236023 | low | 947.4631 | low | 0.014242 | no | 92.5473 | no | 0.75551 | 0.75 | no | 0.79909 | no | 0.70633 | yes | 0.84095 | no | 4.579e-31 | nan | 19 | -0.0025558 | yes | 41 | -0.0033381 | yes | OK | 0.15644 | yes | no outlying indexes |  |  | No NTC on run
140905_M02641_0017_000000000-AA3FN | M02641 | 2014-09-05 | 75 | 75 | in date | in date | in date | 90.0019 | no | 23949422 | low | 1015.424 | very low | 0.055448 | yes | 89.6438 | no | 0.95625 | 0.85 | no | 0.96176 | no | 0.95019 | no | 0.99584 | no | 0.99831 | nan | 0 | -0.00042771 | no | 0 | -0.00021245 | no | OK | 0.47404 | yes | ['14M07917'] | [] | ['14M07917'] | NTC OK
140917_M02641_0018_000000000-AA3H7 | M02641 | 2014-09-17 | 75 | 75 | in date | in date | in date | 91.5859 | no | 24528146 | low | 1028.5468 | very low | 0.035153 | no | 91.5237 | no | 0.96229 | 0.85 | no | 0.96999 | no | 0.95452 | no | 0.84095 | no | 0.88947 | nan | 0 | -0.00031376 | no | 0 | -0.00016688 | no | OK | 0.27666 | yes | ['14M00640'] | ['14M00640'] | [] | NTC OK
141024_M02641_0024_000000000-A8P84 | M02641 | 2014-10-24 | 151 | 151 | in date | in date | in date | 94.9748 | no | 15281418 | low | 794.7029 | very low | 0.031108 | no | 95.0899 | no | 0.94073 | 0.8 | no | 0.95744 | no | 0.92359 | no | 0.84095 | no | 1.9766e-08 | 4.7689e-23 | 0 | -0.00055662 | yes | 0 | -0.00062808 | yes | OK | 0.077865 | yes | no outlying indexes |  |  | NTC OK
141113_M02641_0025_000000000-A8RTY | M02641 | 2014-11-13 | 151 | 151 | in date | in date | in date | 97.3253 | no | 9834883 | low | 502.4441 | very low | 0.049911 | no | 97.4887 | no | 0.91483 | 0.8 | no | 0.92802 | no | 0.90038 | no | 0.95413 | no | 0.00015291 | 0 | 0 | -0.00132 | yes | 0 | -0.0014718 | yes | OK | 0.73682 | yes | no outlying indexes |  |  | NTC OK
141125_M00766_0062_000000000-A7BKW | M00766 | 2014-11-25 | 151 | 151 | out of date | in date | in date | 94.983 | no | 14809543 | low | 762.9367 | very low | 0.018589 | no | 95.0964 | no | 0.89379 | 0.8 | no | 0.96312 | no | 0.82977 | no | 0.24079 | no | 2.4117e-86 | nan | 0 | -0.00041911 | no | 4 | -0.00042946 | no | OK | 0.33497 | yes | no outlying indexes |  |  | NTC OK
141222_M00766_0067_000000000-ACCB3 | M00766 | 2014-12-22 | 151 | 151 | in date | in date | in date | 67.7653 | yes | 14253970 | low | 760.392 | very low | 0.029815 | no | 68.1874 | yes | 0.82139 | 0.8 | no | 0.85456 | no | 0.78426 | yes | 0.95413 | no | 3.8283e-06 | nan | 0 | -0.0021203 | yes | 11 | -0.0013537 | yes | OK | 0.27283 | yes | no outlying indexes |  |  | NTC OK
150112_M00766_0068_000000000-A8PD8 | M00766 | 2015-01-12 | 151 | 151 | in date | in date | in date | 94.817 | no | 15773684 | low | 797.5388 | very low | 0.028582 | no | 94.9895 | no | 0.92451 | 0.8 | no | 0.93876 | no | 0.90863 | no | 0.99584 | no | 0.00017877 | 1.2387e-09 | 0 | -0.00090845 | yes | 0 | -0.00084942 | yes | OK | 0.20577 | yes | no outlying indexes |  |  | NTC OK
150114_M00766_0069_000000000-AA68B | M00766 | 2015-01-14 | 151 | 151 | in date | in date | in date | 77.8262 | yes | 20704375 | low | 1148.0758 | OK | 0.024397 | no | 77.3315 | yes | 0.86381 | 0.8 | no | 0.91949 | no | 0.81333 | no | 0.95413 | no | 2.6748e-38 | nan | 0 | -0.00060681 | yes | 4 | -0.00057587 | yes | OK | 0.18923 | yes | ['14M09058'] | ['14M09058'] | [] | NTC OK
150116_M00766_0070_000000000-ACC3G | M00766 | 2015-01-16 | 151 | 151 | in date | in date | in date | 86.297 | no | 14147272 | low | 752.8015 | very low | 0.040986 | no | 86.2322 | no | 0.93362 | 0.8 | no | 0.95795 | no | 0.91769 | no | 0.95413 | no | 0.027224 | 4.5409e-07 | 0 | -0.00014326 | no | 3 | 0.00022032 | no | OK | 0.37649 | yes | ['14M11762-2' '3M0956-2' '3M0956' '14M11762'] | ['14M11762-2' '3M0956-2' '3M0956' '14M11762'] | [] | NTC OK
150127_M00766_0072_000000000-ABRK7 | M00766 | 2015-01-27 | 151 | 151 | in date | in date | out of date | 94.2571 | no | 17854832 | low | 926.0782 | low | 0.029511 | no | 94.3435 | no | 0.93257 | 0.8 | no | 0.95668 | no | 0.90691 | no | 0.67794 | no | 1.1008e-18 | 1.0313e-37 | 0 | -0.00054076 | yes | 0 | -0.00067649 | yes | OK | 0.23558 | yes | ['14M12186'] | [] | ['14M12186'] | NTC OK
150130_M00766_0073_000000000-ACBVB | M00766 | 2015-01-30 | 251 | 251 | in date | in date | in date | 81.2781 | no | 23929178 | low | 1312.9896 | very high | 0.017972 | no | 81.1332 | yes | 0.66236 | 0.75 | yes | 0.70216 | yes | 0.61962 | yes | 0.84095 | no | 6.7473e-31 | nan | 47 | -0.0033611 | yes | 64 | -0.0034493 | yes | OK | 0.3138 | yes | no outlying indexes |  |  | No NTC on run
150205_M02641_0029_000000000-ACCE2 | M02641 | 2015-02-05 | 151 | 151 | in date | in date | in date | 87.5711 | no | 18965301 | low | 1006.645 | OK | 0.019074 | no | 87.6822 | no | 0.8298 | 0.8 | no | 0.8663 | no | 0.79202 | yes | 0.99584 | no | 1.5184e-12 | 8.2211e-115 | 0 | -0.0022136 | yes | 2 | -0.0018972 | yes | OK | 0.25963 | yes | ['13M12234' '15M00575'] | ['13M12234' '15M00575'] | [] | NTC OK
150220_M00766_0077_000000000-AA2U8 | M00766 | 2015-02-20 | 26 | 26 | in date | in date | in date | 11.2989 | yes | 3422864 | low | 175.3978 | very low | 0.069387 | yes | 0 | yes | 0.63211 | 0.9 | yes | 0.88444 | yes | 0.47035 | yes | 0.95413 | no | 6.5681e-12 | 7.5094e-13 | 0 | 0 | no | 26 | 0 | no | High | 0.6497 | yes | no outlying indexes |  |  | No NTC on run
150227_M00766_0079_000000000-ACPR9 | M00766 | 2015-02-27 | 75 | 75 | in date | in date | in date | 91.9723 | no | 29954337 | low | 1225.6634 | OK | 0.019386 | no | 91.9816 | no | 0.96464 | 0.85 | no | 0.97372 | no | 0.95482 | no | 0.84095 | no | 0.028842 | 2.0571e-07 | 0 | -0.00030049 | no | 0 | -0.00039375 | no | OK | 0.1812 | yes | no outlying indexes |  |  | NTC OK
150227_M02641_0032_000000000-ACNF3 | M02641 | 2015-02-27 | 75 | 75 | in date | in date | in date | 90.2941 | no | 25317785 | low | 1066.3007 | very low | 0.076375 | yes | 90.5194 | no | 0.95732 | 0.85 | no | 0.9695 | no | 0.94388 | no | 0.84095 | no | 0.0023737 | nan | 0 | -0.00027283 | no | 0 | -0.00038603 | no | OK | 0.80447 | yes | ['14M08864'] | [] | ['14M08864'] | NTC OK
150302_M02641_0033_000000000-ACC9K | M02641 | 2015-03-02 | 151 | 151 | in date | in date | in date | 86.4847 | no | 20383236 | low | 1091.5584 | OK | 0.0091317 | no | 86.4229 | no | 0.8794 | 0.8 | no | 0.92333 | no | 0.83429 | no | 0.95413 | no | 4.0711e-39 | 8.175e-86 | 0 | -0.00085422 | yes | 0 | -0.00080079 | yes | OK | 0.50664 | yes | no outlying indexes |  |  | NTC OK
150306_M00766_0080_000000000-ACNEP | M00766 | 2015-03-06 | 75 | 75 | in date | in date | in date | 94.106 | no | 26017554 | low | 1059.9167 | very low | 0.02142 | no | 94.169 | no | 0.97008 | 0.85 | no | 0.97849 | no | 0.96173 | no | 0.84095 | no | 0.078515 | 0.00029715 | 0 | -0.00028625 | no | 0 | -0.00023866 | no | OK | 0.20405 | yes | ['13M11429'] | ['13M11429'] | [] | NTC OK
150306_M02641_0034_000000000-ACRUF | M02641 | 2015-03-06 | 75 | 75 | in date | in date | in date | 91.3955 | no | 24975924 | low | 1049.1583 | very low | 0.0067808 | no | 91.297 | no | 0.95124 | 0.85 | no | 0.9701 | no | 0.93153 | no | 0.95413 | no | 3.4327e-07 | nan | 0 | -0.00040766 | no | 0 | -0.00082529 | yes | OK | 0.17399 | yes | no outlying indexes |  |  | NTC OK
150313_M02641_0037_000000000-ACC8R | M02641 | 2015-03-13 | 151 | 151 | in date | in date | in date | 94.5096 | no | 17647098 | low | 906.6463 | low | 0.017228 | no | 93.7282 | no | 0.94213 | 0.8 | no | 0.9595 | no | 0.9242 | no | 0.67794 | no | 2.7478e-10 | nan | 0 | -0.00051188 | yes | 0 | -0.00060154 | yes | OK | 0.11578 | yes | no outlying indexes |  |  | NTC OK
150325_M00766_0084_000000000-AD8KN | M00766 | 2015-03-25 | 75 | 75 | in date | in date | in date | 93.0443 | no | 26726137 | low | 1096.1198 | very low | 0.025085 | no | 93.5238 | no | 0.96479 | 0.85 | no | 0.97449 | no | 0.95451 | no | 0.84095 | no | 0.047651 | 0.00076142 | 0 | -0.00034315 | no | 0 | -0.00025609 | no | OK | 0.617 | yes | no outlying indexes |  |  | NTC OK
150325_M02641_0038_000000000-AD8KY | M02641 | 2015-03-25 | 75 | 75 | in date | in date | in date | 92.8928 | no | 28292154 | low | 1150.8931 | low | 0.016396 | no | 92.9934 | no | 0.96863 | 0.85 | no | 0.97441 | no | 0.9626 | no | 0.95413 | no | 0.97963 | nan | 0 | -0.00030202 | no | 0 | -0.00017631 | no | OK | 0.27087 | yes | no outlying indexes |  |  | NTC OK
150401_M00766_0086_000000000-ACBU5 | M00766 | 2015-04-01 | 151 | 151 | in date | in date | in date | 90.4494 | no | 18801773 | low | 986.6817 | low | 0.029182 | no | 90.7591 | no | 0.8401 | 0.8 | no | 0.87195 | no | 0.80659 | no | 0.99584 | no | 6.8105e-11 | 1.2646e-33 | 0 | -0.0022957 | yes | 6 | -0.0022224 | yes | High | 2.2561 | yes | ['15M00594' '15M00616'] | [] | ['15M00594' '15M00616'] | NTC OK
150407_M00766_0088_000000000-ACCAT | M00766 | 2015-04-07 | 151 | 151 | in date | in date | in date | 86.0053 | no | 20081955 | low | 1075.4261 | OK | 0.031951 | no | 86.1466 | no | 0.8006 | 0.8 | no | 0.8519 | no | 0.74788 | yes | 0.95413 | no | 6.3802e-32 | nan | 0 | -0.002369 | yes | 12 | -0.001854 | yes | OK | 0.62039 | yes | no outlying indexes |  |  | NTC OK
150410_M00766_0089_000000000-ACBU4 | M00766 | 2015-04-10 | 151 | 151 | in date | in date | out of date | 94.4563 | no | 16861157 | low | 867.0412 | low | 0.021851 | no | 94.4152 | no | 0.93636 | 0.8 | no | 0.9672 | no | 0.90877 | no | 0.50766 | no | 1.738e-17 | nan | 0 | -0.00043159 | no | 0 | -0.001257 | yes | OK | 0.24869 | yes | ['14M04128'] | ['14M04128'] | [] | NTC OK
150423_M02641_0041_000000000-ACN0P | M02641 | 2015-04-23 | 75 | 75 | in date | in date | in date | 91.9691 | no | 30009055 | low | 1228.4129 | OK | 0.027957 | no | 92.2905 | no | 0.9501 | 0.85 | no | 0.96725 | no | 0.93306 | no | 0.84095 | no | 2.6066e-06 | 3.4467e-12 | 0 | -0.00047255 | no | 0 | -0.00049099 | no | OK | 0.53918 | yes | ['15M03290'] | [] | ['15M03290'] | NTC OK
150429_M02641_0043_000000000-AD8KJ | M02641 | 2015-04-29 | 75 | 75 | in date | in date | in date | 77.8177 | yes | 32906358 | low | 1415.0412 | high | 0.018771 | no | 78.178 | yes | 0.8983 | 0.85 | no | 0.92026 | no | 0.87594 | no | 0.84095 | no | 0.00033968 | nan | 0 | -0.00089215 | yes | 0 | -0.00067513 | yes | OK | 0.50312 | yes | no outlying indexes |  |  | NTC OK
150430_M02641_0044_000000000-AD8VC | M02641 | 2015-04-30 | 75 | 75 | in date | in date | in date | 86.9368 | no | 27790413 | low | 1196.1245 | low | 0.022911 | no | 87.4623 | no | 0.92905 | 0.85 | no | 0.95387 | no | 0.90368 | no | 0.84095 | no | 1.2394e-10 | 1.7232e-11 | 0 | -0.00060018 | yes | 0 | -0.0005903 | yes | OK | 0.4048 | yes | no outlying indexes |  |  | NTC OK
150501_M00766_0093_000000000-AF9MH | M00766 | 2015-05-01 | 75 | 75 | in date | in date | in date | 87.2253 | no | 34467195 | low | 1450.6461 | high | 0.022012 | no | 87.0197 | no | 0.94139 | 0.85 | no | 0.95639 | no | 0.92569 | no | 0.84095 | no | 0.00057945 | 2.9759e-05 | 0 | -0.00056762 | yes | 0 | -0.00052432 | yes | OK | 0.38312 | yes | ['14M11012'] | [] | ['14M11012'] | NTC OK
150501_M02641_0045_000000000-AD6UA | M02641 | 2015-05-01 | 75 | 75 | in date | in date | in date | 80.8519 | no | 25913017 | low | 1131.0451 | low | 0.015126 | no | 81.3041 | yes | 0.93089 | 0.85 | no | 0.92992 | no | 0.93269 | no | 0.99998 | no | 0.013057 | nan | 0 | -0.00054803 | yes | 0 | -0.00041428 | no | OK | 0.4481 | yes | no outlying indexes |  |  | NTC OK
150507_M00766_0095_000000000-AF7N3 | M00766 | 2015-05-07 | 75 | 75 | in date | in date | in date | 92.6658 | no | 27410104 | low | 1119.7122 | low | 0.036603 | no | 93.5327 | no | 0.96159 | 0.85 | no | 0.9741 | no | 0.94953 | no | 0.84095 | no | 0.00099419 | 8.626e-08 | 0 | -0.00036389 | no | 0 | -0.00042845 | no | OK | 0.29241 | yes | ['14M10600'] | [] | ['14M10600'] | NTC OK
150507_M02641_0046_000000000-AF7N5 | M02641 | 2015-05-07 | 75 | 75 | in date | in date | in date | 92.9775 | no | 27595623 | low | 1124.3595 | low | 0.040298 | no | 93.1424 | no | 0.95227 | 0.85 | no | 0.96981 | no | 0.93556 | no | 0.84095 | no | 9.3853e-07 | 5.8544e-20 | 0 | -0.00053503 | yes | 0 | -0.0001815 | no | OK | 0.39453 | yes | no outlying indexes |  |  | NTC OK
150515_M00766_0098_000000000-AF9ND | M00766 | 2015-05-15 | 75 | 75 | in date | in date | in date | 96.8787 | no | 15942131 | low | 638.6353 | very low | 0.030068 | no | 96.9762 | no | 0.97995 | 0.85 | no | 0.98705 | no | 0.97446 | no | 0.84095 | no | 0.57572 | nan | 0 | -0.00024399 | no | 0 | -0.00016041 | no | OK | 0.33654 | yes | no outlying indexes |  |  | NTC OK
150518_M00766_0100_000000000-AATJ7 | M00766 | 2015-05-18 | 151 | 151 | in date | in date | in date | 89.6123 | no | 7467630 | low | 351.2212 | very low | 0.03759 | no | 88.7216 | no | 0.94106 | 0.8 | no | 0.95572 | no | 0.92587 | no | 0.84095 | no | 0.55116 | 0.39443 | 0 | -0.00037129 | no | 0 | -0.00084743 | yes | OK | 2.6175 | yes | ['Faststart' '360Taq4mM' '360Taq2mM'] | [] | ['Faststart' '360Taq4mM' '360Taq2mM'] | NTC OK
150521_M00766_0102_000000000-AF9MV | M00766 | 2015-05-21 | 75 | 75 | in date | in date | in date | 89.6458 | no | 30563261 | low | 1273.6592 | OK | 0.043729 | no | 90.9553 | no | 0.95232 | 0.85 | no | 0.96492 | no | 0.93927 | no | 0.84095 | no | 0.0097477 | nan | 0 | -0.00050706 | yes | 0 | -0.00046026 | no | OK | 0.3581 | yes | ['14M09383'] | [] | ['14M09383'] | NTC OK
150526_M02641_0053_000000000-AFHDV | M02641 | 2015-05-26 | 200 | 200 | in date | in date | in date | 48.4208 | yes | 7559009 | low | 310.8307 | very low | 0.22965 | yes | 50.5562 | yes | 0.64122 | 0.7 | yes | 0.51285 | yes | 0.75917 | no | 0.50766 | no | 1.0758e-55 | 2.1338e-299 | 146 | -0.0039621 | yes | 1 | -0.0018861 | yes | OK | 0.33972 | yes | ['14M12202'] | [] | ['14M12202'] | NTC OK
150528_M00766_0105_000000000-AF9MJ | M00766 | 2015-05-28 | 75 | 75 | in date | in date | in date | 91.4093 | no | 30434248 | low | 1259.146 | OK | 0.033027 | no | 91.4154 | no | 0.9591 | 0.85 | no | 0.96908 | no | 0.94896 | no | 0.84095 | no | 0.011641 | 5.539e-05 | 0 | -0.00043918 | no | 0 | -0.00032915 | no | OK | 0.43106 | yes | ['14M12059'] | [] | ['14M12059'] | NTC OK
150529_M00766_0106_000000000-AEVP5 | M00766 | 2015-05-29 | 151 | 151 | in date | in date | in date | 88.3589 | no | 15502868 | low | 806.7578 | very low | 0.032653 | no | 89.3692 | no | 0.91331 | 0.8 | no | 0.92662 | no | 0.89824 | no | 0.99584 | no | 0.34729 | 5.8531e-07 | 0 | -0.00082214 | yes | 0 | -0.0010282 | yes | OK | 0.32149 | yes | ['13M70706' '13M70790'] | ['13M70706' '13M70790'] | [] | NTC OK
150605_M02641_0058_000000000-AFF15 | M02641 | 2015-06-05 | 75 | 75 | in date | in date | in date | 86.9846 | no | 20446753 | low | 822.314 | very low | 0.026851 | no | 87.5712 | no | 0.93738 | 0.85 | no | 0.94849 | no | 0.92669 | no | 0.95413 | no | 0.71981 | nan | 0 | -0.00062227 | yes | 0 | -0.00041516 | no | OK | 0.38966 | yes | ['14M08356'] | ['14M08356'] | [] | NTC OK
150610_M02641_0059_000000000-AFN4H | M02641 | 2015-06-10 | 75 | 75 | in date | in date | in date | 82.3028 | no | 25297112 | low | 1022.9657 | very low | 0.024105 | no | 83.0805 | yes | 0.91102 | 0.85 | no | 0.93192 | no | 0.89027 | no | 0.84095 | no | 0.00015742 | 7.1708e-18 | 0 | -0.00077999 | yes | 0 | -0.00063882 | yes | OK | 0.52264 | yes | ['15M03072'] | [] | ['15M03072'] | NTC OK
150612_M00766_0112_000000000-AFMW5 | M00766 | 2015-06-12 | 75 | 75 | in date | in date | in date | 85.2766 | no | 34132691 | low | 1448.1255 | high | 0.019585 | no | 85.0541 | no | 0.93156 | 0.85 | no | 0.95081 | no | 0.91105 | no | 0.84095 | no | 3.193e-06 | 6.9549e-35 | 0 | -0.0006103 | yes | 0 | -0.00064344 | yes | OK | 0.42363 | yes | no outlying indexes |  |  | NTC OK
150612_M02641_0061_000000000-AFMVT | M02641 | 2015-06-12 | 75 | 75 | in date | in date | in date | 79.341 | yes | 25070490 | low | 1016.9582 | very low | 0.03704 | no | 80.1022 | yes | 0.90347 | 0.85 | no | 0.92645 | no | 0.88019 | no | 0.84095 | no | 3.0151e-05 | 3.0045e-11 | 0 | -0.00078584 | yes | 0 | -0.00069582 | yes | OK | 0.44048 | yes | ['15M04137'] | [] | ['15M04137'] | NTC OK
150615_M00766_0113_000000000-AFBH3 | M00766 | 2015-06-15 | 251 | 251 | in date | in date | in date | 89.384 | no | 20944521 | low | 1111.4379 | OK | 0.037506 | no | 89.0846 | no | 0.71363 | 0.75 | yes | 0.81632 | no | 0.60535 | yes | 0.50766 | no | 3.3983e-154 | 5.608e-166 | 24 | -0.0023213 | yes | 53 | -0.004186 | yes | OK | 0.11445 | yes | ['15M01148'] | [] | ['15M01148'] | NTC OK
150619_M00766_0114_000000000-AFMYN | M00766 | 2015-06-19 | 75 | 75 | in date | in date | in date | 92.1565 | no | 28125244 | low | 1158.4526 | low | 0.034411 | no | 93.2083 | no | 0.95863 | 0.85 | no | 0.97242 | no | 0.94406 | no | 0.84095 | no | 0.00010309 | nan | 0 | -0.00041056 | no | 0 | -0.00058758 | yes | OK | 0.33916 | yes | no outlying indexes |  |  | NTC OK
150622_M00766_0115_000000000-AFN2H | M00766 | 2015-06-22 | 75 | 75 | in date | in date | in date | 93.2875 | no | 26234040 | low | 1075.6113 | very low | 0.020578 | no | 93.4073 | no | 0.96472 | 0.85 | no | 0.97473 | no | 0.95557 | no | 0.84095 | no | 0.048507 | 0.0078432 | 0 | -0.00038391 | no | 0 | -0.0002776 | no | OK | 0.38333 | yes | ['15M05252'] | [] | ['15M05252'] | NTC OK
150623_M00766_0116_000000000-AFLEW | M00766 | 2015-06-23 | 151 | 151 | in date | in date | in date | 87.7026 | no | 20733081 | low | 1104.0961 | OK | 0.034923 | no | 86.8671 | no | 0.84466 | 0.8 | no | 0.89266 | no | 0.79387 | yes | 0.84095 | no | 1.473e-43 | 7.7036e-97 | 0 | -0.0019268 | yes | 9 | -0.0023623 | yes | OK | 0.42771 | yes | no outlying indexes |  |  | NTC OK
150706_M00766_0118_000000000-AF512 | M00766 | 2015-07-06 | 200 | 200 | in date | in date | in date | 94.4229 | no | 16447526 | low | 844.1499 | very low | 0.016368 | no | 94.5523 | no | 0.86041 | 0.75 | no | 0.89416 | no | 0.82397 | no | 0.84095 | no | 3.1972e-28 | nan | 1 | -0.001537 | yes | 1 | -0.0023974 | yes | OK | 0.17724 | yes | no outlying indexes |  |  | NTC OK
150729_M02641_0008_000000000-AGET9 | M02641 | 2015-07-29 | 200 | 200 | in date | in date | in date | 91.862 | no | 30439243 | low | 1258.6104 | OK | 0.012085 | no | 92.3982 | no | 0.84428 | 0.7 | no | 0.87341 | no | 0.81136 | no | 0.67794 | no | 1.7873e-34 | 2.8556e-270 | 1 | -0.0016885 | yes | 2 | -0.0024684 | yes | OK | 0.27015 | yes | ['15M05450'] | [] | ['15M05450'] | NTC OK
150730_M00766_0120_000000000-AFN4E | M00766 | 2015-07-30 | 75 | 75 | in date | in date | in date | 94.0611 | no | 23945923 | low | 972.2454 | very low | 0.052603 | yes | 94.2795 | no | 0.9647 | 0.85 | no | 0.97668 | no | 0.95322 | no | 0.84095 | no | 0.0076657 | nan | 0 | -0.00038193 | no | 0 | -0.0003459 | no | OK | 0.34994 | yes | no outlying indexes |  |  | NTC OK
150731_M00766_0121_000000000-AFMYC | M00766 | 2015-07-31 | 75 | 75 | in date | in date | in date | 89.5212 | no | 27660854 | low | 1152.6219 | low | 0.036004 | no | 89.6942 | no | 0.94787 | 0.85 | no | 0.95439 | no | 0.94056 | no | 0.99584 | no | 0.98595 | nan | 0 | -0.00062082 | yes | 0 | -0.00041404 | no | OK | 0.82507 | yes | ['15M07986'] | [] | ['15M07986'] | NTC OK
150731_M02641_0009_000000000-AFN3M | M02641 | 2015-07-31 | 75 | 75 | in date | in date | in date | 90.2753 | no | 32967275 | low | 1364.8961 | OK | 0.020875 | no | 90.6855 | no | 0.9485 | 0.85 | no | 0.96322 | no | 0.93353 | no | 0.84095 | no | 0.00029106 | 1.2672e-06 | 0 | -0.00052951 | yes | 0 | -0.00044329 | no | OK | 0.29214 | yes | no outlying indexes |  |  | NTC OK
150807_M02641_0012_000000000-AFMYT | M02641 | 2015-08-07 | 75 | 75 | in date | in date | in date | 96.8172 | no | 16466109 | low | 661.2988 | very low | 0.043831 | no | 97.2034 | no | 0.97599 | 0.85 | no | 0.98258 | no | 0.97147 | no | 0.95413 | no | 0.81203 | 0.23825 | 0 | -0.00028401 | no | 0 | -0.0001526 | no | OK | 0.35887 | yes | ['15M08559'] | [] | ['15M08559'] | NTC OK
150810_M00766_0123_000000000-AGTUY | M00766 | 2015-08-10 | 151 | 151 | in date | in date | in date | 89.2945 | no | 22283310 | low | 1188.9462 | OK | 0.016194 | no | 88.9362 | no | 0.92177 | 0.8 | no | 0.95497 | no | 0.88475 | no | 0.84095 | no | 2.7093e-32 | nan | 0 | -0.00035243 | no | 0 | -0.00083315 | yes | OK | 0.36536 | yes | ['15M07650'] | [] | ['15M07650'] | NTC OK
150811_M02641_0014_000000000-AGK0F | M02641 | 2015-08-11 | 151 | 151 | in date | in date | in date | 92.4253 | no | 22253985 | low | 1160.0949 | OK | 0.016889 | no | 92.3445 | no | 0.90106 | 0.8 | no | 0.9283 | no | 0.87132 | no | 0.95413 | no | 7.8772e-25 | nan | 0 | -0.0009934 | yes | 0 | -0.0015021 | yes | OK | 0.17315 | yes | no outlying indexes |  |  | NTC OK
150819_M02641_0016_000000000-AGJHU | M02641 | 2015-08-19 | 75 | 75 | in date | in date | in date | 94.1214 | no | 26775268 | low | 1083.954 | very low | 0.032934 | no | 94.4666 | no | 0.96462 | 0.85 | no | 0.97519 | no | 0.95489 | no | 0.84095 | no | 0.013871 | 2.5966e-09 | 0 | -0.00041278 | no | 0 | -0.00025665 | no | OK | 0.35264 | yes | no outlying indexes |  |  | NTC OK
150820_M00766_0126_000000000-AGHUU | M00766 | 2015-08-20 | 75 | 75 | in date | in date | in date | 92.7672 | no | 28026234 | low | 1148.9642 | low | 0.020853 | no | 93.2904 | no | 0.96152 | 0.85 | no | 0.97405 | no | 0.94902 | no | 0.84095 | no | 0.00031602 | nan | 0 | -0.00033273 | no | 0 | -0.0004084 | no | OK | 0.70143 | yes | no outlying indexes |  |  | NTC OK
150820_M02641_0017_000000000-AGJJA | M02641 | 2015-08-20 | 75 | 75 | in date | in date | in date | 93.4116 | no | 26592471 | low | 1089.8361 | very low | 0.015153 | no | 93.7457 | no | 0.96481 | 0.85 | no | 0.97351 | no | 0.95717 | no | 0.84095 | no | 0.072701 | 0.016306 | 0 | -0.00030311 | no | 0 | -0.00031655 | no | OK | 0.26272 | yes | no outlying indexes |  |  | NTC OK
150821_M02641_0018_000000000-AGJNU | M02641 | 2015-08-21 | 151 | 151 | in date | in date | in date | 96.3232 | no | 11702340 | low | 597.5554 | very low | 0.030041 | no | 96.5218 | no | 0.96101 | 0.8 | no | 0.97032 | no | 0.95086 | no | 0.95413 | no | 0.0070316 | nan | 0 | -0.00034988 | no | 0 | -0.000607 | yes | OK | 0.23539 | yes | no outlying indexes |  |  | NTC OK
150902_M00766_0128_000000000-AFLDG | M00766 | 2015-09-02 | 151 | 151 | in date | in date | in date | 89.3025 | no | 21077655 | low | 1129.9527 | OK | 0.015246 | no | 89.1889 | no | 0.83273 | 0.8 | no | 0.88185 | no | 0.78 | yes | 0.67794 | no | 7.5008e-47 | 3.9097e-271 | 2 | -0.0022787 | yes | 13 | -0.0028073 | yes | OK | 0.46847 | yes | no outlying indexes |  |  | NTC OK
150904_M02641_0022_000000000-AH991 | M02641 | 2015-09-04 | 200 | 200 | in date | in date | in date | 90.7156 | no | 32586889 | low | 1348.7337 | OK | 0.0084128 | no | 90.8839 | no | 0.80842 | 0.7 | no | 0.83792 | no | 0.774 | no | 0.50766 | no | 6.5299e-46 | 4.3193e-60 | 33 | -0.002385 | yes | 45 | -0.0029542 | yes | OK | 0.26214 | yes | ['15M05232' '15M02845'] | [] | ['15M05232' '15M02845'] | NTC OK
150910_M02641_0023_000000000-AFLHH | M02641 | 2015-09-10 | 151 | 151 | in date | in date | in date | 89.5554 | no | 20794433 | low | 1100.488 | OK | 0.02384 | no | 88.8672 | no | 0.85359 | 0.8 | no | 0.88668 | no | 0.81835 | no | 0.95413 | no | 2.3039e-18 | nan | 0 | -0.0021309 | yes | 9 | -0.0024176 | yes | OK | 0.92371 | yes | ['15M04253'] | [] | ['15M04253'] | NTC OK
150911_M00766_0131_000000000-AGLDT | M00766 | 2015-09-11 | 75 | 75 | in date | in date | in date | 91.8669 | no | 29894066 | low | 1233.1842 | OK | 0.017257 | no | 92.0012 | no | 0.95954 | 0.85 | no | 0.97065 | no | 0.94872 | no | 0.84095 | no | 0.0036196 | 0.00029275 | 0 | -0.0004291 | no | 0 | -0.00031438 | no | OK | 0.30964 | yes | no outlying indexes |  |  | NTC OK
150911_M02641_0024_AF50J | M02641 | 2015-09-11 | 151 | 151 | out of date | in date | in date | 91.3886 | no | 20193785 | low | 1043.6462 | OK | 0.026601 | no | 91.1381 | no | 0.85976 | 0.8 | no | 0.89102 | no | 0.82694 | no | 0.84095 | no | 4.1997e-18 | nan | 0 | -0.0020335 | yes | 5 | -0.0021409 | yes | OK | 0.64106 | yes | ['15M04270'] | [] | ['15M04270'] | NTC OK
150914_M00766_0132_000000000-AF41F | M00766 | 2015-09-14 | 151 | 151 | in date | in date | in date | 86.8892 | no | 3991745 | low | 184.4777 | very low | 0.012611 | no | 87.677 | no | 0.95124 | 0.8 | no | 0.96221 | no | 0.94056 | no | 0.95413 | no | 0.17881 | 0.0092674 | 0 | -0.00035999 | no | 0 | -0.00086178 | yes | OK | 0.34857 | yes | ['15M07808'] | ['15M07808'] | [] | NTC OK
150925_M02641_0028_000000000-AGK08 | M02641 | 2015-09-25 | 75 | 75 | in date | in date | in date | 95.0618 | no | 24906667 | low | 1010.387 | very low | 0.01775 | no | 95.4096 | no | 0.96734 | 0.85 | no | 0.97861 | no | 0.95784 | no | 0.84095 | no | 0.0070482 | 1.5449e-09 | 0 | -0.00035963 | no | 0 | -0.00022766 | no | OK | 0.33049 | yes | no outlying indexes |  |  | NTC OK
150928_M02641_0029_000000000-AGHV5 | M02641 | 2015-09-28 | 75 | 75 | in date | in date | in date | 93.773 | no | 28292454 | low | 1144.9595 | low | 0.021423 | no | 93.5295 | no | 0.96266 | 0.85 | no | 0.97493 | no | 0.95191 | no | 0.84095 | no | 0.012643 | nan | 0 | -0.00041356 | no | 0 | -0.00026611 | no | OK | 0.2835 | yes | no outlying indexes |  |  | NTC OK
150930_M02641_0031_000000000-AGJGT | M02641 | 2015-09-30 | 75 | 75 | in date | in date | in date | 93.3234 | no | 29303547 | low | 1206.8802 | OK | 0.049805 | no | 93.5156 | no | 0.9573 | 0.85 | no | 0.9715 | no | 0.94372 | no | 0.84095 | no | 7.6841e-05 | nan | 0 | -0.00045184 | no | 0 | -0.00030011 | no | OK | 0.51261 | yes | no outlying indexes |  |  | NTC OK
151001_M02641_0032_000000000-AGLE0 | M02641 | 2015-10-01 | 75 | 75 | in date | in date | in date | 87.3795 | no | 35725439 | low | 1504.6951 | very high | 0.017939 | no | 86.8081 | no | 0.9359 | 0.85 | no | 0.95659 | no | 0.91522 | no | 0.84095 | no | 5.6512e-08 | nan | 0 | -0.00058153 | yes | 0 | -0.00045814 | no | OK | 0.38331 | yes | no outlying indexes |  |  | NTC OK
151002_M02641_0033_000000000-AFLDA | M02641 | 2015-10-02 | 151 | 151 | in date | in date | in date | 88.0983 | no | 23364962 | low | 1229.9955 | high | 0.042476 | no | 87.9251 | no | 0.90553 | 0.8 | no | 0.937 | no | 0.87138 | no | 0.84095 | no | 1.4773e-23 | 1.9239e-34 | 0 | -0.00058099 | yes | 0 | -0.0010877 | yes | OK | 0.54428 | yes | ['8M4079'] | [] | ['8M4079'] | NTC OK
151009_M00766_0140_000000000-AGK0B | M00766 | 2015-10-09 | 75 | 75 | in date | in date | in date | 87.3055 | no | 24687917 | low | 1048.2614 | very low | 0.024799 | no | 87.0808 | no | 0.95594 | 0.85 | no | 0.96314 | no | 0.94902 | no | 0.99584 | no | 0.95838 | 0.16034 | 0 | -0.00045147 | no | 0 | -0.00036923 | no | OK | 0.4225 | yes | ['15M10481'] | ['15M10481'] | [] | NTC OK
151009_M02641_0036_000000000-AFN30 | M02641 | 2015-10-09 | 151 | 151 | in date | in date | in date | 91.4518 | no | 21458630 | low | 1123.3539 | OK | 0.024973 | no | 91.2766 | no | 0.86713 | 0.8 | no | 0.89737 | no | 0.83419 | no | 0.95413 | no | 2.2346e-20 | 0 | 0 | -0.0020801 | yes | 10 | -0.0026428 | yes | OK | 0.89911 | yes | ['15M09439' '15M09554'] | [] | ['15M09439' '15M09554'] | NTC OK
151012_M02641_0037_000000000-AFNA5 | M02641 | 2015-10-12 | 151 | 151 | in date | in date | in date | 93.0707 | no | 19894718 | low | 1031.818 | OK | 0.019111 | no | 93.4892 | no | 0.88194 | 0.8 | no | 0.90915 | no | 0.85233 | no | 0.95413 | no | 1.2493e-20 | 1.0755e-37 | 0 | -0.0019675 | yes | 8 | -0.0026129 | yes | OK | 0.45894 | yes | ['15M09548'] | [] | ['15M09548'] | NTC OK
151019_M02641_0040_000000000-AJDVH | M02641 | 2015-10-19 | 151 | 151 | in date | in date | in date | 90.6324 | no | 24425659 | OK | 1287.4209 | very high | 0.019011 | no | 90.7816 | no | 0.92696 | 0.8 | no | 0.94837 | no | 0.90505 | no | 0.95413 | no | 6.5136e-18 | nan | 0 | -0.00048911 | no | 0 | -0.00086419 | yes | OK | 0.45517 | yes | ['15M10923'] | ['15M10923'] | [] | NTC OK
151021_M00766_0144_000000000-AG028 | M00766 | 2015-10-21 | 75 | 75 | in date | in date | in date | 85.0537 | no | 24774347 | low | 1085.7652 | very low | 0.013905 | no | 84.6897 | yes | 0.94129 | 0.85 | no | 0.95646 | no | 0.92591 | no | 0.84095 | no | 0.0030374 | 4.0367e-11 | 0 | -0.00055587 | yes | 0 | -0.00055374 | yes | OK | 0.46617 | yes | ['15M10881'] | ['15M10881'] | [] | NTC OK
151022_M02641_0042_000000000-AJ5B5 | M02641 | 2015-10-22 | 200 | 200 | in date | in date | in date | 89.6462 | no | 28861426 | low | 1217.8503 | OK | 0.010571 | no | 89.8396 | no | 0.67585 | 0.7 | yes | 0.7507 | no | 0.59122 | yes | 0.67794 | no | 4.6203e-88 | 7.0766e-116 | 12 | -0.0033975 | yes | 75 | -0.0048318 | yes | OK | 0.26348 | yes | ['15M09575' '15M10221'] | [] | ['15M09575' '15M10221'] | NTC OK
151026_M00766_0148_000000000-AFLVV | M00766 | 2015-10-26 | 151 | 151 | in date | in date | in date | 88.3189 | no | 18895886 | low | 1013.3256 | OK | 0.032209 | no | 89.0093 | no | 0.93493 | 0.8 | no | 0.95661 | no | 0.91268 | no | 0.84095 | no | 9.4968e-15 | 5.2193e-112 | 0 | -0.00031813 | no | 0 | -0.00073043 | yes | OK | 0.28355 | yes | ['15M10912'] | [] | ['15M10912'] | NTC OK
151102_M02641_0045_000000000-AFN3H | M02641 | 2015-11-02 | 151 | 151 | in date | in date | in date | 94.4283 | no | 15063309 | low | 779.0627 | very low | 0.048767 | no | 94.8292 | no | 0.95449 | 0.8 | no | 0.96887 | no | 0.93956 | no | 0.84095 | no | 2.4447e-05 | 2.8037e-38 | 0 | -0.00027595 | no | 0 | -0.00058311 | yes | OK | 0.18495 | yes | ['15M10676' '15M10427'] | ['15M10427'] | ['15M10676'] | NTC OK
151104_M00766_0151_000000000-AK8EW | M00766 | 2015-11-04 | 151 | 151 | in date | in date | in date | 82.696 | no | 16143415 | low | 883.0439 | low | 0.069792 | yes | 83.2962 | yes | 0.91934 | 0.8 | no | 0.94177 | no | 0.89677 | no | 0.95413 | no | 1.0753e-07 | 5.9859e-42 | 0 | -0.00035735 | no | 0 | -0.00067331 | yes | OK | 0.40846 | yes | no outlying indexes |  |  | NTC OK
151106_M00766_0152_000000000-AJ5W5 | M00766 | 2015-11-06 | 151 | 151 | in date | in date | in date | 89.6104 | no | 16168604 | low | 863.0096 | low | 0.024136 | no | 90.7829 | no | 0.9317 | 0.8 | no | 0.95041 | no | 0.91232 | no | 0.95413 | no | 3.0502e-07 | 1.2764e-25 | 0 | -0.00049715 | no | 0 | -0.00054569 | yes | OK | 0.26734 | yes | ['15M11860' '15M11870'] | ['15M11860'] | ['15M11870'] | NTC OK
151111_M02641_0049_000000000-AJD47 | M02641 | 2015-11-11 | 200 | 200 | in date | in date | in date | 92.8535 | no | 30208881 | low | 1238.2172 | OK | 0.021332 | no | 93.316 | no | 0.83991 | 0.7 | no | 0.87018 | no | 0.80592 | no | 0.50766 | no | 8.1049e-40 | 6.164e-109 | 14 | -0.0019638 | yes | 19 | -0.0027647 | yes | OK | 0.28743 | yes | ['15M10790'] | [] | ['15M10790'] | NTC OK
151116_M00766_0154_000000000-AJF4C | M00766 | 2015-11-16 | 151 | 151 | in date | in date | in date | 87.9376 | no | 13835502 | low | 734.2783 | very low | 0.054577 | yes | 88.5491 | no | 0.9313 | 0.8 | no | 0.95109 | no | 0.90966 | no | 0.95413 | no | 5.1553e-08 | nan | 0 | -0.00033626 | no | 0 | -0.00061836 | yes | OK | 0.25181 | yes | ['15M12301'] | ['15M12301'] | [] | NTC OK
151119_M02641_0052_000000000-AJJ6B | M02641 | 2015-11-19 | 75 | 75 | in date | in date | out of date | 94.5733 | no | 24895626 | low | 1016.266 | very low | 0.029336 | no | 94.9966 | no | 0.96439 | 0.85 | no | 0.97782 | no | 0.95208 | no | 0.84095 | no | 0.00044318 | 1.1267e-06 | 0 | -0.00028501 | no | 0 | -0.00037122 | no | OK | 0.33896 | yes | no outlying indexes |  |  | NTC OK
151125_M00766_0158_000000000-AJJC1 | M00766 | 2015-11-25 | 151 | 151 | in date | in date | in date | 78.1687 | yes | 3563033 | low | 150.7739 | very low | 0.031274 | no | 78.1863 | yes | 0.9558 | 0.8 | no | 0.96712 | no | 0.94724 | no | 0.95413 | no | 0.17297 | 0.010728 | 0 | -0.00022353 | no | 0 | -0.00043897 | no | OK | 0.29368 | yes | ['15M12680' '15M12866'] | ['15M12680'] | ['15M12866'] | NTC OK
151130_M00766_0163_000000000-AJGC9 | M00766 | 2015-11-30 | 151 | 151 | in date | in date | in date | 79.0357 | yes | 21816702 | low | 1234.7264 | high | 0.043221 | no | 79.6269 | yes | 0.91879 | 0.8 | no | 0.94269 | no | 0.8946 | no | 0.95413 | no | 4.5777e-17 | 3.5682e-28 | 0 | -0.0002497 | no | 0 | -0.00046641 | no | OK | 0.37872 | yes | ['15M12680' '15M12866'] | ['15M12680'] | ['15M12866'] | NTC OK
151204_M02641_0056_000000000-AJDC6 | M02641 | 2015-12-04 | 200 | 200 | in date | in date | in date | 93.7017 | no | 25575432 | low | 1042.2805 | very low | 0.036678 | no | 94.2736 | no | 0.85421 | 0.7 | no | 0.88062 | no | 0.82404 | no | 0.50766 | no | 2.3201e-29 | nan | 13 | -0.0018496 | yes | 19 | -0.0025865 | yes | OK | 0.1513 | yes | ['15M04443'] | [] | ['15M04443'] | NTC OK
151208_M02641_0057_000000000-AJHRF | M02641 | 2015-12-08 | 151 | 151 | in date | in date | in date | 97.1211 | no | 4496775 | low | 228.8877 | very low | 0.02523 | no | 97.1074 | no | 0.94501 | 0.8 | no | 0.95817 | no | 0.93055 | no | 0.95413 | no | 0.083233 | 0.003379 | 0 | -0.0007036 | yes | 0 | -0.00058399 | yes | OK | 0.58595 | yes | ['15M12287'] | [] | ['15M12287'] | NTC OK
151210_M00766_0168_000000000-AJDDH | M00766 | 2015-12-10 | 75 | 75 | in date | in date | in date | 91.376 | no | 26145512 | low | 1086.9972 | very low | 0.021546 | no | 90.8931 | no | 0.956 | 0.85 | no | 0.96684 | no | 0.94501 | no | 0.84095 | no | 0.10224 | 0.014367 | 0 | -0.00051081 | yes | 0 | -0.0004594 | no | OK | 0.4035 | yes | ['15M13946'] | ['15M13946'] | [] | NTC OK
151211_M00766_0169_000000000-AK5DW | M00766 | 2015-12-11 | 300 | 300 | in date | in date | in date | 93.0016 | no | 21872195 | low | 891.8405 | very low | 0.023515 | no | 93.6566 | no | 0.54826 | 0.7 | yes | 0.59108 | yes | 0.48586 | yes | 0.24079 | no | 3.3919e-41 | 2.4542e-59 | 169 | -0.0035153 | yes | 181 | -0.004229 | yes | OK | 0.21949 | yes | ['14M10927b'] | [] | ['14M10927b'] | NTC OK
151218_M00766_0171_000000000-AK5FL | M00766 | 2015-12-18 | 226 | 226 | in date | in date | in date | 92.0182 | no | 20535429 | low | 846.6865 | very low | 0.020936 | no | 91.9926 | no | 0.57401 | 0.7 | yes | 0.58212 | yes | 0.54171 | yes | 0.67794 | no | 3.7819e-14 | 8.2184e-29 | 24 | -0.005182 | yes | 97 | -0.0055294 | yes | OK | 0.25591 | yes | no outlying indexes |  |  | NTC OK
160104_M02641_0062_000000000-AL603 | M02641 | 2016-01-04 | 101 | 101 | in date | in date | in date | 93.4734 | no | 19858211 | low | 1030.1337 | OK | 0.01816 | no | 93.4478 | no | 0.95548 | 0.8 | no | 0.97027 | no | 0.94346 | no | 0.84095 | no | 1.6436e-05 | 1.1366e-31 | 0 | -0.00025903 | no | 0 | -0.00019241 | no | OK | 0.31626 | yes | ['15M11705'] | [] | ['15M11705'] | NTC OK
160105_M00766_0174_000000000-AL4LB | M00766 | 2016-01-05 | 151 | 151 | in date | in date | in date | 91.2681 | no | 14218489 | low | 749.2227 | very low | 0.036601 | no | 91.4658 | no | 0.93032 | 0.8 | no | 0.95273 | no | 0.9068 | no | 0.95413 | no | 1.8718e-11 | nan | 0 | -0.00053469 | yes | 0 | -0.00071583 | yes | OK | 0.63317 | yes | ['15M12287'] | [] | ['15M12287'] | NTC OK
160107_M02641_0063_000000000-AJDC9 | M02641 | 2016-01-07 | 75 | 75 | in date | in date | in date | 94.1624 | no | 24963706 | low | 1022.5057 | very low | 0.017799 | no | 94.5508 | no | 0.9694 | 0.85 | no | 0.97637 | no | 0.96261 | no | 0.84095 | no | 0.22798 | nan | 0 | -0.00029644 | no | 0 | -0.00032503 | no | OK | 0.51228 | yes | no outlying indexes |  |  | NTC OK
160108_M02641_0064_000000000-ALNH2 | M02641 | 2016-01-08 | 200 | 200 | in date | in date | in date | 91.9702 | no | 21242610 | low | 1109.3897 | OK | 0.022049 | no | 92.3443 | no | 0.8579 | 0.75 | no | 0.90283 | no | 0.81012 | no | 0.67794 | no | 5.5542e-59 | nan | 0 | -0.0013823 | yes | 1 | -0.0025063 | yes | OK | 0.31714 | yes | ['15M11705'] | [] | ['15M11705'] | NTC OK
160115_M00766_0178_000000000-ALKJV | M00766 | 2016-01-15 | 151 | 151 | in date | in date | in date | 83.9557 | no | 15235349 | low | 831.5431 | very low | 0.066527 | yes | 83.4762 | yes | 0.91226 | 0.8 | no | 0.93609 | no | 0.8844 | no | 0.95413 | no | 5.102e-09 | nan | 0 | -0.00057903 | yes | 0 | -0.0011478 | yes | OK | 0.31466 | yes | ['15M13506'] | ['15M13506'] | [] | NTC OK
160118_M02641_0066_000000000-AJETF | M02641 | 2016-01-18 | 75 | 75 | in date | in date | in date | 98.1898 | no | 9321517 | low | 373.3045 | very low | 0.028008 | no | 98.3733 | no | 0.97492 | 0.85 | no | 0.97945 | no | 0.97382 | no | 0.99998 | no | 1 | 0.5283 | 0 | -0.00022747 | no | 0 | -0.000157 | no | OK | 0.35791 | yes | ['15M13946'] | ['15M13946'] | [] | NTC OK
160119_M02641_0067_000000000-ALN4P | M02641 | 2016-01-19 | 151 | 151 | in date | in date | in date | 87.2125 | no | 18949772 | low | 995.7574 | low | 0.091472 | yes | 86.1573 | no | 0.9035 | 0.8 | no | 0.92594 | no | 0.87789 | no | 0.95413 | no | 2.2467e-08 | nan | 0 | -0.00078266 | yes | 0 | -0.0012195 | yes | OK | 0.31982 | yes | ['15M14816'] | [] | ['15M14816'] | NTC OK
160127_M00766_0001_000000000-ALR2L | M00766 | 2016-01-27 | 151 | 151 | in date | in date | in date | 94.4219 | no | 16728171 | low | 869.4499 | low | 0.024493 | no | 94.8131 | no | 0.94892 | 0.8 | no | 0.95934 | no | 0.93843 | no | 0.95413 | no | 5.3787e-06 | nan | 0 | -0.00045839 | no | 0 | -0.00066696 | yes | OK | 0.97347 | yes | no outlying indexes |  |  | NTC OK
160127_M02641_0069_000000000-AMDH4 | M02641 | 2016-01-27 | 75 | 75 | in date | in date | in date | 76.69 | yes | 29747487 | low | 1328.6387 | OK | 0.096919 | yes | 78.2455 | yes | 0.91416 | 0.85 | no | 0.93371 | no | 0.89378 | no | 0.84095 | no | 0.00062682 | nan | 0 | -0.00074637 | yes | 0 | -0.00037367 | no | OK | 0.52107 | yes | no outlying indexes |  |  | NTC OK
160128_M00766_0002_000000000-ALAJ1 | M00766 | 2016-01-28 | 75 | 75 | in date | in date | in date | 88.2213 | no | 28975202 | low | 1238.6174 | OK | 0.0080531 | no | 88.3898 | no | 0.95207 | 0.85 | no | 0.96129 | no | 0.94181 | no | 0.84095 | no | 0.078659 | 0.012869 | 0 | -0.00042109 | no | 0 | -0.00041041 | no | OK | 0.33634 | yes | no outlying indexes |  |  | NTC OK
160204_M00766_0004_000000000-AL4H0 | M00766 | 2016-02-04 | 151 | 151 | in date | in date | in date | 95.4825 | no | 14166224 | low | 726.9693 | very low | 0.021398 | no | 95.3691 | no | 0.95549 | 0.8 | no | 0.97197 | no | 0.93852 | no | 0.84095 | no | 3.1187e-14 | 2.8941e-31 | 0 | -0.00029231 | no | 0 | -0.00059973 | yes | OK | 0.56692 | yes | ['15M13181'] | [] | ['15M13181'] | NTC OK
160204_M02641_0072_000000000-AK6CE | M02641 | 2016-02-04 | 75 | 75 | in date | in date | in date | 91.6672 | no | 25918178 | low | 1060.0101 | very low | 0.015091 | no | 91.5669 | no | 0.95246 | 0.85 | no | 0.96475 | no | 0.93912 | no | 0.84095 | no | 0.00023857 | 2.281e-41 | 0 | -0.00053931 | yes | 0 | -0.00055384 | yes | OK | 0.32617 | yes | no outlying indexes |  |  | NTC OK
160205_M02641_0073_000000000-ALY9Y | M02641 | 2016-02-05 | 75 | 75 | in date | in date | in date | 81.6149 | no | 30119640 | low | 1331.3992 | OK | 0.025391 | no | 80.9054 | yes | 0.92844 | 0.85 | no | 0.94565 | no | 0.90982 | no | 0.84095 | no | 0.023743 | nan | 0 | -0.0004287 | no | 0 | -0.00027127 | no | OK | 0.37123 | yes | no outlying indexes |  |  | NTC OK
160208_M00766_0006_000000000-AMF4G | M00766 | 2016-02-08 | 75 | 75 | in date | in date | in date | 85.7346 | no | 32902642 | low | 1404.3574 | high | 0.015626 | no | 85.8175 | no | 0.93494 | 0.85 | no | 0.95198 | no | 0.91641 | no | 0.84095 | no | 1.9762e-05 | 5.3645e-08 | 0 | -0.00062874 | yes | 0 | -0.00051979 | yes | OK | 0.40774 | yes | ['16M00692'] | ['16M00692'] | [] | NTC OK
160208_M02641_0075_000000000-AMF40 | M02641 | 2016-02-08 | 75 | 75 | in date | in date | in date | 79.3999 | yes | 30335883 | low | 1362.3448 | OK | 0.022595 | no | 80.0194 | yes | 0.9085 | 0.85 | no | 0.93381 | no | 0.88187 | no | 0.84095 | no | 9.8314e-08 | 2.2541e-295 | 0 | -0.00078696 | yes | 0 | -0.00037771 | no | OK | 0.359 | yes | ['16M00691'] | [] | ['16M00691'] | NTC OK
160209_M02641_0076_000000000-AMF44 | M02641 | 2016-02-09 | 75 | 75 | in date | in date | in date | 89.416 | no | 23928513 | low | 1023.4734 | very low | 0.02047 | no | 89.4341 | no | 0.95438 | 0.85 | no | 0.96439 | no | 0.94402 | no | 0.84095 | no | 0.15811 | 4.2223e-80 | 0 | -0.00036997 | no | 0 | -0.00030272 | no | OK | 0.28461 | yes | ['15M15069'] | [] | ['15M15069'] | NTC OK
160210_M02641_0077_000000000-AME81 | M02641 | 2016-02-10 | 75 | 75 | in date | in date | in date | 93.9536 | no | 19770059 | low | 818.1394 | very low | 0.024181 | no | 94.3148 | no | 0.97229 | 0.85 | no | 0.97917 | no | 0.96582 | no | 0.95413 | no | 0.58184 | 6.9811e-05 | 0 | -0.00021377 | no | 0 | -0.00022485 | no | OK | 0.54079 | yes | no outlying indexes |  |  | NTC OK
160212_M00766_0008_000000000-AL607 | M00766 | 2016-02-12 | 151 | 151 | in date | in date | in date | 87.5244 | no | 20878061 | low | 1123.9732 | OK | 0.026805 | no | 87.1194 | no | 0.85079 | 0.8 | no | 0.8926 | no | 0.80667 | no | 0.84095 | no | 5.1098e-39 | nan | 0 | -0.002166 | yes | 10 | -0.0030426 | yes | OK | 0.42813 | yes | ['15M12553' '15M12790'] | ['15M12553' '15M12790'] | [] | NTC OK
160216_M02641_0079_000000000-AL75M | M02641 | 2016-02-16 | 151 | 151 | in date | in date | in date | 92.5527 | no | 17259426 | low | 905.0069 | low | 0.019188 | no | 93.4136 | no | 0.9397 | 0.8 | no | 0.96218 | no | 0.91691 | no | 0.67794 | no | 4.7356e-16 | nan | 0 | -0.00030351 | no | 0 | -0.00073347 | yes | OK | 0.44782 | yes | ['15M13455' '15M13442'] | ['15M13455' '15M13442'] | [] | NTC OK
160218_M00766_0010_000000000-AMF48 | M00766 | 2016-02-18 | 75 | 75 | in date | in date | in date | 90.3938 | no | 26783167 | low | 1115.6764 | low | 0.019331 | no | 90.2226 | no | 0.95211 | 0.85 | no | 0.96432 | no | 0.93856 | no | 0.84095 | no | 0.0028637 | 0.00027102 | 0 | -0.00054184 | yes | 0 | -0.00048552 | no | OK | 0.32057 | yes | ['15M15396'] | [] | ['15M15396'] | NTC OK
160218_M02641_0080_000000000-AMF3Y | M02641 | 2016-02-18 | 75 | 75 | in date | in date | in date | 92.7164 | no | 22919923 | low | 955.1644 | very low | 0.010077 | no | 92.7435 | no | 0.95689 | 0.85 | no | 0.97361 | no | 0.93955 | no | 0.84095 | no | 0.00019591 | nan | 0 | -0.00031985 | no | 0 | -0.00022652 | no | OK | 0.48739 | yes | no outlying indexes |  |  | NTC OK
| 63 |  |  | NUM FAILS | 3 | 0 | 4 |  | 11 |  | 120 |  | 89 |  | 10 |  | 19 |  |  | 9 |  | 5 |  | 16 |  | 0 |  |  | 16 |  | 66 | 33 |  | 73 | 5 |  | 121 | 66 | 26 | 44 | 0

## OutOfDate
flow cell out of date | pr2 out of date | reagent kit out of date | Fails (manual) | Fail Method 1 (auto)
no | no | no | 0 | 130417_M00766_0007_000000000-A3R1F
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 130613_M00766_0018_000000000-A4FN3 | 130613_M00766_0018_000000000-A4FN3
130708_M00766_0023_000000000-A20R8 | no | 130708_M00766_0023_000000000-A20R8 | 0 | 0
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 140205_M00766_0021_000000000-A7RVR | 140205_M00766_0021_000000000-A7RVR
no | no | no | 140603_M00766_0037_000000000-A7WU9 | 140603_M00766_0037_000000000-A7WU9
no | no | no | 0 | 140620_M02641_0004_000000000-A8R7V
no | no | no | 140625_M00766_0041_000000000-A8P78 | 140625_M00766_0041_000000000-A8P78
no | no | no | 0 | 140715_M02641_0008_000000000-A8P81
no | no | no | 140804_M02641_0011_000000000-AAD6A | 0
no | no | no | 0 | 140905_M02641_0017_000000000-AA3FN
no | no | no | 0 | 140917_M02641_0018_000000000-AA3H7
no | no | no | 0 | 141024_M02641_0024_000000000-A8P84
no | no | no | 0 | 141113_M02641_0025_000000000-A8RTY
141125_M00766_0062_000000000-A7BKW | no | no | 0 | 141125_M00766_0062_000000000-A7BKW
no | no | no | 141222_M00766_0067_000000000-ACCB3 | 141222_M00766_0067_000000000-ACCB3
no | no | no | 0 | 150112_M00766_0068_000000000-A8PD8
no | no | no | 150114_M00766_0069_000000000-AA68B | 150114_M00766_0069_000000000-AA68B
no | no | no | 0 | 150116_M00766_0070_000000000-ACC3G
no | no | 150127_M00766_0072_000000000-ABRK7 | 0 | 0
no | no | no | 150130_M00766_0073_000000000-ACBVB | 150130_M00766_0073_000000000-ACBVB
no | no | no | 150205_M02641_0029_000000000-ACCE2 | 0
no | no | no | 150220_M00766_0077_000000000-AA2U8 | 150220_M00766_0077_000000000-AA2U8
no | no | no | 0 | 0
no | no | no | 0 | 150227_M02641_0032_000000000-ACNF3
no | no | no | 0 | 0
no | no | no | 0 | 150306_M00766_0080_000000000-ACNEP
no | no | no | 0 | 150306_M02641_0034_000000000-ACRUF
no | no | no | 0 | 0
no | no | no | 0 | 150325_M00766_0084_000000000-AD8KN
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 150407_M00766_0088_000000000-ACCAT | 0
no | no | 150410_M00766_0089_000000000-ACBU4 | 0 | 0
no | no | no | 0 | 0
no | no | no | 150429_M02641_0043_000000000-AD8KJ | 150429_M02641_0043_000000000-AD8KJ
no | no | no | 0 | 0
no | no | no | 150501_M00766_0093_000000000-AF9MH | 0
no | no | no | 0 | 150501_M02641_0045_000000000-AD6UA
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 0 | 150515_M00766_0098_000000000-AF9ND
no | no | no | 0 | 150518_M00766_0100_000000000-AATJ7
no | no | no | 0 | 0
no | no | no | 150526_M02641_0053_000000000-AFHDV | 150526_M02641_0053_000000000-AFHDV
no | no | no | 0 | 0
no | no | no | 0 | 150529_M00766_0106_000000000-AEVP5
no | no | no | 0 | 150605_M02641_0058_000000000-AFF15
no | no | no | 0 | 150610_M02641_0059_000000000-AFN4H
no | no | no | 150612_M00766_0112_000000000-AFMW5 | 0
no | no | no | 0 | 150612_M02641_0061_000000000-AFMVT
no | no | no | 150615_M00766_0113_000000000-AFBH3 | 150615_M00766_0113_000000000-AFBH3
no | no | no | 0 | 0
no | no | no | 0 | 150622_M00766_0115_000000000-AFN2H
no | no | no | 0 | 0
no | no | no | 0 | 150706_M00766_0118_000000000-AF512
no | no | no | 0 | 0
no | no | no | 0 | 150730_M00766_0120_000000000-AFN4E
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 0 | 150807_M02641_0012_000000000-AFMYT
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 0 | 150819_M02641_0016_000000000-AGJHU
no | no | no | 0 | 0
no | no | no | 0 | 150820_M02641_0017_000000000-AGJJA
no | no | no | 0 | 150821_M02641_0018_000000000-AGJNU
no | no | no | 150902_M00766_0128_000000000-AFLDG | 0
no | no | no | 150904_M02641_0022_000000000-AH991 | 0
no | no | no | 150910_M02641_0023_000000000-AFLHH | 0
no | no | no | 0 | 0
150911_M02641_0024_AF50J | no | no | 0 | 0
no | no | no | 0 | 150914_M00766_0132_000000000-AF41F
no | no | no | 0 | 150925_M02641_0028_000000000-AGK08
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 151001_M02641_0032_000000000-AGLE0 | 151001_M02641_0032_000000000-AGLE0
no | no | no | 0 | 0
no | no | no | 0 | 151009_M00766_0140_000000000-AGK0B
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 0 | 151019_M02641_0040_000000000-AJDVH
no | no | no | 0 | 151021_M00766_0144_000000000-AG028
no | no | no | 151022_M02641_0042_000000000-AJ5B5 | 151022_M02641_0042_000000000-AJ5B5
no | no | no | 0 | 0
no | no | no | 0 | 151102_M02641_0045_000000000-AFN3H
no | no | no | 0 | 151104_M00766_0151_000000000-AK8EW
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 0 | 151116_M00766_0154_000000000-AJF4C
no | no | 151119_M02641_0052_000000000-AJJ6B | 0 | 151119_M02641_0052_000000000-AJJ6B
no | no | no | 0 | 151125_M00766_0158_000000000-AJJC1
no | no | no | 151130_M00766_0163_000000000-AJGC9 | 151130_M00766_0163_000000000-AJGC9
no | no | no | 0 | 151204_M02641_0056_000000000-AJDC6
no | no | no | 0 | 151208_M02641_0057_000000000-AJHRF
no | no | no | 0 | 151210_M00766_0168_000000000-AJDDH
no | no | no | 151211_M00766_0169_000000000-AK5DW | 151211_M00766_0169_000000000-AK5DW
no | no | no | 151218_M00766_0171_000000000-AK5FL | 151218_M00766_0171_000000000-AK5FL
no | no | no | 0 | 0
no | no | no | 0 | 160105_M00766_0174_000000000-AL4LB
no | no | no | 0 | 160107_M02641_0063_000000000-AJDC9
no | no | no | 0 | 0
no | no | no | 0 | 160115_M00766_0178_000000000-ALKJV
no | no | no | 0 | 160118_M02641_0066_000000000-AJETF
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 160127_M02641_0069_000000000-AMDH4 | 160127_M02641_0069_000000000-AMDH4
no | no | no | 0 | 0
no | no | no | 0 | 160204_M00766_0004_000000000-AL4H0
no | no | no | 0 | 160204_M02641_0072_000000000-AK6CE
no | no | no | 160205_M02641_0073_000000000-ALY9Y | 160205_M02641_0073_000000000-ALY9Y
no | no | no | 160208_M00766_0006_000000000-AMF4G | 0
no | no | no | 160208_M02641_0075_000000000-AMF40 | 160208_M02641_0075_000000000-AMF40
no | no | no | 0 | 160209_M02641_0076_000000000-AMF44
no | no | no | 0 | 160210_M02641_0077_000000000-AME81
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 0 | 0
no | no | no | 0 | 160218_M02641_0080_000000000-AMF3Y

## AnalyData
RunID | InstrumentID | RunDate | NumberOfCyclesR1 | NumberofCyclesR2 | %ReadsPassingFilter | %ReadsPassingFilterOutsideThreshold? | NumberofReadsPassingFilter | ReadCountWithinThreshold? | MeanClusterDensityPassingFilter | MeanClusterDensityWithinThreshold | COVClusterDensity | COVClusterDensityExceedsThreshold? | %MedianClusterDensityPassingFilter | GapBetweenClusterDensitiesPassingFilterAndNotExceedsThreshold? | ProportionOverQ30 | IlluminaQ30Threshold | ProportionOverQ30ExceedsThreshold? | ProportionOverQ30Read1 | ProportionOverQ30ExceedsThresholdRead1 | ProportionOverQ30Read2 | ProportionOverQ30ExceedsThresholdRead2 | K-SPValue | K-SPValueSignificant? | K-SUnpackPvalue | M-WUnpackPvalue | NumberOfCyclesExceedingThresholdUnderQ30Read1 | SlopeOfLineOfBestFitRead1 | SlopeOfLineOfBestFitRead1ExceedsThreshold? | NumberOfCyclesExceedingThresholdUnderQ30Read2 | SlopeOfLineOfBestFitRead2 | SlopeOfLineOfBestFitRead2ExceedsThreshold? | Fail? | Fail2? | Fail? No CD | Fail2? No CD | R1vsR2PropQ30 | R1vsR2K-S | NumCycleR1Threshold | NumCycleR2Threshold | R1vsR2CyclesQ30 | R1vsR2bestfit | KEY
130417_M00766_0007_000000000-A3R1F | M00766 | 2013-04-17 | 151 | 151 | 93.6874 | no | 14918743 | low | 773.0359 | very low | 0.024644 | no | 93.6673 | no | 0.97122 | 0.8 | no | 0.98017 | no | 0.96226 | no | 0.84095 | no | 0.0022684 | 4.2948e-06 | 0 | -0.00013884 | no | 0 | -0.00025711 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | incorrectly passed by auto QC
130508_M00766_0013_000000000-A3RB5 | M00766 | 2013-05-08 | 151 | 151 | 89.695 | no | 17060727 | low | 897.222 | low | 0.013058 | no | 89.6111 | no | 0.92341 | 0.8 | no | 0.94982 | no | 0.89785 | no | 0.67794 | no | 1.7122e-13 | 9.1751e-65 | 0 | -0.00060464 | yes | 0 | -0.00096571 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | incorrectly failed by auto QC
130510_M00766_0014_000000000-A3PFJ | M00766 | 2013-05-10 | 156 | 151 | 88.5142 | no | 18378480 | low | 995.2515 | low | 0.013484 | no | 88.6611 | no | 0.95249 | 0.75 | no | 0.96867 | no | 0.93502 | no | 0.95413 | no | 3.3216e-08 | nan | 0 | -0.0002029 | no | 0 | -0.00047138 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
130613_M00766_0018_000000000-A4FN3 | M00766 | 2013-06-13 | 251 | 251 | 62.2349 | yes | 14759376 | low | 877.4421 | low | 0.0052548 | no | 61.6486 | yes | 0.68299 | 0.75 | yes | 0.7656 | no | 0.58473 | yes | 0.67794 | no | 5.4172e-57 | 1.5425e-86 | 28 | -0.0020629 | yes | 55 | -0.0033012 | yes | fail | fail | fail | fail | diff | pass | exceeded | exceeded | same | same
130708_M00766_0023_000000000-A20R8 | M00766 | 2013-07-08 | 151 | 151 | 90.9082 | no | 8144995 | None | 872.0913 | None | 0.015455 | no | 90.7757 | no | 0.86589 | 0.85 | no | 0.91799 | no | 0.81047 | yes | 0.67794 | no | 9.6929e-19 | 6.9078e-114 | 0 | -0.0013571 | yes | 0 | -0.0030958 | yes | pass | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same
130902_M00766_0045_000000000-A5BVD | M00766 | 2013-09-02 | 156 | 151 | 92.4439 | no | 17107054 | low | 889.9218 | low | 0.020943 | no | 92.3467 | no | 0.95012 | 0.75 | no | 0.96778 | no | 0.93094 | no | 0.84095 | no | 1.4736e-06 | nan | 0 | -0.00028935 | no | 1 | -0.00072347 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
130909_M00766_0047_000000000-A5PEH | M00766 | 2013-09-09 | 151 | 151 | 91.8202 | no | 17818894 | low | 943.064 | low | 0.013531 | no | 91.9043 | no | 0.91659 | 0.8 | no | 0.93844 | no | 0.89323 | no | 0.95413 | no | 2.3926e-11 | 3.6095e-18 | 0 | -0.00083727 | yes | 0 | -0.00095948 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
140205_M00766_0021_000000000-A7RVR | M00766 | 2014-02-05 | 251 | 251 | 96.2303 | no | 10640092 | low | 549.8594 | very low | 0.015526 | no | 96.2427 | no | 0.77554 | 0.75 | no | 0.82459 | no | 0.7232 | yes | 0.84095 | no | 3.5805e-28 | 1.8115e-29 | 1 | -0.0023111 | yes | 65 | -0.0033538 | yes | fail | fail | pass | fail | diff | pass | not exceeded | exceeded | diff | same
140603_M00766_0037_000000000-A7WU9 | M00766 | 2014-06-03 | 251 | 251 | 80.6392 | no | 20279789 | low | 1121.2903 | OK | 0.020185 | no | 81.2363 | yes | 0.67262 | 0.75 | yes | 0.87096 | no | 0.46478 | yes | 0.24079 | no | 0 | 0 | 0 | -0.0011435 | yes | 73 | -0.0033779 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same
140620_M02641_0004_000000000-A8R7V | M02641 | 2014-06-20 | 151 | 151 | 95.2324 | no | 15895393 | low | 815.6811 | very low | 0.010793 | no | 95.1913 | no | 0.92041 | 0.8 | no | 0.96394 | no | 0.87468 | no | 0.35842 | no | 5.5248e-61 | nan | 0 | -0.0004185 | no | 0 | -0.00070599 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
140625_M00766_0041_000000000-A8P78 | M00766 | 2014-06-25 | 251 | 251 | 93.6991 | no | 12306226 | low | 638.3155 | very low | 0.015947 | no | 93.5516 | no | 0.80794 | 0.75 | no | 0.83721 | no | 0.77401 | no | 0.95413 | no | 1.446e-13 | nan | 15 | -0.0023266 | yes | 34 | -0.0029216 | yes | fail | fail | pass | pass | same | pass | not exceeded | exceeded | diff | same
140715_M02641_0008_000000000-A8P81 | M02641 | 2014-07-15 | 151 | 151 | 98.3026 | no | 5446264 | low | 277.542 | very low | 0.020978 | no | 98.4264 | no | 0.94495 | 0.8 | no | 0.95412 | no | 0.93542 | no | 0.99584 | no | 0.593 | 0.077423 | 0 | -0.00077434 | yes | 0 | -0.00075567 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
140804_M02641_0011_000000000-AAD6A | M02641 | 2014-08-04 | 251 | 251 | 92.8278 | no | 18236023 | low | 947.4631 | low | 0.014242 | no | 92.5473 | no | 0.75551 | 0.75 | no | 0.79909 | no | 0.70633 | yes | 0.84095 | no | 4.579e-31 | nan | 19 | -0.0025558 | yes | 41 | -0.0033381 | yes | pass | fail | pass | fail | diff | pass | not exceeded | exceeded | diff | same
140905_M02641_0017_000000000-AA3FN | M02641 | 2014-09-05 | 75 | 75 | 90.0019 | no | 23949422 | low | 1015.424 | very low | 0.055448 | yes | 89.6438 | no | 0.95625 | 0.85 | no | 0.96176 | no | 0.95019 | no | 0.99584 | no | 0.99831 | nan | 0 | -0.00042771 | no | 0 | -0.00021245 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
140917_M02641_0018_000000000-AA3H7 | M02641 | 2014-09-17 | 75 | 75 | 91.5859 | no | 24528146 | low | 1028.5468 | very low | 0.035153 | no | 91.5237 | no | 0.96229 | 0.85 | no | 0.96999 | no | 0.95452 | no | 0.84095 | no | 0.88947 | nan | 0 | -0.00031376 | no | 0 | -0.00016688 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
141024_M02641_0024_000000000-A8P84 | M02641 | 2014-10-24 | 151 | 151 | 94.9748 | no | 15281418 | low | 794.7029 | very low | 0.031108 | no | 95.0899 | no | 0.94073 | 0.8 | no | 0.95744 | no | 0.92359 | no | 0.84095 | no | 1.9766e-08 | 4.7689e-23 | 0 | -0.00055662 | yes | 0 | -0.00062808 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
141113_M02641_0025_000000000-A8RTY | M02641 | 2014-11-13 | 151 | 151 | 97.3253 | no | 9834883 | low | 502.4441 | very low | 0.049911 | no | 97.4887 | no | 0.91483 | 0.8 | no | 0.92802 | no | 0.90038 | no | 0.95413 | no | 0.00015291 | 0 | 0 | -0.00132 | yes | 0 | -0.0014718 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
141125_M00766_0062_000000000-A7BKW | M00766 | 2014-11-25 | 151 | 151 | 94.983 | no | 14809543 | low | 762.9367 | very low | 0.018589 | no | 95.0964 | no | 0.89379 | 0.8 | no | 0.96312 | no | 0.82977 | no | 0.24079 | no | 2.4117e-86 | nan | 0 | -0.00041911 | no | 4 | -0.00042946 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
141222_M00766_0067_000000000-ACCB3 | M00766 | 2014-12-22 | 151 | 151 | 67.7653 | yes | 14253970 | low | 760.392 | very low | 0.029815 | no | 68.1874 | yes | 0.82139 | 0.8 | no | 0.85456 | no | 0.78426 | yes | 0.95413 | no | 3.8283e-06 | nan | 0 | -0.0021203 | yes | 11 | -0.0013537 | yes | fail | fail | fail | fail | diff | pass | not exceeded | not exceeded | same | same
150112_M00766_0068_000000000-A8PD8 | M00766 | 2015-01-12 | 151 | 151 | 94.817 | no | 15773684 | low | 797.5388 | very low | 0.028582 | no | 94.9895 | no | 0.92451 | 0.8 | no | 0.93876 | no | 0.90863 | no | 0.99584 | no | 0.00017877 | 1.2387e-09 | 0 | -0.00090845 | yes | 0 | -0.00084942 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150114_M00766_0069_000000000-AA68B | M00766 | 2015-01-14 | 151 | 151 | 77.8262 | yes | 20704375 | low | 1148.0758 | OK | 0.024397 | no | 77.3315 | yes | 0.86381 | 0.8 | no | 0.91949 | no | 0.81333 | no | 0.95413 | no | 2.6748e-38 | nan | 0 | -0.00060681 | yes | 4 | -0.00057587 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150116_M00766_0070_000000000-ACC3G | M00766 | 2015-01-16 | 151 | 151 | 86.297 | no | 14147272 | low | 752.8015 | very low | 0.040986 | no | 86.2322 | no | 0.93362 | 0.8 | no | 0.95795 | no | 0.91769 | no | 0.95413 | no | 0.027224 | 4.5409e-07 | 0 | -0.00014326 | no | 3 | 0.00022032 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150127_M00766_0072_000000000-ABRK7 | M00766 | 2015-01-27 | 151 | 151 | 94.2571 | no | 17854832 | low | 926.0782 | low | 0.029511 | no | 94.3435 | no | 0.93257 | 0.8 | no | 0.95668 | no | 0.90691 | no | 0.67794 | no | 1.1008e-18 | 1.0313e-37 | 0 | -0.00054076 | yes | 0 | -0.00067649 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150130_M00766_0073_000000000-ACBVB | M00766 | 2015-01-30 | 251 | 251 | 81.2781 | no | 23929178 | low | 1312.9896 | very high | 0.017972 | no | 81.1332 | yes | 0.66236 | 0.75 | yes | 0.70216 | yes | 0.61962 | yes | 0.84095 | no | 6.7473e-31 | nan | 47 | -0.0033611 | yes | 64 | -0.0034493 | yes | fail | fail | fail | fail | same | pass | exceeded | exceeded | same | same
150205_M02641_0029_000000000-ACCE2 | M02641 | 2015-02-05 | 151 | 151 | 87.5711 | no | 18965301 | low | 1006.645 | OK | 0.019074 | no | 87.6822 | no | 0.8298 | 0.8 | no | 0.8663 | no | 0.79202 | yes | 0.99584 | no | 1.5184e-12 | 8.2211e-115 | 0 | -0.0022136 | yes | 2 | -0.0018972 | yes | pass | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same
150220_M00766_0077_000000000-AA2U8 | M00766 | 2015-02-20 | 26 | 26 | 11.2989 | yes | 3422864 | low | 175.3978 | very low | 0.069387 | yes | 0 | yes | 0.63211 | 0.9 | yes | 0.88444 | yes | 0.47035 | yes | 0.95413 | no | 6.5681e-12 | 7.5094e-13 | 0 | 0 | no | 26 | 0 | no | fail | fail | fail | fail | same | pass | not exceeded | exceeded | diff | same
150227_M00766_0079_000000000-ACPR9 | M00766 | 2015-02-27 | 75 | 75 | 91.9723 | no | 29954337 | low | 1225.6634 | OK | 0.019386 | no | 91.9816 | no | 0.96464 | 0.85 | no | 0.97372 | no | 0.95482 | no | 0.84095 | no | 0.028842 | 2.0571e-07 | 0 | -0.00030049 | no | 0 | -0.00039375 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150227_M02641_0032_000000000-ACNF3 | M02641 | 2015-02-27 | 75 | 75 | 90.2941 | no | 25317785 | low | 1066.3007 | very low | 0.076375 | yes | 90.5194 | no | 0.95732 | 0.85 | no | 0.9695 | no | 0.94388 | no | 0.84095 | no | 0.0023737 | nan | 0 | -0.00027283 | no | 0 | -0.00038603 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150302_M02641_0033_000000000-ACC9K | M02641 | 2015-03-02 | 151 | 151 | 86.4847 | no | 20383236 | low | 1091.5584 | OK | 0.0091317 | no | 86.4229 | no | 0.8794 | 0.8 | no | 0.92333 | no | 0.83429 | no | 0.95413 | no | 4.0711e-39 | 8.175e-86 | 0 | -0.00085422 | yes | 0 | -0.00080079 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150306_M00766_0080_000000000-ACNEP | M00766 | 2015-03-06 | 75 | 75 | 94.106 | no | 26017554 | low | 1059.9167 | very low | 0.02142 | no | 94.169 | no | 0.97008 | 0.85 | no | 0.97849 | no | 0.96173 | no | 0.84095 | no | 0.078515 | 0.00029715 | 0 | -0.00028625 | no | 0 | -0.00023866 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150306_M02641_0034_000000000-ACRUF | M02641 | 2015-03-06 | 75 | 75 | 91.3955 | no | 24975924 | low | 1049.1583 | very low | 0.0067808 | no | 91.297 | no | 0.95124 | 0.85 | no | 0.9701 | no | 0.93153 | no | 0.95413 | no | 3.4327e-07 | nan | 0 | -0.00040766 | no | 0 | -0.00082529 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150313_M02641_0037_000000000-ACC8R | M02641 | 2015-03-13 | 151 | 151 | 94.5096 | no | 17647098 | low | 906.6463 | low | 0.017228 | no | 93.7282 | no | 0.94213 | 0.8 | no | 0.9595 | no | 0.9242 | no | 0.67794 | no | 2.7478e-10 | nan | 0 | -0.00051188 | yes | 0 | -0.00060154 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150325_M00766_0084_000000000-AD8KN | M00766 | 2015-03-25 | 75 | 75 | 93.0443 | no | 26726137 | low | 1096.1198 | very low | 0.025085 | no | 93.5238 | no | 0.96479 | 0.85 | no | 0.97449 | no | 0.95451 | no | 0.84095 | no | 0.047651 | 0.00076142 | 0 | -0.00034315 | no | 0 | -0.00025609 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150325_M02641_0038_000000000-AD8KY | M02641 | 2015-03-25 | 75 | 75 | 92.8928 | no | 28292154 | low | 1150.8931 | low | 0.016396 | no | 92.9934 | no | 0.96863 | 0.85 | no | 0.97441 | no | 0.9626 | no | 0.95413 | no | 0.97963 | nan | 0 | -0.00030202 | no | 0 | -0.00017631 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150401_M00766_0086_000000000-ACBU5 | M00766 | 2015-04-01 | 151 | 151 | 90.4494 | no | 18801773 | low | 986.6817 | low | 0.029182 | no | 90.7591 | no | 0.8401 | 0.8 | no | 0.87195 | no | 0.80659 | no | 0.99584 | no | 6.8105e-11 | 1.2646e-33 | 0 | -0.0022957 | yes | 6 | -0.0022224 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150407_M00766_0088_000000000-ACCAT | M00766 | 2015-04-07 | 151 | 151 | 86.0053 | no | 20081955 | low | 1075.4261 | OK | 0.031951 | no | 86.1466 | no | 0.8006 | 0.8 | no | 0.8519 | no | 0.74788 | yes | 0.95413 | no | 6.3802e-32 | nan | 0 | -0.002369 | yes | 12 | -0.001854 | yes | pass | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same
150410_M00766_0089_000000000-ACBU4 | M00766 | 2015-04-10 | 151 | 151 | 94.4563 | no | 16861157 | low | 867.0412 | low | 0.021851 | no | 94.4152 | no | 0.93636 | 0.8 | no | 0.9672 | no | 0.90877 | no | 0.50766 | no | 1.738e-17 | nan | 0 | -0.00043159 | no | 0 | -0.001257 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150423_M02641_0041_000000000-ACN0P | M02641 | 2015-04-23 | 75 | 75 | 91.9691 | no | 30009055 | low | 1228.4129 | OK | 0.027957 | no | 92.2905 | no | 0.9501 | 0.85 | no | 0.96725 | no | 0.93306 | no | 0.84095 | no | 2.6066e-06 | 3.4467e-12 | 0 | -0.00047255 | no | 0 | -0.00049099 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150429_M02641_0043_000000000-AD8KJ | M02641 | 2015-04-29 | 75 | 75 | 77.8177 | yes | 32906358 | low | 1415.0412 | high | 0.018771 | no | 78.178 | yes | 0.8983 | 0.85 | no | 0.92026 | no | 0.87594 | no | 0.84095 | no | 0.00033968 | nan | 0 | -0.00089215 | yes | 0 | -0.00067513 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150430_M02641_0044_000000000-AD8VC | M02641 | 2015-04-30 | 75 | 75 | 86.9368 | no | 27790413 | low | 1196.1245 | low | 0.022911 | no | 87.4623 | no | 0.92905 | 0.85 | no | 0.95387 | no | 0.90368 | no | 0.84095 | no | 1.2394e-10 | 1.7232e-11 | 0 | -0.00060018 | yes | 0 | -0.0005903 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150501_M00766_0093_000000000-AF9MH | M00766 | 2015-05-01 | 75 | 75 | 87.2253 | no | 34467195 | low | 1450.6461 | high | 0.022012 | no | 87.0197 | no | 0.94139 | 0.85 | no | 0.95639 | no | 0.92569 | no | 0.84095 | no | 0.00057945 | 2.9759e-05 | 0 | -0.00056762 | yes | 0 | -0.00052432 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150501_M02641_0045_000000000-AD6UA | M02641 | 2015-05-01 | 75 | 75 | 80.8519 | no | 25913017 | low | 1131.0451 | low | 0.015126 | no | 81.3041 | yes | 0.93089 | 0.85 | no | 0.92992 | no | 0.93269 | no | 0.99998 | no | 0.013057 | nan | 0 | -0.00054803 | yes | 0 | -0.00041428 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff
150507_M00766_0095_000000000-AF7N3 | M00766 | 2015-05-07 | 75 | 75 | 92.6658 | no | 27410104 | low | 1119.7122 | low | 0.036603 | no | 93.5327 | no | 0.96159 | 0.85 | no | 0.9741 | no | 0.94953 | no | 0.84095 | no | 0.00099419 | 8.626e-08 | 0 | -0.00036389 | no | 0 | -0.00042845 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150507_M02641_0046_000000000-AF7N5 | M02641 | 2015-05-07 | 75 | 75 | 92.9775 | no | 27595623 | low | 1124.3595 | low | 0.040298 | no | 93.1424 | no | 0.95227 | 0.85 | no | 0.96981 | no | 0.93556 | no | 0.84095 | no | 9.3853e-07 | 5.8544e-20 | 0 | -0.00053503 | yes | 0 | -0.0001815 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150515_M00766_0098_000000000-AF9ND | M00766 | 2015-05-15 | 75 | 75 | 96.8787 | no | 15942131 | low | 638.6353 | very low | 0.030068 | no | 96.9762 | no | 0.97995 | 0.85 | no | 0.98705 | no | 0.97446 | no | 0.84095 | no | 0.57572 | nan | 0 | -0.00024399 | no | 0 | -0.00016041 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150518_M00766_0100_000000000-AATJ7 | M00766 | 2015-05-18 | 151 | 151 | 89.6123 | no | 7467630 | low | 351.2212 | very low | 0.03759 | no | 88.7216 | no | 0.94106 | 0.8 | no | 0.95572 | no | 0.92587 | no | 0.84095 | no | 0.55116 | 0.39443 | 0 | -0.00037129 | no | 0 | -0.00084743 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150521_M00766_0102_000000000-AF9MV | M00766 | 2015-05-21 | 75 | 75 | 89.6458 | no | 30563261 | low | 1273.6592 | OK | 0.043729 | no | 90.9553 | no | 0.95232 | 0.85 | no | 0.96492 | no | 0.93927 | no | 0.84095 | no | 0.0097477 | nan | 0 | -0.00050706 | yes | 0 | -0.00046026 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150526_M02641_0053_000000000-AFHDV | M02641 | 2015-05-26 | 200 | 200 | 48.4208 | yes | 7559009 | low | 310.8307 | very low | 0.22965 | yes | 50.5562 | yes | 0.64122 | 0.7 | yes | 0.51285 | yes | 0.75917 | no | 0.50766 | no | 1.0758e-55 | 2.1338e-299 | 146 | -0.0039621 | yes | 1 | -0.0018861 | yes | fail | fail | fail | fail | diff | pass | exceeded | not exceeded | diff | same
150528_M00766_0105_000000000-AF9MJ | M00766 | 2015-05-28 | 75 | 75 | 91.4093 | no | 30434248 | low | 1259.146 | OK | 0.033027 | no | 91.4154 | no | 0.9591 | 0.85 | no | 0.96908 | no | 0.94896 | no | 0.84095 | no | 0.011641 | 5.539e-05 | 0 | -0.00043918 | no | 0 | -0.00032915 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150529_M00766_0106_000000000-AEVP5 | M00766 | 2015-05-29 | 151 | 151 | 88.3589 | no | 15502868 | low | 806.7578 | very low | 0.032653 | no | 89.3692 | no | 0.91331 | 0.8 | no | 0.92662 | no | 0.89824 | no | 0.99584 | no | 0.34729 | 5.8531e-07 | 0 | -0.00082214 | yes | 0 | -0.0010282 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150605_M02641_0058_000000000-AFF15 | M02641 | 2015-06-05 | 75 | 75 | 86.9846 | no | 20446753 | low | 822.314 | very low | 0.026851 | no | 87.5712 | no | 0.93738 | 0.85 | no | 0.94849 | no | 0.92669 | no | 0.95413 | no | 0.71981 | nan | 0 | -0.00062227 | yes | 0 | -0.00041516 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150610_M02641_0059_000000000-AFN4H | M02641 | 2015-06-10 | 75 | 75 | 82.3028 | no | 25297112 | low | 1022.9657 | very low | 0.024105 | no | 83.0805 | yes | 0.91102 | 0.85 | no | 0.93192 | no | 0.89027 | no | 0.84095 | no | 0.00015742 | 7.1708e-18 | 0 | -0.00077999 | yes | 0 | -0.00063882 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150612_M00766_0112_000000000-AFMW5 | M00766 | 2015-06-12 | 75 | 75 | 85.2766 | no | 34132691 | low | 1448.1255 | high | 0.019585 | no | 85.0541 | no | 0.93156 | 0.85 | no | 0.95081 | no | 0.91105 | no | 0.84095 | no | 3.193e-06 | 6.9549e-35 | 0 | -0.0006103 | yes | 0 | -0.00064344 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150612_M02641_0061_000000000-AFMVT | M02641 | 2015-06-12 | 75 | 75 | 79.341 | yes | 25070490 | low | 1016.9582 | very low | 0.03704 | no | 80.1022 | yes | 0.90347 | 0.85 | no | 0.92645 | no | 0.88019 | no | 0.84095 | no | 3.0151e-05 | 3.0045e-11 | 0 | -0.00078584 | yes | 0 | -0.00069582 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150615_M00766_0113_000000000-AFBH3 | M00766 | 2015-06-15 | 251 | 251 | 89.384 | no | 20944521 | low | 1111.4379 | OK | 0.037506 | no | 89.0846 | no | 0.71363 | 0.75 | yes | 0.81632 | no | 0.60535 | yes | 0.50766 | no | 3.3983e-154 | 5.608e-166 | 24 | -0.0023213 | yes | 53 | -0.004186 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same
150619_M00766_0114_000000000-AFMYN | M00766 | 2015-06-19 | 75 | 75 | 92.1565 | no | 28125244 | low | 1158.4526 | low | 0.034411 | no | 93.2083 | no | 0.95863 | 0.85 | no | 0.97242 | no | 0.94406 | no | 0.84095 | no | 0.00010309 | nan | 0 | -0.00041056 | no | 0 | -0.00058758 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150622_M00766_0115_000000000-AFN2H | M00766 | 2015-06-22 | 75 | 75 | 93.2875 | no | 26234040 | low | 1075.6113 | very low | 0.020578 | no | 93.4073 | no | 0.96472 | 0.85 | no | 0.97473 | no | 0.95557 | no | 0.84095 | no | 0.048507 | 0.0078432 | 0 | -0.00038391 | no | 0 | -0.0002776 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150623_M00766_0116_000000000-AFLEW | M00766 | 2015-06-23 | 151 | 151 | 87.7026 | no | 20733081 | low | 1104.0961 | OK | 0.034923 | no | 86.8671 | no | 0.84466 | 0.8 | no | 0.89266 | no | 0.79387 | yes | 0.84095 | no | 1.473e-43 | 7.7036e-97 | 0 | -0.0019268 | yes | 9 | -0.0023623 | yes | pass | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same
150706_M00766_0118_000000000-AF512 | M00766 | 2015-07-06 | 200 | 200 | 94.4229 | no | 16447526 | low | 844.1499 | very low | 0.016368 | no | 94.5523 | no | 0.86041 | 0.75 | no | 0.89416 | no | 0.82397 | no | 0.84095 | no | 3.1972e-28 | nan | 1 | -0.001537 | yes | 1 | -0.0023974 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150729_M02641_0008_000000000-AGET9 | M02641 | 2015-07-29 | 200 | 200 | 91.862 | no | 30439243 | low | 1258.6104 | OK | 0.012085 | no | 92.3982 | no | 0.84428 | 0.7 | no | 0.87341 | no | 0.81136 | no | 0.67794 | no | 1.7873e-34 | 2.8556e-270 | 1 | -0.0016885 | yes | 2 | -0.0024684 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150730_M00766_0120_000000000-AFN4E | M00766 | 2015-07-30 | 75 | 75 | 94.0611 | no | 23945923 | low | 972.2454 | very low | 0.052603 | yes | 94.2795 | no | 0.9647 | 0.85 | no | 0.97668 | no | 0.95322 | no | 0.84095 | no | 0.0076657 | nan | 0 | -0.00038193 | no | 0 | -0.0003459 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150731_M00766_0121_000000000-AFMYC | M00766 | 2015-07-31 | 75 | 75 | 89.5212 | no | 27660854 | low | 1152.6219 | low | 0.036004 | no | 89.6942 | no | 0.94787 | 0.85 | no | 0.95439 | no | 0.94056 | no | 0.99584 | no | 0.98595 | nan | 0 | -0.00062082 | yes | 0 | -0.00041404 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150731_M02641_0009_000000000-AFN3M | M02641 | 2015-07-31 | 75 | 75 | 90.2753 | no | 32967275 | low | 1364.8961 | OK | 0.020875 | no | 90.6855 | no | 0.9485 | 0.85 | no | 0.96322 | no | 0.93353 | no | 0.84095 | no | 0.00029106 | 1.2672e-06 | 0 | -0.00052951 | yes | 0 | -0.00044329 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150807_M02641_0012_000000000-AFMYT | M02641 | 2015-08-07 | 75 | 75 | 96.8172 | no | 16466109 | low | 661.2988 | very low | 0.043831 | no | 97.2034 | no | 0.97599 | 0.85 | no | 0.98258 | no | 0.97147 | no | 0.95413 | no | 0.81203 | 0.23825 | 0 | -0.00028401 | no | 0 | -0.0001526 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150810_M00766_0123_000000000-AGTUY | M00766 | 2015-08-10 | 151 | 151 | 89.2945 | no | 22283310 | low | 1188.9462 | OK | 0.016194 | no | 88.9362 | no | 0.92177 | 0.8 | no | 0.95497 | no | 0.88475 | no | 0.84095 | no | 2.7093e-32 | nan | 0 | -0.00035243 | no | 0 | -0.00083315 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150811_M02641_0014_000000000-AGK0F | M02641 | 2015-08-11 | 151 | 151 | 92.4253 | no | 22253985 | low | 1160.0949 | OK | 0.016889 | no | 92.3445 | no | 0.90106 | 0.8 | no | 0.9283 | no | 0.87132 | no | 0.95413 | no | 7.8772e-25 | nan | 0 | -0.0009934 | yes | 0 | -0.0015021 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150819_M02641_0016_000000000-AGJHU | M02641 | 2015-08-19 | 75 | 75 | 94.1214 | no | 26775268 | low | 1083.954 | very low | 0.032934 | no | 94.4666 | no | 0.96462 | 0.85 | no | 0.97519 | no | 0.95489 | no | 0.84095 | no | 0.013871 | 2.5966e-09 | 0 | -0.00041278 | no | 0 | -0.00025665 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150820_M00766_0126_000000000-AGHUU | M00766 | 2015-08-20 | 75 | 75 | 92.7672 | no | 28026234 | low | 1148.9642 | low | 0.020853 | no | 93.2904 | no | 0.96152 | 0.85 | no | 0.97405 | no | 0.94902 | no | 0.84095 | no | 0.00031602 | nan | 0 | -0.00033273 | no | 0 | -0.0004084 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150820_M02641_0017_000000000-AGJJA | M02641 | 2015-08-20 | 75 | 75 | 93.4116 | no | 26592471 | low | 1089.8361 | very low | 0.015153 | no | 93.7457 | no | 0.96481 | 0.85 | no | 0.97351 | no | 0.95717 | no | 0.84095 | no | 0.072701 | 0.016306 | 0 | -0.00030311 | no | 0 | -0.00031655 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150821_M02641_0018_000000000-AGJNU | M02641 | 2015-08-21 | 151 | 151 | 96.3232 | no | 11702340 | low | 597.5554 | very low | 0.030041 | no | 96.5218 | no | 0.96101 | 0.8 | no | 0.97032 | no | 0.95086 | no | 0.95413 | no | 0.0070316 | nan | 0 | -0.00034988 | no | 0 | -0.000607 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150902_M00766_0128_000000000-AFLDG | M00766 | 2015-09-02 | 151 | 151 | 89.3025 | no | 21077655 | low | 1129.9527 | OK | 0.015246 | no | 89.1889 | no | 0.83273 | 0.8 | no | 0.88185 | no | 0.78 | yes | 0.67794 | no | 7.5008e-47 | 3.9097e-271 | 2 | -0.0022787 | yes | 13 | -0.0028073 | yes | pass | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same
150904_M02641_0022_000000000-AH991 | M02641 | 2015-09-04 | 200 | 200 | 90.7156 | no | 32586889 | low | 1348.7337 | OK | 0.0084128 | no | 90.8839 | no | 0.80842 | 0.7 | no | 0.83792 | no | 0.774 | no | 0.50766 | no | 6.5299e-46 | 4.3193e-60 | 33 | -0.002385 | yes | 45 | -0.0029542 | yes | pass | pass | pass | pass | same | pass | exceeded | exceeded | same | same
150910_M02641_0023_000000000-AFLHH | M02641 | 2015-09-10 | 151 | 151 | 89.5554 | no | 20794433 | low | 1100.488 | OK | 0.02384 | no | 88.8672 | no | 0.85359 | 0.8 | no | 0.88668 | no | 0.81835 | no | 0.95413 | no | 2.3039e-18 | nan | 0 | -0.0021309 | yes | 9 | -0.0024176 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150911_M00766_0131_000000000-AGLDT | M00766 | 2015-09-11 | 75 | 75 | 91.8669 | no | 29894066 | low | 1233.1842 | OK | 0.017257 | no | 92.0012 | no | 0.95954 | 0.85 | no | 0.97065 | no | 0.94872 | no | 0.84095 | no | 0.0036196 | 0.00029275 | 0 | -0.0004291 | no | 0 | -0.00031438 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150911_M02641_0024_AF50J | M02641 | 2015-09-11 | 151 | 151 | 91.3886 | no | 20193785 | low | 1043.6462 | OK | 0.026601 | no | 91.1381 | no | 0.85976 | 0.8 | no | 0.89102 | no | 0.82694 | no | 0.84095 | no | 4.1997e-18 | nan | 0 | -0.0020335 | yes | 5 | -0.0021409 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150914_M00766_0132_000000000-AF41F | M00766 | 2015-09-14 | 151 | 151 | 86.8892 | no | 3991745 | low | 184.4777 | very low | 0.012611 | no | 87.677 | no | 0.95124 | 0.8 | no | 0.96221 | no | 0.94056 | no | 0.95413 | no | 0.17881 | 0.0092674 | 0 | -0.00035999 | no | 0 | -0.00086178 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150925_M02641_0028_000000000-AGK08 | M02641 | 2015-09-25 | 75 | 75 | 95.0618 | no | 24906667 | low | 1010.387 | very low | 0.01775 | no | 95.4096 | no | 0.96734 | 0.85 | no | 0.97861 | no | 0.95784 | no | 0.84095 | no | 0.0070482 | 1.5449e-09 | 0 | -0.00035963 | no | 0 | -0.00022766 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150928_M02641_0029_000000000-AGHV5 | M02641 | 2015-09-28 | 75 | 75 | 93.773 | no | 28292454 | low | 1144.9595 | low | 0.021423 | no | 93.5295 | no | 0.96266 | 0.85 | no | 0.97493 | no | 0.95191 | no | 0.84095 | no | 0.012643 | nan | 0 | -0.00041356 | no | 0 | -0.00026611 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150930_M02641_0031_000000000-AGJGT | M02641 | 2015-09-30 | 75 | 75 | 93.3234 | no | 29303547 | low | 1206.8802 | OK | 0.049805 | no | 93.5156 | no | 0.9573 | 0.85 | no | 0.9715 | no | 0.94372 | no | 0.84095 | no | 7.6841e-05 | nan | 0 | -0.00045184 | no | 0 | -0.00030011 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151001_M02641_0032_000000000-AGLE0 | M02641 | 2015-10-01 | 75 | 75 | 87.3795 | no | 35725439 | low | 1504.6951 | very high | 0.017939 | no | 86.8081 | no | 0.9359 | 0.85 | no | 0.95659 | no | 0.91522 | no | 0.84095 | no | 5.6512e-08 | nan | 0 | -0.00058153 | yes | 0 | -0.00045814 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151002_M02641_0033_000000000-AFLDA | M02641 | 2015-10-02 | 151 | 151 | 88.0983 | no | 23364962 | low | 1229.9955 | high | 0.042476 | no | 87.9251 | no | 0.90553 | 0.8 | no | 0.937 | no | 0.87138 | no | 0.84095 | no | 1.4773e-23 | 1.9239e-34 | 0 | -0.00058099 | yes | 0 | -0.0010877 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151009_M00766_0140_000000000-AGK0B | M00766 | 2015-10-09 | 75 | 75 | 87.3055 | no | 24687917 | low | 1048.2614 | very low | 0.024799 | no | 87.0808 | no | 0.95594 | 0.85 | no | 0.96314 | no | 0.94902 | no | 0.99584 | no | 0.95838 | 0.16034 | 0 | -0.00045147 | no | 0 | -0.00036923 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
151009_M02641_0036_000000000-AFN30 | M02641 | 2015-10-09 | 151 | 151 | 91.4518 | no | 21458630 | low | 1123.3539 | OK | 0.024973 | no | 91.2766 | no | 0.86713 | 0.8 | no | 0.89737 | no | 0.83419 | no | 0.95413 | no | 2.2346e-20 | 0 | 0 | -0.0020801 | yes | 10 | -0.0026428 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151012_M02641_0037_000000000-AFNA5 | M02641 | 2015-10-12 | 151 | 151 | 93.0707 | no | 19894718 | low | 1031.818 | OK | 0.019111 | no | 93.4892 | no | 0.88194 | 0.8 | no | 0.90915 | no | 0.85233 | no | 0.95413 | no | 1.2493e-20 | 1.0755e-37 | 0 | -0.0019675 | yes | 8 | -0.0026129 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151019_M02641_0040_000000000-AJDVH | M02641 | 2015-10-19 | 151 | 151 | 90.6324 | no | 24425659 | OK | 1287.4209 | very high | 0.019011 | no | 90.7816 | no | 0.92696 | 0.8 | no | 0.94837 | no | 0.90505 | no | 0.95413 | no | 6.5136e-18 | nan | 0 | -0.00048911 | no | 0 | -0.00086419 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151021_M00766_0144_000000000-AG028 | M00766 | 2015-10-21 | 75 | 75 | 85.0537 | no | 24774347 | low | 1085.7652 | very low | 0.013905 | no | 84.6897 | yes | 0.94129 | 0.85 | no | 0.95646 | no | 0.92591 | no | 0.84095 | no | 0.0030374 | 4.0367e-11 | 0 | -0.00055587 | yes | 0 | -0.00055374 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
151022_M02641_0042_000000000-AJ5B5 | M02641 | 2015-10-22 | 200 | 200 | 89.6462 | no | 28861426 | low | 1217.8503 | OK | 0.010571 | no | 89.8396 | no | 0.67585 | 0.7 | yes | 0.7507 | no | 0.59122 | yes | 0.67794 | no | 4.6203e-88 | 7.0766e-116 | 12 | -0.0033975 | yes | 75 | -0.0048318 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same
151026_M00766_0148_000000000-AFLVV | M00766 | 2015-10-26 | 151 | 151 | 88.3189 | no | 18895886 | low | 1013.3256 | OK | 0.032209 | no | 89.0093 | no | 0.93493 | 0.8 | no | 0.95661 | no | 0.91268 | no | 0.84095 | no | 9.4968e-15 | 5.2193e-112 | 0 | -0.00031813 | no | 0 | -0.00073043 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151102_M02641_0045_000000000-AFN3H | M02641 | 2015-11-02 | 151 | 151 | 94.4283 | no | 15063309 | low | 779.0627 | very low | 0.048767 | no | 94.8292 | no | 0.95449 | 0.8 | no | 0.96887 | no | 0.93956 | no | 0.84095 | no | 2.4447e-05 | 2.8037e-38 | 0 | -0.00027595 | no | 0 | -0.00058311 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151104_M00766_0151_000000000-AK8EW | M00766 | 2015-11-04 | 151 | 151 | 82.696 | no | 16143415 | low | 883.0439 | low | 0.069792 | yes | 83.2962 | yes | 0.91934 | 0.8 | no | 0.94177 | no | 0.89677 | no | 0.95413 | no | 1.0753e-07 | 5.9859e-42 | 0 | -0.00035735 | no | 0 | -0.00067331 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff
151106_M00766_0152_000000000-AJ5W5 | M00766 | 2015-11-06 | 151 | 151 | 89.6104 | no | 16168604 | low | 863.0096 | low | 0.024136 | no | 90.7829 | no | 0.9317 | 0.8 | no | 0.95041 | no | 0.91232 | no | 0.95413 | no | 3.0502e-07 | 1.2764e-25 | 0 | -0.00049715 | no | 0 | -0.00054569 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151111_M02641_0049_000000000-AJD47 | M02641 | 2015-11-11 | 200 | 200 | 92.8535 | no | 30208881 | low | 1238.2172 | OK | 0.021332 | no | 93.316 | no | 0.83991 | 0.7 | no | 0.87018 | no | 0.80592 | no | 0.50766 | no | 8.1049e-40 | 6.164e-109 | 14 | -0.0019638 | yes | 19 | -0.0027647 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151116_M00766_0154_000000000-AJF4C | M00766 | 2015-11-16 | 151 | 151 | 87.9376 | no | 13835502 | low | 734.2783 | very low | 0.054577 | yes | 88.5491 | no | 0.9313 | 0.8 | no | 0.95109 | no | 0.90966 | no | 0.95413 | no | 5.1553e-08 | nan | 0 | -0.00033626 | no | 0 | -0.00061836 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff
151119_M02641_0052_000000000-AJJ6B | M02641 | 2015-11-19 | 75 | 75 | 94.5733 | no | 24895626 | low | 1016.266 | very low | 0.029336 | no | 94.9966 | no | 0.96439 | 0.85 | no | 0.97782 | no | 0.95208 | no | 0.84095 | no | 0.00044318 | 1.1267e-06 | 0 | -0.00028501 | no | 0 | -0.00037122 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
151125_M00766_0158_000000000-AJJC1 | M00766 | 2015-11-25 | 151 | 151 | 78.1687 | yes | 3563033 | low | 150.7739 | very low | 0.031274 | no | 78.1863 | yes | 0.9558 | 0.8 | no | 0.96712 | no | 0.94724 | no | 0.95413 | no | 0.17297 | 0.010728 | 0 | -0.00022353 | no | 0 | -0.00043897 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
151130_M00766_0163_000000000-AJGC9 | M00766 | 2015-11-30 | 151 | 151 | 79.0357 | yes | 21816702 | low | 1234.7264 | high | 0.043221 | no | 79.6269 | yes | 0.91879 | 0.8 | no | 0.94269 | no | 0.8946 | no | 0.95413 | no | 4.5777e-17 | 3.5682e-28 | 0 | -0.0002497 | no | 0 | -0.00046641 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
151204_M02641_0056_000000000-AJDC6 | M02641 | 2015-12-04 | 200 | 200 | 93.7017 | no | 25575432 | low | 1042.2805 | very low | 0.036678 | no | 94.2736 | no | 0.85421 | 0.7 | no | 0.88062 | no | 0.82404 | no | 0.50766 | no | 2.3201e-29 | nan | 13 | -0.0018496 | yes | 19 | -0.0025865 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
151208_M02641_0057_000000000-AJHRF | M02641 | 2015-12-08 | 151 | 151 | 97.1211 | no | 4496775 | low | 228.8877 | very low | 0.02523 | no | 97.1074 | no | 0.94501 | 0.8 | no | 0.95817 | no | 0.93055 | no | 0.95413 | no | 0.083233 | 0.003379 | 0 | -0.0007036 | yes | 0 | -0.00058399 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
151210_M00766_0168_000000000-AJDDH | M00766 | 2015-12-10 | 75 | 75 | 91.376 | no | 26145512 | low | 1086.9972 | very low | 0.021546 | no | 90.8931 | no | 0.956 | 0.85 | no | 0.96684 | no | 0.94501 | no | 0.84095 | no | 0.10224 | 0.014367 | 0 | -0.00051081 | yes | 0 | -0.0004594 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151211_M00766_0169_000000000-AK5DW | M00766 | 2015-12-11 | 300 | 300 | 93.0016 | no | 21872195 | low | 891.8405 | very low | 0.023515 | no | 93.6566 | no | 0.54826 | 0.7 | yes | 0.59108 | yes | 0.48586 | yes | 0.24079 | no | 3.3919e-41 | 2.4542e-59 | 169 | -0.0035153 | yes | 181 | -0.004229 | yes | fail | fail | fail | fail | same | pass | exceeded | exceeded | same | same
151218_M00766_0171_000000000-AK5FL | M00766 | 2015-12-18 | 226 | 226 | 92.0182 | no | 20535429 | low | 846.6865 | very low | 0.020936 | no | 91.9926 | no | 0.57401 | 0.7 | yes | 0.58212 | yes | 0.54171 | yes | 0.67794 | no | 3.7819e-14 | 8.2184e-29 | 24 | -0.005182 | yes | 97 | -0.0055294 | yes | fail | fail | fail | fail | same | pass | exceeded | exceeded | same | same
160104_M02641_0062_000000000-AL603 | M02641 | 2016-01-04 | 101 | 101 | 93.4734 | no | 19858211 | low | 1030.1337 | OK | 0.01816 | no | 93.4478 | no | 0.95548 | 0.8 | no | 0.97027 | no | 0.94346 | no | 0.84095 | no | 1.6436e-05 | 1.1366e-31 | 0 | -0.00025903 | no | 0 | -0.00019241 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160105_M00766_0174_000000000-AL4LB | M00766 | 2016-01-05 | 151 | 151 | 91.2681 | no | 14218489 | low | 749.2227 | very low | 0.036601 | no | 91.4658 | no | 0.93032 | 0.8 | no | 0.95273 | no | 0.9068 | no | 0.95413 | no | 1.8718e-11 | nan | 0 | -0.00053469 | yes | 0 | -0.00071583 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
160107_M02641_0063_000000000-AJDC9 | M02641 | 2016-01-07 | 75 | 75 | 94.1624 | no | 24963706 | low | 1022.5057 | very low | 0.017799 | no | 94.5508 | no | 0.9694 | 0.85 | no | 0.97637 | no | 0.96261 | no | 0.84095 | no | 0.22798 | nan | 0 | -0.00029644 | no | 0 | -0.00032503 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
160108_M02641_0064_000000000-ALNH2 | M02641 | 2016-01-08 | 200 | 200 | 91.9702 | no | 21242610 | low | 1109.3897 | OK | 0.022049 | no | 92.3443 | no | 0.8579 | 0.75 | no | 0.90283 | no | 0.81012 | no | 0.67794 | no | 5.5542e-59 | nan | 0 | -0.0013823 | yes | 1 | -0.0025063 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160115_M00766_0178_000000000-ALKJV | M00766 | 2016-01-15 | 151 | 151 | 83.9557 | no | 15235349 | low | 831.5431 | very low | 0.066527 | yes | 83.4762 | yes | 0.91226 | 0.8 | no | 0.93609 | no | 0.8844 | no | 0.95413 | no | 5.102e-09 | nan | 0 | -0.00057903 | yes | 0 | -0.0011478 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
160118_M02641_0066_000000000-AJETF | M02641 | 2016-01-18 | 75 | 75 | 98.1898 | no | 9321517 | low | 373.3045 | very low | 0.028008 | no | 98.3733 | no | 0.97492 | 0.85 | no | 0.97945 | no | 0.97382 | no | 0.99998 | no | 1 | 0.5283 | 0 | -0.00022747 | no | 0 | -0.000157 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
160119_M02641_0067_000000000-ALN4P | M02641 | 2016-01-19 | 151 | 151 | 87.2125 | no | 18949772 | low | 995.7574 | low | 0.091472 | yes | 86.1573 | no | 0.9035 | 0.8 | no | 0.92594 | no | 0.87789 | no | 0.95413 | no | 2.2467e-08 | nan | 0 | -0.00078266 | yes | 0 | -0.0012195 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
160127_M00766_0001_000000000-ALR2L | M00766 | 2016-01-27 | 151 | 151 | 94.4219 | no | 16728171 | low | 869.4499 | low | 0.024493 | no | 94.8131 | no | 0.94892 | 0.8 | no | 0.95934 | no | 0.93843 | no | 0.95413 | no | 5.3787e-06 | nan | 0 | -0.00045839 | no | 0 | -0.00066696 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
160127_M02641_0069_000000000-AMDH4 | M02641 | 2016-01-27 | 75 | 75 | 76.69 | yes | 29747487 | low | 1328.6387 | OK | 0.096919 | yes | 78.2455 | yes | 0.91416 | 0.85 | no | 0.93371 | no | 0.89378 | no | 0.84095 | no | 0.00062682 | nan | 0 | -0.00074637 | yes | 0 | -0.00037367 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff
160128_M00766_0002_000000000-ALAJ1 | M00766 | 2016-01-28 | 75 | 75 | 88.2213 | no | 28975202 | low | 1238.6174 | OK | 0.0080531 | no | 88.3898 | no | 0.95207 | 0.85 | no | 0.96129 | no | 0.94181 | no | 0.84095 | no | 0.078659 | 0.012869 | 0 | -0.00042109 | no | 0 | -0.00041041 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160204_M00766_0004_000000000-AL4H0 | M00766 | 2016-02-04 | 151 | 151 | 95.4825 | no | 14166224 | low | 726.9693 | very low | 0.021398 | no | 95.3691 | no | 0.95549 | 0.8 | no | 0.97197 | no | 0.93852 | no | 0.84095 | no | 3.1187e-14 | 2.8941e-31 | 0 | -0.00029231 | no | 0 | -0.00059973 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
160204_M02641_0072_000000000-AK6CE | M02641 | 2016-02-04 | 75 | 75 | 91.6672 | no | 25918178 | low | 1060.0101 | very low | 0.015091 | no | 91.5669 | no | 0.95246 | 0.85 | no | 0.96475 | no | 0.93912 | no | 0.84095 | no | 0.00023857 | 2.281e-41 | 0 | -0.00053931 | yes | 0 | -0.00055384 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
160205_M02641_0073_000000000-ALY9Y | M02641 | 2016-02-05 | 75 | 75 | 81.6149 | no | 30119640 | low | 1331.3992 | OK | 0.025391 | no | 80.9054 | yes | 0.92844 | 0.85 | no | 0.94565 | no | 0.90982 | no | 0.84095 | no | 0.023743 | nan | 0 | -0.0004287 | no | 0 | -0.00027127 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
160208_M00766_0006_000000000-AMF4G | M00766 | 2016-02-08 | 75 | 75 | 85.7346 | no | 32902642 | low | 1404.3574 | high | 0.015626 | no | 85.8175 | no | 0.93494 | 0.85 | no | 0.95198 | no | 0.91641 | no | 0.84095 | no | 1.9762e-05 | 5.3645e-08 | 0 | -0.00062874 | yes | 0 | -0.00051979 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160208_M02641_0075_000000000-AMF40 | M02641 | 2016-02-08 | 75 | 75 | 79.3999 | yes | 30335883 | low | 1362.3448 | OK | 0.022595 | no | 80.0194 | yes | 0.9085 | 0.85 | no | 0.93381 | no | 0.88187 | no | 0.84095 | no | 9.8314e-08 | 2.2541e-295 | 0 | -0.00078696 | yes | 0 | -0.00037771 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff
160209_M02641_0076_000000000-AMF44 | M02641 | 2016-02-09 | 75 | 75 | 89.416 | no | 23928513 | low | 1023.4734 | very low | 0.02047 | no | 89.4341 | no | 0.95438 | 0.85 | no | 0.96439 | no | 0.94402 | no | 0.84095 | no | 0.15811 | 4.2223e-80 | 0 | -0.00036997 | no | 0 | -0.00030272 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
160210_M02641_0077_000000000-AME81 | M02641 | 2016-02-10 | 75 | 75 | 93.9536 | no | 19770059 | low | 818.1394 | very low | 0.024181 | no | 94.3148 | no | 0.97229 | 0.85 | no | 0.97917 | no | 0.96582 | no | 0.95413 | no | 0.58184 | 6.9811e-05 | 0 | -0.00021377 | no | 0 | -0.00022485 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
160212_M00766_0008_000000000-AL607 | M00766 | 2016-02-12 | 151 | 151 | 87.5244 | no | 20878061 | low | 1123.9732 | OK | 0.026805 | no | 87.1194 | no | 0.85079 | 0.8 | no | 0.8926 | no | 0.80667 | no | 0.84095 | no | 5.1098e-39 | nan | 0 | -0.002166 | yes | 10 | -0.0030426 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160216_M02641_0079_000000000-AL75M | M02641 | 2016-02-16 | 151 | 151 | 92.5527 | no | 17259426 | low | 905.0069 | low | 0.019188 | no | 93.4136 | no | 0.9397 | 0.8 | no | 0.96218 | no | 0.91691 | no | 0.67794 | no | 4.7356e-16 | nan | 0 | -0.00030351 | no | 0 | -0.00073347 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
160218_M00766_0010_000000000-AMF48 | M00766 | 2016-02-18 | 75 | 75 | 90.3938 | no | 26783167 | low | 1115.6764 | low | 0.019331 | no | 90.2226 | no | 0.95211 | 0.85 | no | 0.96432 | no | 0.93856 | no | 0.84095 | no | 0.0028637 | 0.00027102 | 0 | -0.00054184 | yes | 0 | -0.00048552 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
160218_M02641_0080_000000000-AMF3Y | M02641 | 2016-02-18 | 75 | 75 | 92.7164 | no | 22919923 | low | 955.1644 | very low | 0.010077 | no | 92.7435 | no | 0.95689 | 0.85 | no | 0.97361 | no | 0.93955 | no | 0.84095 | no | 0.00019591 | nan | 0 | -0.00031985 | no | 0 | -0.00022652 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
| 63 |  |  | NUM FAILS |  | 11 |  | 120 |  | 89 |  | 10 |  | 19 |  |  | 9 |  | 5 |  | 16 |  | 0 |  |  | 16 |  | 66 | 33 |  | 73 | 70 | 76 | 28 | 35 | 13 | 0 |  |  | 8 | 29
|  |  |  | % FAILS |  | 9.0164 |  | 98.3607 |  | 72.9508 |  | 8.1967 |  | 15.5738 |  |  | 7.377 |  | 4.0984 |  | 13.1148 |  | 0 |  |  | 13.1148 |  | 54.0984 | 27.0492 |  | 59.8361 | 52 | 46 | 94 | 87
Exclude this one as useless
57
46.7213
[31 empty] | Overall | Method 1 | Method 2 | Method 1 (no CD) | Method 2 (no CD)
[31 empty] | Pass | 52 | 46 | 94 | 87
[31 empty] | Fail | 70 | 76 | 28 | 35

## BadRunsEye
RunID | InstrumentID | RunDate | Pass/Fail | Indexing Comments | General Comments
130417_M00766_0007_000000000-A3R1F | M00766 | 2013-04-17 | Pass
130508_M00766_0013_000000000-A3RB5 | M00766 | 2013-05-08 | Pass
130510_M00766_0014_000000000-A3PFJ | M00766 | 2013-05-10 | Pass | Some samples very low
130613_M00766_0018_000000000-A4FN3 | M00766 | 2013-06-13 | Fail |  | Looks overclustered (High cluster density; lowish quality)
130708_M00766_0023_000000000-A20R8 | M00766 | 2013-07-08 | Pass | Some samples very low
130902_M00766_0045_000000000-A5BVD | M00766 | 2013-09-02 | Pass | Some samples very low
130909_M00766_0047_000000000-A5PEH | M00766 | 2013-09-09 | Pass
140205_M00766_0021_000000000-A7RVR | M00766 | 2014-02-05 | Fail |  | Low >Q30; Low cluster density
140603_M00766_0037_000000000-A7WU9 | M00766 | 2014-06-03 | Fail | One sample very low | Low >Q30; High cluster density; Smearing in heatmap
140620_M02641_0004_000000000-A8R7V | M02641 | 2014-06-20 | Pass
140625_M00766_0041_000000000-A8P78 | M00766 | 2014-06-25 | Fail |  | Low cluster density; Borderline >Q30; Smearing in heatmap
140715_M02641_0008_000000000-A8P81 | M02641 | 2014-07-15 | Borderline Pass |  | Low cluster density
140804_M02641_0011_000000000-AAD6A | M02641 | 2014-08-04 | Fail |  | Low >Q30; Smearing in heatmap
140905_M02641_0017_000000000-AA3FN | M02641 | 2014-09-05 | Pass
140917_M02641_0018_000000000-AA3H7 | M02641 | 2014-09-17 | Pass
141024_M02641_0024_000000000-A8P84 | M02641 | 2014-10-24 | Pass
141113_M02641_0025_000000000-A8RTY | M02641 | 2014-11-13 | Borderline Pass | Some samples very low | Low cluster density
141125_M00766_0062_000000000-A7BKW | M00766 | 2014-11-25 | Pass |  | Read 2 not looking too good
141222_M00766_0067_000000000-ACCB3 | M00766 | 2014-12-22 | Fail | Some samples a little low | Borderline >Q30; Smearing in heatmap; Large difference in cluster density passing filter and cluster density
150112_M00766_0068_000000000-A8PD8 | M00766 | 2015-01-12 | Pass
150114_M00766_0069_000000000-AA68B | M00766 | 2015-01-14 | Fail | One sample a little low | High cluster density; Smearing in heatmap; Large difference in cluster density passing filter and cluster density
150116_M00766_0070_000000000-ACC3G | M00766 | 2015-01-16 | Pass | Some samples a little low
150127_M00766_0072_000000000-ABRK7 | M00766 | 2015-01-27 | Pass
150130_M00766_0073_000000000-ACBVB | M00766 | 2015-01-30 | Fail |  | High cluster density; Low >Q30; Smearing in heatmap
150205_M02641_0029_000000000-ACCE2 | M02641 | 2015-02-05 | Borderline Fail | Some samples very low | Borderline >Q30; Smearing in heatmap
150220_M00766_0077_000000000-AA2U8 | M00766 | 2015-02-20 | Fail | One sample very low | Low >Q30; High cluster density; Smearing in heatmap; Large difference in cluster density passing filter and cluster density
150227_M00766_0079_000000000-ACPR9 | M00766 | 2015-02-27 | Pass
150227_M02641_0032_000000000-ACNF3 | M02641 | 2015-02-27 | Pass | One sample very low
150302_M02641_0033_000000000-ACC9K | M02641 | 2015-03-02 | Borderline Pass |  | All tests borderline
150306_M00766_0080_000000000-ACNEP | M00766 | 2015-03-06 | Pass
150306_M02641_0034_000000000-ACRUF | M02641 | 2015-03-06 | Pass
150313_M02641_0037_000000000-ACC8R | M02641 | 2015-03-13 | Pass
150325_M00766_0084_000000000-AD8KN | M00766 | 2015-03-25 | Pass | One sample very low
150325_M02641_0038_000000000-AD8KY | M02641 | 2015-03-25 | Pass
150401_M00766_0086_000000000-ACBU5 | M00766 | 2015-04-01 | Borderline Pass | Some samples very low | Borderline >Q30; A bit of smearing in heatmap
150407_M00766_0088_000000000-ACCAT | M00766 | 2015-04-07 | Borderline Fail | Some samples very low | Borderline >Q30; A bit of smearing in heatmap
150410_M00766_0089_000000000-ACBU4 | M00766 | 2015-04-10 | Pass
150423_M02641_0041_000000000-ACN0P | M02641 | 2015-04-23 | Pass | One sample very high
150429_M02641_0043_000000000-AD8KJ | M02641 | 2015-04-29 | Fail | A couple of samples low | A bit of smearing in heatmap; reasonable difference in cluster density passing filter and cluster density
150430_M02641_0044_000000000-AD8VC | M02641 | 2015-04-30 | Pass
150501_M00766_0093_000000000-AF9MH | M00766 | 2015-05-01 | Borderline Fail |  | High cluster density
150501_M02641_0045_000000000-AD6UA | M02641 | 2015-05-01 | Borderline Pass | One sample very low | Reasonable difference in cluster density passing filter and cluster density
150507_M00766_0095_000000000-AF7N3 | M00766 | 2015-05-07 | Pass
150507_M02641_0046_000000000-AF7N5 | M02641 | 2015-05-07 | Pass
150515_M00766_0098_000000000-AF9ND | M00766 | 2015-05-15 | Borderline Pass |  | Low cluster density
150518_M00766_0100_000000000-AATJ7 | M00766 | 2015-05-18 | Borderline Pass | Most samples very low | Low cluster density; Minimal smearing in heatmap
150521_M00766_0102_000000000-AF9MV | M00766 | 2015-05-21 | Pass
150526_M02641_0053_000000000-AFHDV | M02641 | 2015-05-26 | Fail | One sample high | Low cluster density; Low >Q30; Smearing in heatmap; Large spread in boxplots, but difference not very high
150528_M00766_0105_000000000-AF9MJ | M00766 | 2015-05-28 | Pass | One sample low, One sample high
150529_M00766_0106_000000000-AEVP5 | M00766 | 2015-05-29 | Pass | All samples low, Some samples very low
150605_M02641_0058_000000000-AFF15 | M02641 | 2015-06-05 | Borderline Pass |  | Low cluster density
150610_M02641_0059_000000000-AFN4H | M02641 | 2015-06-10 | Pass | One sample very low
150612_M00766_0112_000000000-AFMW5 | M00766 | 2015-06-12 | Borderline Fail | One sample low | High cluster density
150612_M02641_0061_000000000-AFMVT | M02641 | 2015-06-12 | Pass
150615_M00766_0113_000000000-AFBH3 | M00766 | 2015-06-15 | Fail |  | Low >Q30; Smearing in heatmap
150619_M00766_0114_000000000-AFMYN | M00766 | 2015-06-19 | Pass | One sample a little low
150622_M00766_0115_000000000-AFN2H | M00766 | 2015-06-22 | Pass | One sample a little low | Borderline >Q30; Smearing in heatmap but longer reads
150623_M00766_0116_000000000-AFLEW | M00766 | 2015-06-23 | Pass | One sample very low | Borderline >Q30; Smearing in heatmap but longer reads
150706_M00766_0118_000000000-AF512 | M00766 | 2015-07-06 | Pass |  | Borderline >Q30; Smearing in heatmap but longer reads
150729_M02641_0008_000000000-AGET9 | M02641 | 2015-07-29 | Pass
150730_M00766_0120_000000000-AFN4E | M00766 | 2015-07-30 | Pass | One sample very low
150731_M00766_0121_000000000-AFMYC | M00766 | 2015-07-31 | Pass | One sample very high
150731_M02641_0009_000000000-AFN3M | M02641 | 2015-07-31 | Pass | Some samples high
150807_M02641_0012_000000000-AFMYT | M02641 | 2015-08-07 | Borderline Pass | One sample high | Low cluster density
150810_M00766_0123_000000000-AGTUY | M00766 | 2015-08-10 | Pass | One sample very low
150811_M02641_0014_000000000-AGK0F | M02641 | 2015-08-11 | Pass |  | Cluster density may be a little high
150819_M02641_0016_000000000-AGJHU | M02641 | 2015-08-19 | Pass |  | Cluster density may be a little low
150820_M00766_0126_000000000-AGHUU | M00766 | 2015-08-20 | Pass | Some samples very low
150820_M02641_0017_000000000-AGJJA | M02641 | 2015-08-20 | Pass |  | Cluster density may be a little low
150821_M02641_0018_000000000-AGJNU | M02641 | 2015-08-21 | Borderline Pass |  | Low cluster density
150902_M00766_0128_000000000-AFLDG | M00766 | 2015-09-02 | Borderline Fail | Some samples very low | Borderline >Q30; Smearing in heatmap
150904_M02641_0022_000000000-AH991 | M02641 | 2015-09-04 | Borderline Fail | One sample very low | Borderline >Q30; Smearing in heatmap (longer reads)
150910_M02641_0023_000000000-AFLHH | M02641 | 2015-09-10 | Borderline Fail | One sample very high | Cluster density a little high; Smearing in heatmap
150911_M00766_0131_000000000-AGLDT | M00766 | 2015-09-11 | Pass
150911_M02641_0024_AF50J | M02641 | 2015-09-11 | Borderline Pass | One sample very high | Borderline >Q30; Some small smearing in heatmap towards ends of reads
150914_M00766_0132_000000000-AF41F | M00766 | 2015-09-14 | Borderline Pass | One sample very low | Very low cluster density
150925_M02641_0028_000000000-AGK08 | M02641 | 2015-09-25 | Pass
150928_M02641_0029_000000000-AGHV5 | M02641 | 2015-09-28 | Pass
150930_M02641_0031_000000000-AGJGT | M02641 | 2015-09-30 | Pass | One sample very low
151001_M02641_0032_000000000-AGLE0 | M02641 | 2015-10-01 | Borderline Fail |  | High cluster density
151002_M02641_0033_000000000-AFLDA | M02641 | 2015-10-02 | Borderline Pass | One sample very low | Cluster density a little high
151009_M00766_0140_000000000-AGK0B | M00766 | 2015-10-09 | Pass
151009_M02641_0036_000000000-AFN30 | M02641 | 2015-10-09 | Borderline Pass | Some samples high | Borderline >Q30; Slight smearing in heatmap; Cluster density a little high
151012_M02641_0037_000000000-AFNA5 | M02641 | 2015-10-12 | Pass | One sample very high | Borderline >Q30; Slight smearing in heatmap
151019_M02641_0040_000000000-AJDVH | M02641 | 2015-10-19 | Borderline Pass | Some samples very low | High cluster density
151021_M00766_0144_000000000-AG028 | M00766 | 2015-10-21 | Pass
151022_M02641_0042_000000000-AJ5B5 | M02641 | 2015-10-22 | Fail | Some samples a little high | Low >Q30; Smearing in heatmap
151026_M00766_0148_000000000-AFLVV | M00766 | 2015-10-26 | Pass | One sample very high
151102_M02641_0045_000000000-AFN3H | M02641 | 2015-11-02 | Pass
151104_M00766_0151_000000000-AK8EW | M00766 | 2015-11-04 | Pass | Some samples low and high
151106_M00766_0152_000000000-AJ5W5 | M00766 | 2015-11-06 | Pass | One sample low, One sample high
151111_M02641_0049_000000000-AJD47 | M02641 | 2015-11-11 | Pass | One sample low, One sample high | Borderline >Q30; Smearing in heatmap but longer reads
151116_M00766_0154_000000000-AJF4C | M00766 | 2015-11-16 | Pass | One sample low
151119_M02641_0052_000000000-AJJ6B | M02641 | 2015-11-19 | Pass |  | Cluster density may be a little low
151125_M00766_0158_000000000-AJJC1 | M00766 | 2015-11-25 | Borderline Pass | One sample low, One sample high | Very low cluster density
151130_M00766_0163_000000000-AJGC9 | M00766 | 2015-11-30 | Borderline Fail | One sample low, One sample high | High cluster density; Reasonable difference in cluster density passing filter and cluster density
151204_M02641_0056_000000000-AJDC6 | M02641 | 2015-12-04 | Borderline Pass | One sample slightly high | Cluster density a little low; Borderline >Q30; Smearing in heatmap but longer reads
151208_M02641_0057_000000000-AJHRF | M02641 | 2015-12-08 | Borderline Pass |  | Very low cluster density
151210_M00766_0168_000000000-AJDDH | M00766 | 2015-12-10 | Pass | One sample high
151211_M00766_0169_000000000-AK5DW | M00766 | 2015-12-11 | Fail | One sample a little high | Low >Q30; Smearing in heatmap
151218_M00766_0171_000000000-AK5FL | M00766 | 2015-12-18 | Fail | High NTCs | Low >Q30; Smearing in heatmap; Cluster density a little low
160104_M02641_0062_000000000-AL603 | M02641 | 2016-01-04 | Pass
160105_M00766_0174_000000000-AL4LB | M00766 | 2016-01-05 | Pass | Some samples low, One sample high
160107_M02641_0063_000000000-AJDC9 | M02641 | 2016-01-07 | Pass | One sample high
160108_M02641_0064_000000000-ALNH2 | M02641 | 2016-01-08 | Pass | One sample high | Cluster density a little high; Borderline >Q30; Smearing in heatmap but longer reads
160115_M00766_0178_000000000-ALKJV | M00766 | 2016-01-15 | Pass | One sample very low
160118_M02641_0066_000000000-AJETF | M02641 | 2016-01-18 | Borderline Pass | One sample low, One sample high | Very low cluster density
160119_M02641_0067_000000000-ALN4P | M02641 | 2016-01-19 | Pass
160127_M00766_0001_000000000-ALR2L | M00766 | 2016-01-27 | Pass | Some samples low, Some samples high
160127_M02641_0069_000000000-AMDH4 | M02641 | 2016-01-27 | Borderline Fail | One sample very low, Some samples high | Very high cluster density; Reasonable difference in cluster density passing filter and cluster density
160128_M00766_0002_000000000-ALAJ1 | M00766 | 2016-01-28 | Pass | Some samples high
160204_M00766_0004_000000000-AL4H0 | M00766 | 2016-02-04 | Pass | Some samples very low, One sample very high | Cluster density a little low
160204_M02641_0072_000000000-AK6CE | M02641 | 2016-02-04 | Pass | One sample a little low
160205_M02641_0073_000000000-ALY9Y | M02641 | 2016-02-05 | Borderline Fail | Some samples a little high | Cluster density a little high; Reasonable difference in cluster density passing filter and cluster density
160208_M00766_0006_000000000-AMF4G | M00766 | 2016-02-08 | Borderline Fail | One sample low; Some samples high | Cluster density a little high; Reasonable difference in cluster density passing filter and cluster density
160208_M02641_0075_000000000-AMF40 | M02641 | 2016-02-08 | Borderline Fail | Some samples high | Cluster density a little high; Reasonable difference in cluster density passing filter and cluster density
160209_M02641_0076_000000000-AMF44 | M02641 | 2016-02-09 | Pass | One sample high
160210_M02641_0077_000000000-AME81 | M02641 | 2016-02-10 | Borderline Pass | One sample low, Some samples high | Cluster density a little low
160212_M00766_0008_000000000-AL607 | M00766 | 2016-02-12 | Borderline Pass | Some samples low | Cluster density a little high; Borderline >Q30; Smearing in heatmap but longer reads
160216_M02641_0079_000000000-AL75M | M02641 | 2016-02-16 | Pass | One sample very low, Some samples high
160218_M00766_0010_000000000-AMF48 | M00766 | 2016-02-18 | Pass | One sample high
160218_M02641_0080_000000000-AMF3Y | M02641 | 2016-02-18 | Pass | Some samples low, Some samples high | Cluster density a little low
|  | Pass | 73
|  | Borderline Pass | 21
|  | Fail | 15
|  | Borderline Fail | 13
|  | Passes | 94
|  | Fails | 28
|  | Indexing | 73

## FailsPassesAuto
Pass Method 1 | Fail Method 1 | Pass Method 2 | Fail Method 2
0 | 130417_M00766_0007_000000000-A3R1F | 0 | 130417_M00766_0007_000000000-A3R1F | INDEX($A$2:$A$123,SMALL(IF($A$2:$A$123<>0,ROW($A$2:$A$123),ROW(2:2)),2))
130508_M00766_0013_000000000-A3RB5 | 0 | 130508_M00766_0013_000000000-A3RB5 | 0
130510_M00766_0014_000000000-A3PFJ | 0 | 130510_M00766_0014_000000000-A3PFJ | 0
0 | 130613_M00766_0018_000000000-A4FN3 | 0 | 130613_M00766_0018_000000000-A4FN3
130708_M00766_0023_000000000-A20R8 | 0 | 0 | 130708_M00766_0023_000000000-A20R8
130902_M00766_0045_000000000-A5BVD | 0 | 130902_M00766_0045_000000000-A5BVD | 0
130909_M00766_0047_000000000-A5PEH | 0 | 130909_M00766_0047_000000000-A5PEH | 0
0 | 140205_M00766_0021_000000000-A7RVR | 0 | 140205_M00766_0021_000000000-A7RVR
0 | 140603_M00766_0037_000000000-A7WU9 | 0 | 140603_M00766_0037_000000000-A7WU9
0 | 140620_M02641_0004_000000000-A8R7V | 0 | 140620_M02641_0004_000000000-A8R7V
0 | 140625_M00766_0041_000000000-A8P78 | 0 | 140625_M00766_0041_000000000-A8P78
0 | 140715_M02641_0008_000000000-A8P81 | 0 | 140715_M02641_0008_000000000-A8P81
140804_M02641_0011_000000000-AAD6A | 0 | 0 | 140804_M02641_0011_000000000-AAD6A
0 | 140905_M02641_0017_000000000-AA3FN | 0 | 140905_M02641_0017_000000000-AA3FN
0 | 140917_M02641_0018_000000000-AA3H7 | 0 | 140917_M02641_0018_000000000-AA3H7
0 | 141024_M02641_0024_000000000-A8P84 | 0 | 141024_M02641_0024_000000000-A8P84
0 | 141113_M02641_0025_000000000-A8RTY | 0 | 141113_M02641_0025_000000000-A8RTY
0 | 141125_M00766_0062_000000000-A7BKW | 0 | 141125_M00766_0062_000000000-A7BKW
0 | 141222_M00766_0067_000000000-ACCB3 | 0 | 141222_M00766_0067_000000000-ACCB3
0 | 150112_M00766_0068_000000000-A8PD8 | 0 | 150112_M00766_0068_000000000-A8PD8
0 | 150114_M00766_0069_000000000-AA68B | 0 | 150114_M00766_0069_000000000-AA68B
0 | 150116_M00766_0070_000000000-ACC3G | 0 | 150116_M00766_0070_000000000-ACC3G
150127_M00766_0072_000000000-ABRK7 | 0 | 150127_M00766_0072_000000000-ABRK7 | 0
0 | 150130_M00766_0073_000000000-ACBVB | 0 | 150130_M00766_0073_000000000-ACBVB
150205_M02641_0029_000000000-ACCE2 | 0 | 0 | 150205_M02641_0029_000000000-ACCE2
0 | 150220_M00766_0077_000000000-AA2U8 | 0 | 150220_M00766_0077_000000000-AA2U8
150227_M00766_0079_000000000-ACPR9 | 0 | 150227_M00766_0079_000000000-ACPR9 | 0
0 | 150227_M02641_0032_000000000-ACNF3 | 0 | 150227_M02641_0032_000000000-ACNF3
150302_M02641_0033_000000000-ACC9K | 0 | 150302_M02641_0033_000000000-ACC9K | 0
0 | 150306_M00766_0080_000000000-ACNEP | 0 | 150306_M00766_0080_000000000-ACNEP
0 | 150306_M02641_0034_000000000-ACRUF | 0 | 150306_M02641_0034_000000000-ACRUF
150313_M02641_0037_000000000-ACC8R | 0 | 150313_M02641_0037_000000000-ACC8R | 0
0 | 150325_M00766_0084_000000000-AD8KN | 0 | 150325_M00766_0084_000000000-AD8KN
150325_M02641_0038_000000000-AD8KY | 0 | 150325_M02641_0038_000000000-AD8KY | 0
150401_M00766_0086_000000000-ACBU5 | 0 | 150401_M00766_0086_000000000-ACBU5 | 0
150407_M00766_0088_000000000-ACCAT | 0 | 0 | 150407_M00766_0088_000000000-ACCAT
150410_M00766_0089_000000000-ACBU4 | 0 | 150410_M00766_0089_000000000-ACBU4 | 0
150423_M02641_0041_000000000-ACN0P | 0 | 150423_M02641_0041_000000000-ACN0P | 0
0 | 150429_M02641_0043_000000000-AD8KJ | 0 | 150429_M02641_0043_000000000-AD8KJ
150430_M02641_0044_000000000-AD8VC | 0 | 150430_M02641_0044_000000000-AD8VC | 0
150501_M00766_0093_000000000-AF9MH | 0 | 150501_M00766_0093_000000000-AF9MH | 0
0 | 150501_M02641_0045_000000000-AD6UA | 0 | 150501_M02641_0045_000000000-AD6UA
150507_M00766_0095_000000000-AF7N3 | 0 | 150507_M00766_0095_000000000-AF7N3 | 0
150507_M02641_0046_000000000-AF7N5 | 0 | 150507_M02641_0046_000000000-AF7N5 | 0
0 | 150515_M00766_0098_000000000-AF9ND | 0 | 150515_M00766_0098_000000000-AF9ND
0 | 150518_M00766_0100_000000000-AATJ7 | 0 | 150518_M00766_0100_000000000-AATJ7
150521_M00766_0102_000000000-AF9MV | 0 | 150521_M00766_0102_000000000-AF9MV | 0
0 | 150526_M02641_0053_000000000-AFHDV | 0 | 150526_M02641_0053_000000000-AFHDV
150528_M00766_0105_000000000-AF9MJ | 0 | 150528_M00766_0105_000000000-AF9MJ | 0
0 | 150529_M00766_0106_000000000-AEVP5 | 0 | 150529_M00766_0106_000000000-AEVP5
0 | 150605_M02641_0058_000000000-AFF15 | 0 | 150605_M02641_0058_000000000-AFF15
0 | 150610_M02641_0059_000000000-AFN4H | 0 | 150610_M02641_0059_000000000-AFN4H
150612_M00766_0112_000000000-AFMW5 | 0 | 150612_M00766_0112_000000000-AFMW5 | 0
0 | 150612_M02641_0061_000000000-AFMVT | 0 | 150612_M02641_0061_000000000-AFMVT
0 | 150615_M00766_0113_000000000-AFBH3 | 0 | 150615_M00766_0113_000000000-AFBH3
150619_M00766_0114_000000000-AFMYN | 0 | 150619_M00766_0114_000000000-AFMYN | 0
0 | 150622_M00766_0115_000000000-AFN2H | 0 | 150622_M00766_0115_000000000-AFN2H
150623_M00766_0116_000000000-AFLEW | 0 | 0 | 150623_M00766_0116_000000000-AFLEW
0 | 150706_M00766_0118_000000000-AF512 | 0 | 150706_M00766_0118_000000000-AF512
150729_M02641_0008_000000000-AGET9 | 0 | 150729_M02641_0008_000000000-AGET9 | 0
0 | 150730_M00766_0120_000000000-AFN4E | 0 | 150730_M00766_0120_000000000-AFN4E
150731_M00766_0121_000000000-AFMYC | 0 | 150731_M00766_0121_000000000-AFMYC | 0
150731_M02641_0009_000000000-AFN3M | 0 | 150731_M02641_0009_000000000-AFN3M | 0
0 | 150807_M02641_0012_000000000-AFMYT | 0 | 150807_M02641_0012_000000000-AFMYT
150810_M00766_0123_000000000-AGTUY | 0 | 150810_M00766_0123_000000000-AGTUY | 0
150811_M02641_0014_000000000-AGK0F | 0 | 150811_M02641_0014_000000000-AGK0F | 0
0 | 150819_M02641_0016_000000000-AGJHU | 0 | 150819_M02641_0016_000000000-AGJHU
150820_M00766_0126_000000000-AGHUU | 0 | 150820_M00766_0126_000000000-AGHUU | 0
0 | 150820_M02641_0017_000000000-AGJJA | 0 | 150820_M02641_0017_000000000-AGJJA
0 | 150821_M02641_0018_000000000-AGJNU | 0 | 150821_M02641_0018_000000000-AGJNU
150902_M00766_0128_000000000-AFLDG | 0 | 0 | 150902_M00766_0128_000000000-AFLDG
150904_M02641_0022_000000000-AH991 | 0 | 150904_M02641_0022_000000000-AH991 | 0
150910_M02641_0023_000000000-AFLHH | 0 | 150910_M02641_0023_000000000-AFLHH | 0
150911_M00766_0131_000000000-AGLDT | 0 | 150911_M00766_0131_000000000-AGLDT | 0
150911_M02641_0024_AF50J | 0 | 150911_M02641_0024_AF50J | 0
0 | 150914_M00766_0132_000000000-AF41F | 0 | 150914_M00766_0132_000000000-AF41F
0 | 150925_M02641_0028_000000000-AGK08 | 0 | 150925_M02641_0028_000000000-AGK08
150928_M02641_0029_000000000-AGHV5 | 0 | 150928_M02641_0029_000000000-AGHV5 | 0
150930_M02641_0031_000000000-AGJGT | 0 | 150930_M02641_0031_000000000-AGJGT | 0
0 | 151001_M02641_0032_000000000-AGLE0 | 0 | 151001_M02641_0032_000000000-AGLE0
151002_M02641_0033_000000000-AFLDA | 0 | 151002_M02641_0033_000000000-AFLDA | 0
0 | 151009_M00766_0140_000000000-AGK0B | 0 | 151009_M00766_0140_000000000-AGK0B
151009_M02641_0036_000000000-AFN30 | 0 | 151009_M02641_0036_000000000-AFN30 | 0
151012_M02641_0037_000000000-AFNA5 | 0 | 151012_M02641_0037_000000000-AFNA5 | 0
0 | 151019_M02641_0040_000000000-AJDVH | 0 | 151019_M02641_0040_000000000-AJDVH
0 | 151021_M00766_0144_000000000-AG028 | 0 | 151021_M00766_0144_000000000-AG028
0 | 151022_M02641_0042_000000000-AJ5B5 | 0 | 151022_M02641_0042_000000000-AJ5B5
151026_M00766_0148_000000000-AFLVV | 0 | 151026_M00766_0148_000000000-AFLVV | 0
0 | 151102_M02641_0045_000000000-AFN3H | 0 | 151102_M02641_0045_000000000-AFN3H
0 | 151104_M00766_0151_000000000-AK8EW | 0 | 151104_M00766_0151_000000000-AK8EW
151106_M00766_0152_000000000-AJ5W5 | 0 | 151106_M00766_0152_000000000-AJ5W5 | 0
151111_M02641_0049_000000000-AJD47 | 0 | 151111_M02641_0049_000000000-AJD47 | 0
0 | 151116_M00766_0154_000000000-AJF4C | 0 | 151116_M00766_0154_000000000-AJF4C
0 | 151119_M02641_0052_000000000-AJJ6B | 0 | 151119_M02641_0052_000000000-AJJ6B
0 | 151125_M00766_0158_000000000-AJJC1 | 0 | 151125_M00766_0158_000000000-AJJC1
0 | 151130_M00766_0163_000000000-AJGC9 | 0 | 151130_M00766_0163_000000000-AJGC9
0 | 151204_M02641_0056_000000000-AJDC6 | 0 | 151204_M02641_0056_000000000-AJDC6
0 | 151208_M02641_0057_000000000-AJHRF | 0 | 151208_M02641_0057_000000000-AJHRF
0 | 151210_M00766_0168_000000000-AJDDH | 0 | 151210_M00766_0168_000000000-AJDDH
0 | 151211_M00766_0169_000000000-AK5DW | 0 | 151211_M00766_0169_000000000-AK5DW
0 | 151218_M00766_0171_000000000-AK5FL | 0 | 151218_M00766_0171_000000000-AK5FL
160104_M02641_0062_000000000-AL603 | 0 | 160104_M02641_0062_000000000-AL603 | 0
0 | 160105_M00766_0174_000000000-AL4LB | 0 | 160105_M00766_0174_000000000-AL4LB
0 | 160107_M02641_0063_000000000-AJDC9 | 0 | 160107_M02641_0063_000000000-AJDC9
160108_M02641_0064_000000000-ALNH2 | 0 | 160108_M02641_0064_000000000-ALNH2 | 0
0 | 160115_M00766_0178_000000000-ALKJV | 0 | 160115_M00766_0178_000000000-ALKJV
0 | 160118_M02641_0066_000000000-AJETF | 0 | 160118_M02641_0066_000000000-AJETF
0 | 0 | 0 | 0
160127_M00766_0001_000000000-ALR2L | 0 | 160127_M00766_0001_000000000-ALR2L | 0
0 | 160127_M02641_0069_000000000-AMDH4 | 0 | 160127_M02641_0069_000000000-AMDH4
160128_M00766_0002_000000000-ALAJ1 | 0 | 160128_M00766_0002_000000000-ALAJ1 | 0
0 | 160204_M00766_0004_000000000-AL4H0 | 0 | 160204_M00766_0004_000000000-AL4H0
0 | 160204_M02641_0072_000000000-AK6CE | 0 | 160204_M02641_0072_000000000-AK6CE
0 | 160205_M02641_0073_000000000-ALY9Y | 0 | 160205_M02641_0073_000000000-ALY9Y
160208_M00766_0006_000000000-AMF4G | 0 | 160208_M00766_0006_000000000-AMF4G | 0
0 | 160208_M02641_0075_000000000-AMF40 | 0 | 160208_M02641_0075_000000000-AMF40
0 | 160209_M02641_0076_000000000-AMF44 | 0 | 160209_M02641_0076_000000000-AMF44
0 | 160210_M02641_0077_000000000-AME81 | 0 | 160210_M02641_0077_000000000-AME81
160212_M00766_0008_000000000-AL607 | 0 | 160212_M00766_0008_000000000-AL607 | 0
160216_M02641_0079_000000000-AL75M | 0 | 160216_M02641_0079_000000000-AL75M | 0
160218_M00766_0010_000000000-AMF48 | 0 | 160218_M00766_0010_000000000-AMF48 | 0
0 | 160218_M02641_0080_000000000-AMF3Y | 0 | 160218_M02641_0080_000000000-AMF3Y

## FailsPassesManual
Pass | Fail | Borderline Pass | Borderline Fail | Passes | Fails
130417_M00766_0007_000000000-A3R1F | 0 | 0 | 0 | 130417_M00766_0007_000000000-A3R1F | 0
130508_M00766_0013_000000000-A3RB5 | 0 | 0 | 0 | 130508_M00766_0013_000000000-A3RB5 | 0
130510_M00766_0014_000000000-A3PFJ | 0 | 0 | 0 | 130510_M00766_0014_000000000-A3PFJ | 0
0 | 130613_M00766_0018_000000000-A4FN3 | 0 | 0 | 0 | 130613_M00766_0018_000000000-A4FN3
130708_M00766_0023_000000000-A20R8 | 0 | 0 | 0 | 130708_M00766_0023_000000000-A20R8 | 0
130902_M00766_0045_000000000-A5BVD | 0 | 0 | 0 | 130902_M00766_0045_000000000-A5BVD | 0
130909_M00766_0047_000000000-A5PEH | 0 | 0 | 0 | 130909_M00766_0047_000000000-A5PEH | 0
0 | 140205_M00766_0021_000000000-A7RVR | 0 | 0 | 0 | 140205_M00766_0021_000000000-A7RVR
0 | 140603_M00766_0037_000000000-A7WU9 | 0 | 0 | 0 | 140603_M00766_0037_000000000-A7WU9
140620_M02641_0004_000000000-A8R7V | 0 | 0 | 0 | 140620_M02641_0004_000000000-A8R7V | 0
0 | 140625_M00766_0041_000000000-A8P78 | 0 | 0 | 0 | 140625_M00766_0041_000000000-A8P78
0 | 0 | 140715_M02641_0008_000000000-A8P81 | 0 | 140715_M02641_0008_000000000-A8P81 | 0
0 | 140804_M02641_0011_000000000-AAD6A | 0 | 0 | 0 | 140804_M02641_0011_000000000-AAD6A
140905_M02641_0017_000000000-AA3FN | 0 | 0 | 0 | 140905_M02641_0017_000000000-AA3FN | 0
140917_M02641_0018_000000000-AA3H7 | 0 | 0 | 0 | 140917_M02641_0018_000000000-AA3H7 | 0
141024_M02641_0024_000000000-A8P84 | 0 | 0 | 0 | 141024_M02641_0024_000000000-A8P84 | 0
0 | 0 | 141113_M02641_0025_000000000-A8RTY | 0 | 141113_M02641_0025_000000000-A8RTY | 0
141125_M00766_0062_000000000-A7BKW | 0 | 0 | 0 | 141125_M00766_0062_000000000-A7BKW | 0
0 | 141222_M00766_0067_000000000-ACCB3 | 0 | 0 | 0 | 141222_M00766_0067_000000000-ACCB3
150112_M00766_0068_000000000-A8PD8 | 0 | 0 | 0 | 150112_M00766_0068_000000000-A8PD8 | 0
0 | 150114_M00766_0069_000000000-AA68B | 0 | 0 | 0 | 150114_M00766_0069_000000000-AA68B
150116_M00766_0070_000000000-ACC3G | 0 | 0 | 0 | 150116_M00766_0070_000000000-ACC3G | 0
150127_M00766_0072_000000000-ABRK7 | 0 | 0 | 0 | 150127_M00766_0072_000000000-ABRK7 | 0
0 | 150130_M00766_0073_000000000-ACBVB | 0 | 0 | 0 | 150130_M00766_0073_000000000-ACBVB
0 | 0 | 0 | 150205_M02641_0029_000000000-ACCE2 | 0 | 150205_M02641_0029_000000000-ACCE2
0 | 150220_M00766_0077_000000000-AA2U8 | 0 | 0 | 0 | 150220_M00766_0077_000000000-AA2U8
150227_M00766_0079_000000000-ACPR9 | 0 | 0 | 0 | 150227_M00766_0079_000000000-ACPR9 | 0
150227_M02641_0032_000000000-ACNF3 | 0 | 0 | 0 | 150227_M02641_0032_000000000-ACNF3 | 0
0 | 0 | 150302_M02641_0033_000000000-ACC9K | 0 | 150302_M02641_0033_000000000-ACC9K | 0
150306_M00766_0080_000000000-ACNEP | 0 | 0 | 0 | 150306_M00766_0080_000000000-ACNEP | 0
150306_M02641_0034_000000000-ACRUF | 0 | 0 | 0 | 150306_M02641_0034_000000000-ACRUF | 0
150313_M02641_0037_000000000-ACC8R | 0 | 0 | 0 | 150313_M02641_0037_000000000-ACC8R | 0
150325_M00766_0084_000000000-AD8KN | 0 | 0 | 0 | 150325_M00766_0084_000000000-AD8KN | 0
150325_M02641_0038_000000000-AD8KY | 0 | 0 | 0 | 150325_M02641_0038_000000000-AD8KY | 0
0 | 0 | 150401_M00766_0086_000000000-ACBU5 | 0 | 150401_M00766_0086_000000000-ACBU5 | 0
0 | 0 | 0 | 150407_M00766_0088_000000000-ACCAT | 0 | 150407_M00766_0088_000000000-ACCAT
150410_M00766_0089_000000000-ACBU4 | 0 | 0 | 0 | 150410_M00766_0089_000000000-ACBU4 | 0
150423_M02641_0041_000000000-ACN0P | 0 | 0 | 0 | 150423_M02641_0041_000000000-ACN0P | 0
0 | 150429_M02641_0043_000000000-AD8KJ | 0 | 0 | 0 | 150429_M02641_0043_000000000-AD8KJ
150430_M02641_0044_000000000-AD8VC | 0 | 0 | 0 | 150430_M02641_0044_000000000-AD8VC | 0
0 | 0 | 0 | 150501_M00766_0093_000000000-AF9MH | 0 | 150501_M00766_0093_000000000-AF9MH
0 | 0 | 150501_M02641_0045_000000000-AD6UA | 0 | 150501_M02641_0045_000000000-AD6UA | 0
150507_M00766_0095_000000000-AF7N3 | 0 | 0 | 0 | 150507_M00766_0095_000000000-AF7N3 | 0
150507_M02641_0046_000000000-AF7N5 | 0 | 0 | 0 | 150507_M02641_0046_000000000-AF7N5 | 0
0 | 0 | 150515_M00766_0098_000000000-AF9ND | 0 | 150515_M00766_0098_000000000-AF9ND | 0
0 | 0 | 150518_M00766_0100_000000000-AATJ7 | 0 | 150518_M00766_0100_000000000-AATJ7 | 0
150521_M00766_0102_000000000-AF9MV | 0 | 0 | 0 | 150521_M00766_0102_000000000-AF9MV | 0
0 | 150526_M02641_0053_000000000-AFHDV | 0 | 0 | 0 | 150526_M02641_0053_000000000-AFHDV
150528_M00766_0105_000000000-AF9MJ | 0 | 0 | 0 | 150528_M00766_0105_000000000-AF9MJ | 0
150529_M00766_0106_000000000-AEVP5 | 0 | 0 | 0 | 150529_M00766_0106_000000000-AEVP5 | 0
0 | 0 | 150605_M02641_0058_000000000-AFF15 | 0 | 150605_M02641_0058_000000000-AFF15 | 0
150610_M02641_0059_000000000-AFN4H | 0 | 0 | 0 | 150610_M02641_0059_000000000-AFN4H | 0
0 | 0 | 0 | 150612_M00766_0112_000000000-AFMW5 | 0 | 150612_M00766_0112_000000000-AFMW5
150612_M02641_0061_000000000-AFMVT | 0 | 0 | 0 | 150612_M02641_0061_000000000-AFMVT | 0
0 | 150615_M00766_0113_000000000-AFBH3 | 0 | 0 | 0 | 150615_M00766_0113_000000000-AFBH3
150619_M00766_0114_000000000-AFMYN | 0 | 0 | 0 | 150619_M00766_0114_000000000-AFMYN | 0
150622_M00766_0115_000000000-AFN2H | 0 | 0 | 0 | 150622_M00766_0115_000000000-AFN2H | 0
150623_M00766_0116_000000000-AFLEW | 0 | 0 | 0 | 150623_M00766_0116_000000000-AFLEW | 0
150706_M00766_0118_000000000-AF512 | 0 | 0 | 0 | 150706_M00766_0118_000000000-AF512 | 0
150729_M02641_0008_000000000-AGET9 | 0 | 0 | 0 | 150729_M02641_0008_000000000-AGET9 | 0
150730_M00766_0120_000000000-AFN4E | 0 | 0 | 0 | 150730_M00766_0120_000000000-AFN4E | 0
150731_M00766_0121_000000000-AFMYC | 0 | 0 | 0 | 150731_M00766_0121_000000000-AFMYC | 0
150731_M02641_0009_000000000-AFN3M | 0 | 0 | 0 | 150731_M02641_0009_000000000-AFN3M | 0
0 | 0 | 150807_M02641_0012_000000000-AFMYT | 0 | 150807_M02641_0012_000000000-AFMYT | 0
150810_M00766_0123_000000000-AGTUY | 0 | 0 | 0 | 150810_M00766_0123_000000000-AGTUY | 0
150811_M02641_0014_000000000-AGK0F | 0 | 0 | 0 | 150811_M02641_0014_000000000-AGK0F | 0
150819_M02641_0016_000000000-AGJHU | 0 | 0 | 0 | 150819_M02641_0016_000000000-AGJHU | 0
150820_M00766_0126_000000000-AGHUU | 0 | 0 | 0 | 150820_M00766_0126_000000000-AGHUU | 0
150820_M02641_0017_000000000-AGJJA | 0 | 0 | 0 | 150820_M02641_0017_000000000-AGJJA | 0
0 | 0 | 150821_M02641_0018_000000000-AGJNU | 0 | 150821_M02641_0018_000000000-AGJNU | 0
0 | 0 | 0 | 150902_M00766_0128_000000000-AFLDG | 0 | 150902_M00766_0128_000000000-AFLDG
0 | 0 | 0 | 150904_M02641_0022_000000000-AH991 | 0 | 150904_M02641_0022_000000000-AH991
0 | 0 | 0 | 150910_M02641_0023_000000000-AFLHH | 0 | 150910_M02641_0023_000000000-AFLHH
150911_M00766_0131_000000000-AGLDT | 0 | 0 | 0 | 150911_M00766_0131_000000000-AGLDT | 0
0 | 0 | 150911_M02641_0024_AF50J | 0 | 150911_M02641_0024_AF50J | 0
0 | 0 | 150914_M00766_0132_000000000-AF41F | 0 | 150914_M00766_0132_000000000-AF41F | 0
150925_M02641_0028_000000000-AGK08 | 0 | 0 | 0 | 150925_M02641_0028_000000000-AGK08 | 0
150928_M02641_0029_000000000-AGHV5 | 0 | 0 | 0 | 150928_M02641_0029_000000000-AGHV5 | 0
150930_M02641_0031_000000000-AGJGT | 0 | 0 | 0 | 150930_M02641_0031_000000000-AGJGT | 0
0 | 0 | 0 | 151001_M02641_0032_000000000-AGLE0 | 0 | 151001_M02641_0032_000000000-AGLE0
0 | 0 | 151002_M02641_0033_000000000-AFLDA | 0 | 151002_M02641_0033_000000000-AFLDA | 0
151009_M00766_0140_000000000-AGK0B | 0 | 0 | 0 | 151009_M00766_0140_000000000-AGK0B | 0
0 | 0 | 151009_M02641_0036_000000000-AFN30 | 0 | 151009_M02641_0036_000000000-AFN30 | 0
151012_M02641_0037_000000000-AFNA5 | 0 | 0 | 0 | 151012_M02641_0037_000000000-AFNA5 | 0
0 | 0 | 151019_M02641_0040_000000000-AJDVH | 0 | 151019_M02641_0040_000000000-AJDVH | 0
151021_M00766_0144_000000000-AG028 | 0 | 0 | 0 | 151021_M00766_0144_000000000-AG028 | 0
0 | 151022_M02641_0042_000000000-AJ5B5 | 0 | 0 | 0 | 151022_M02641_0042_000000000-AJ5B5
151026_M00766_0148_000000000-AFLVV | 0 | 0 | 0 | 151026_M00766_0148_000000000-AFLVV | 0
151102_M02641_0045_000000000-AFN3H | 0 | 0 | 0 | 151102_M02641_0045_000000000-AFN3H | 0
151104_M00766_0151_000000000-AK8EW | 0 | 0 | 0 | 151104_M00766_0151_000000000-AK8EW | 0
151106_M00766_0152_000000000-AJ5W5 | 0 | 0 | 0 | 151106_M00766_0152_000000000-AJ5W5 | 0
151111_M02641_0049_000000000-AJD47 | 0 | 0 | 0 | 151111_M02641_0049_000000000-AJD47 | 0
151116_M00766_0154_000000000-AJF4C | 0 | 0 | 0 | 151116_M00766_0154_000000000-AJF4C | 0
151119_M02641_0052_000000000-AJJ6B | 0 | 0 | 0 | 151119_M02641_0052_000000000-AJJ6B | 0
0 | 0 | 151125_M00766_0158_000000000-AJJC1 | 0 | 151125_M00766_0158_000000000-AJJC1 | 0
0 | 0 | 0 | 151130_M00766_0163_000000000-AJGC9 | 0 | 151130_M00766_0163_000000000-AJGC9
0 | 0 | 151204_M02641_0056_000000000-AJDC6 | 0 | 151204_M02641_0056_000000000-AJDC6 | 0
0 | 0 | 151208_M02641_0057_000000000-AJHRF | 0 | 151208_M02641_0057_000000000-AJHRF | 0
151210_M00766_0168_000000000-AJDDH | 0 | 0 | 0 | 151210_M00766_0168_000000000-AJDDH | 0
0 | 151211_M00766_0169_000000000-AK5DW | 0 | 0 | 0 | 151211_M00766_0169_000000000-AK5DW
0 | 151218_M00766_0171_000000000-AK5FL | 0 | 0 | 0 | 151218_M00766_0171_000000000-AK5FL
160104_M02641_0062_000000000-AL603 | 0 | 0 | 0 | 160104_M02641_0062_000000000-AL603 | 0
160105_M00766_0174_000000000-AL4LB | 0 | 0 | 0 | 160105_M00766_0174_000000000-AL4LB | 0
160107_M02641_0063_000000000-AJDC9 | 0 | 0 | 0 | 160107_M02641_0063_000000000-AJDC9 | 0
160108_M02641_0064_000000000-ALNH2 | 0 | 0 | 0 | 160108_M02641_0064_000000000-ALNH2 | 0
160115_M00766_0178_000000000-ALKJV | 0 | 0 | 0 | 160115_M00766_0178_000000000-ALKJV | 0
0 | 0 | 160118_M02641_0066_000000000-AJETF | 0 | 160118_M02641_0066_000000000-AJETF | 0
160119_M02641_0067_000000000-ALN4P | 0 | 0 | 0 | 160119_M02641_0067_000000000-ALN4P | 0
160127_M00766_0001_000000000-ALR2L | 0 | 0 | 0 | 160127_M00766_0001_000000000-ALR2L | 0
0 | 0 | 0 | 160127_M02641_0069_000000000-AMDH4 | 0 | 160127_M02641_0069_000000000-AMDH4
160128_M00766_0002_000000000-ALAJ1 | 0 | 0 | 0 | 160128_M00766_0002_000000000-ALAJ1 | 0
160204_M00766_0004_000000000-AL4H0 | 0 | 0 | 0 | 160204_M00766_0004_000000000-AL4H0 | 0
160204_M02641_0072_000000000-AK6CE | 0 | 0 | 0 | 160204_M02641_0072_000000000-AK6CE | 0
0 | 0 | 0 | 160205_M02641_0073_000000000-ALY9Y | 0 | 160205_M02641_0073_000000000-ALY9Y
0 | 0 | 0 | 160208_M00766_0006_000000000-AMF4G | 0 | 160208_M00766_0006_000000000-AMF4G
0 | 0 | 0 | 160208_M02641_0075_000000000-AMF40 | 0 | 160208_M02641_0075_000000000-AMF40
160209_M02641_0076_000000000-AMF44 | 0 | 0 | 0 | 160209_M02641_0076_000000000-AMF44 | 0
0 | 0 | 160210_M02641_0077_000000000-AME81 | 0 | 160210_M02641_0077_000000000-AME81 | 0
0 | 0 | 160212_M00766_0008_000000000-AL607 | 0 | 160212_M00766_0008_000000000-AL607 | 0
160216_M02641_0079_000000000-AL75M | 0 | 0 | 0 | 160216_M02641_0079_000000000-AL75M | 0
160218_M00766_0010_000000000-AMF48 | 0 | 0 | 0 | 160218_M00766_0010_000000000-AMF48 | 0
160218_M02641_0080_000000000-AMF3Y | 0 | 0 | 0 | 160218_M02641_0080_000000000-AMF3Y | 0

## CompareManual1Auto
Pass Method 1 (Automated) | Passes | same/diff | run name for lookup | borderline? | auto wrong (to pass)? | false +ves | false -ves | true +ves | true -ves | Fail Method 1 (Automated) | Fails | same/diff | run name for lookup | borderline? | auto wrong? (to fail) | what test(s) triggered?
130417_M00766_0007_000000000-A3R1F | 130417_M00766_0007_000000000-A3R1F | same |  |  |  | no | no | no | yes | 0 | 0 | same
130508_M00766_0013_000000000-A3RB5 | 130508_M00766_0013_000000000-A3RB5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
130510_M00766_0014_000000000-A3PFJ | 130510_M00766_0014_000000000-A3PFJ | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 130613_M00766_0018_000000000-A4FN3 | 130613_M00766_0018_000000000-A4FN3 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered | Under Q30 threshold triggered
130708_M00766_0023_000000000-A20R8 | 130708_M00766_0023_000000000-A20R8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
130902_M00766_0045_000000000-A5BVD | 130902_M00766_0045_000000000-A5BVD | same |  |  |  | no | no | no | yes | 0 | 0 | same
130909_M00766_0047_000000000-A5PEH | 130909_M00766_0047_000000000-A5PEH | same |  |  |  | no | no | no | yes | 0 | 0 | same
140205_M00766_0021_000000000-A7RVR | 0 | diff | 140205_M00766_0021_000000000-A7RVR | no | yes | no | yes | no | no | 0 | 140205_M00766_0021_000000000-A7RVR | diff | 140205_M00766_0021_000000000-A7RVR | no | no
0 | 0 | same |  |  |  | no | no | yes | no | 140603_M00766_0037_000000000-A7WU9 | 140603_M00766_0037_000000000-A7WU9 | same
140620_M02641_0004_000000000-A8R7V | 140620_M02641_0004_000000000-A8R7V | same |  |  |  | no | no | no | yes | 0 | 0 | same
140625_M00766_0041_000000000-A8P78 | 0 | diff | 140625_M00766_0041_000000000-A8P78 | no | yes | no | yes | no | no | 0 | 140625_M00766_0041_000000000-A8P78 | diff | 140625_M00766_0041_000000000-A8P78 | no | no
140715_M02641_0008_000000000-A8P81 | 140715_M02641_0008_000000000-A8P81 | same |  |  |  | no | no | no | yes | 0 | 0 | same
140804_M02641_0011_000000000-AAD6A | 0 | diff | 140804_M02641_0011_000000000-AAD6A | no | yes | no | yes | no | no | 0 | 140804_M02641_0011_000000000-AAD6A | diff | 140804_M02641_0011_000000000-AAD6A | no | no
0 | 140905_M02641_0017_000000000-AA3FN | diff | 140905_M02641_0017_000000000-AA3FN | no | no | yes | no | no | no | 140905_M02641_0017_000000000-AA3FN | 0 | diff | 140905_M02641_0017_000000000-AA3FN | no | yes | COV cluster density triggered
140917_M02641_0018_000000000-AA3H7 | 140917_M02641_0018_000000000-AA3H7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
141024_M02641_0024_000000000-A8P84 | 141024_M02641_0024_000000000-A8P84 | same |  |  |  | no | no | no | yes | 0 | 0 | same
141113_M02641_0025_000000000-A8RTY | 141113_M02641_0025_000000000-A8RTY | same |  |  |  | no | no | no | yes | 0 | 0 | same
141125_M00766_0062_000000000-A7BKW | 141125_M00766_0062_000000000-A7BKW | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 141222_M00766_0067_000000000-ACCB3 | 141222_M00766_0067_000000000-ACCB3 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
150112_M00766_0068_000000000-A8PD8 | 150112_M00766_0068_000000000-A8PD8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150114_M00766_0069_000000000-AA68B | 150114_M00766_0069_000000000-AA68B | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
150116_M00766_0070_000000000-ACC3G | 150116_M00766_0070_000000000-ACC3G | same |  |  |  | no | no | no | yes | 0 | 0 | same
150127_M00766_0072_000000000-ABRK7 | 150127_M00766_0072_000000000-ABRK7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150130_M00766_0073_000000000-ACBVB | 150130_M00766_0073_000000000-ACBVB | same |  |  |  | Cluster density pf gap triggered | Under Q30 threshold triggered
150205_M02641_0029_000000000-ACCE2 | 0 | diff | 150205_M02641_0029_000000000-ACCE2 | no | yes | no | yes | no | no | 0 | 150205_M02641_0029_000000000-ACCE2 | diff | 150205_M02641_0029_000000000-ACCE2 | yes | no
0 | 0 | same |  |  |  | no | no | yes | no | 150220_M00766_0077_000000000-AA2U8 | 150220_M00766_0077_000000000-AA2U8 | same |  |  |  | % reads passing filter triggered | COV cluster density triggered | Cluster density pf gap triggered | Under Q30 threshold triggered
150227_M00766_0079_000000000-ACPR9 | 150227_M00766_0079_000000000-ACPR9 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150227_M02641_0032_000000000-ACNF3 | diff | 150227_M02641_0032_000000000-ACNF3 | no | no | yes | no | no | no | 150227_M02641_0032_000000000-ACNF3 | 0 | diff | 150227_M02641_0032_000000000-ACNF3 | no | yes | COV cluster density triggered
150302_M02641_0033_000000000-ACC9K | 150302_M02641_0033_000000000-ACC9K | same |  |  |  | no | no | no | yes | 0 | 0 | same
150306_M00766_0080_000000000-ACNEP | 150306_M00766_0080_000000000-ACNEP | same |  |  |  | no | no | no | yes | 0 | 0 | same
150306_M02641_0034_000000000-ACRUF | 150306_M02641_0034_000000000-ACRUF | same |  |  |  | no | no | no | yes | 0 | 0 | same
150313_M02641_0037_000000000-ACC8R | 150313_M02641_0037_000000000-ACC8R | same |  |  |  | no | no | no | yes | 0 | 0 | same
150325_M00766_0084_000000000-AD8KN | 150325_M00766_0084_000000000-AD8KN | same |  |  |  | no | no | no | yes | 0 | 0 | same
150325_M02641_0038_000000000-AD8KY | 150325_M02641_0038_000000000-AD8KY | same |  |  |  | no | no | no | yes | 0 | 0 | same
150401_M00766_0086_000000000-ACBU5 | 150401_M00766_0086_000000000-ACBU5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150407_M00766_0088_000000000-ACCAT | 0 | diff | 150407_M00766_0088_000000000-ACCAT | no | yes | no | yes | no | no | 0 | 150407_M00766_0088_000000000-ACCAT | diff | 150407_M00766_0088_000000000-ACCAT | yes | no
150410_M00766_0089_000000000-ACBU4 | 150410_M00766_0089_000000000-ACBU4 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150423_M02641_0041_000000000-ACN0P | 150423_M02641_0041_000000000-ACN0P | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150429_M02641_0043_000000000-AD8KJ | 150429_M02641_0043_000000000-AD8KJ | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
150430_M02641_0044_000000000-AD8VC | 150430_M02641_0044_000000000-AD8VC | same |  |  |  | no | no | no | yes | 0 | 0 | same
150501_M00766_0093_000000000-AF9MH | 0 | diff | 150501_M00766_0093_000000000-AF9MH | no | yes | no | yes | no | no | 0 | 150501_M00766_0093_000000000-AF9MH | diff | 150501_M00766_0093_000000000-AF9MH | yes | no
0 | 150501_M02641_0045_000000000-AD6UA | diff | 150501_M02641_0045_000000000-AD6UA | yes | no | yes | no | no | no | 150501_M02641_0045_000000000-AD6UA | 0 | diff | 150501_M02641_0045_000000000-AD6UA | no | yes | Cluster density pf gap triggered
150507_M00766_0095_000000000-AF7N3 | 150507_M00766_0095_000000000-AF7N3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150507_M02641_0046_000000000-AF7N5 | 150507_M02641_0046_000000000-AF7N5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150515_M00766_0098_000000000-AF9ND | 150515_M00766_0098_000000000-AF9ND | same |  |  |  | no | no | no | yes | 0 | 0 | same
150518_M00766_0100_000000000-AATJ7 | 150518_M00766_0100_000000000-AATJ7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150521_M00766_0102_000000000-AF9MV | 150521_M00766_0102_000000000-AF9MV | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150526_M02641_0053_000000000-AFHDV | 150526_M02641_0053_000000000-AFHDV | same |  |  |  | % reads passing filter triggered | COV cluster density triggered | Cluster density pf gap triggered | Under Q30 threshold triggered
150528_M00766_0105_000000000-AF9MJ | 150528_M00766_0105_000000000-AF9MJ | same |  |  |  | no | no | no | yes | 0 | 0 | same
150529_M00766_0106_000000000-AEVP5 | 150529_M00766_0106_000000000-AEVP5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150605_M02641_0058_000000000-AFF15 | 150605_M02641_0058_000000000-AFF15 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150610_M02641_0059_000000000-AFN4H | diff | 150610_M02641_0059_000000000-AFN4H | no | no | yes | no | no | no | 150610_M02641_0059_000000000-AFN4H | 0 | diff | 150610_M02641_0059_000000000-AFN4H | no | yes | Cluster density pf gap triggered
150612_M00766_0112_000000000-AFMW5 | 0 | diff | 150612_M00766_0112_000000000-AFMW5 | no | yes | no | yes | no | no | 0 | 150612_M00766_0112_000000000-AFMW5 | diff | 150612_M00766_0112_000000000-AFMW5 | yes | no
0 | 150612_M02641_0061_000000000-AFMVT | diff | 150612_M02641_0061_000000000-AFMVT | no | no | yes | no | no | no | 150612_M02641_0061_000000000-AFMVT | 0 | diff | 150612_M02641_0061_000000000-AFMVT | no | yes | % reads passing filter triggered | Cluster density pf gap triggered
0 | 0 | same |  |  |  | no | no | yes | no | 150615_M00766_0113_000000000-AFBH3 | 150615_M00766_0113_000000000-AFBH3 | same |  |  |  | Under Q30 threshold triggered
150619_M00766_0114_000000000-AFMYN | 150619_M00766_0114_000000000-AFMYN | same |  |  |  | no | no | no | yes | 0 | 0 | same
150622_M00766_0115_000000000-AFN2H | 150622_M00766_0115_000000000-AFN2H | same |  |  |  | no | no | no | yes | 0 | 0 | same
150623_M00766_0116_000000000-AFLEW | 150623_M00766_0116_000000000-AFLEW | same |  |  |  | no | no | no | yes | 0 | 0 | same
150706_M00766_0118_000000000-AF512 | 150706_M00766_0118_000000000-AF512 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150729_M02641_0008_000000000-AGET9 | 150729_M02641_0008_000000000-AGET9 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150730_M00766_0120_000000000-AFN4E | diff | 150730_M00766_0120_000000000-AFN4E | no | no | yes | no | no | no | 150730_M00766_0120_000000000-AFN4E | 0 | diff | 150730_M00766_0120_000000000-AFN4E | no | yes | COV cluster density triggered
150731_M00766_0121_000000000-AFMYC | 150731_M00766_0121_000000000-AFMYC | same |  |  |  | no | no | no | yes | 0 | 0 | same
150731_M02641_0009_000000000-AFN3M | 150731_M02641_0009_000000000-AFN3M | same |  |  |  | no | no | no | yes | 0 | 0 | same
150807_M02641_0012_000000000-AFMYT | 150807_M02641_0012_000000000-AFMYT | same |  |  |  | no | no | no | yes | 0 | 0 | same
150810_M00766_0123_000000000-AGTUY | 150810_M00766_0123_000000000-AGTUY | same |  |  |  | no | no | no | yes | 0 | 0 | same
150811_M02641_0014_000000000-AGK0F | 150811_M02641_0014_000000000-AGK0F | same |  |  |  | no | no | no | yes | 0 | 0 | same
150819_M02641_0016_000000000-AGJHU | 150819_M02641_0016_000000000-AGJHU | same |  |  |  | no | no | no | yes | 0 | 0 | same
150820_M00766_0126_000000000-AGHUU | 150820_M00766_0126_000000000-AGHUU | same |  |  |  | no | no | no | yes | 0 | 0 | same
150820_M02641_0017_000000000-AGJJA | 150820_M02641_0017_000000000-AGJJA | same |  |  |  | no | no | no | yes | 0 | 0 | same
150821_M02641_0018_000000000-AGJNU | 150821_M02641_0018_000000000-AGJNU | same |  |  |  | no | no | no | yes | 0 | 0 | same
150902_M00766_0128_000000000-AFLDG | 0 | diff | 150902_M00766_0128_000000000-AFLDG | no | yes | no | yes | no | no | 0 | 150902_M00766_0128_000000000-AFLDG | diff | 150902_M00766_0128_000000000-AFLDG | yes | no
150904_M02641_0022_000000000-AH991 | 0 | diff | 150904_M02641_0022_000000000-AH991 | no | yes | no | yes | no | no | 0 | 150904_M02641_0022_000000000-AH991 | diff | 150904_M02641_0022_000000000-AH991 | yes | no
150910_M02641_0023_000000000-AFLHH | 0 | diff | 150910_M02641_0023_000000000-AFLHH | no | yes | no | yes | no | no | 0 | 150910_M02641_0023_000000000-AFLHH | diff | 150910_M02641_0023_000000000-AFLHH | yes | no
150911_M00766_0131_000000000-AGLDT | 150911_M00766_0131_000000000-AGLDT | same |  |  |  | no | no | no | yes | 0 | 0 | same
150911_M02641_0024_AF50J | 150911_M02641_0024_AF50J | same |  |  |  | no | no | no | yes | 0 | 0 | same
150914_M00766_0132_000000000-AF41F | 150914_M00766_0132_000000000-AF41F | same |  |  |  | no | no | no | yes | 0 | 0 | same
150925_M02641_0028_000000000-AGK08 | 150925_M02641_0028_000000000-AGK08 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150928_M02641_0029_000000000-AGHV5 | 150928_M02641_0029_000000000-AGHV5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150930_M02641_0031_000000000-AGJGT | 150930_M02641_0031_000000000-AGJGT | same |  |  |  | no | no | no | yes | 0 | 0 | same
151001_M02641_0032_000000000-AGLE0 | 0 | diff | 151001_M02641_0032_000000000-AGLE0 | no | yes | no | yes | no | no | 0 | 151001_M02641_0032_000000000-AGLE0 | diff | 151001_M02641_0032_000000000-AGLE0 | yes | no
151002_M02641_0033_000000000-AFLDA | 151002_M02641_0033_000000000-AFLDA | same |  |  |  | no | no | no | yes | 0 | 0 | same
151009_M00766_0140_000000000-AGK0B | 151009_M00766_0140_000000000-AGK0B | same |  |  |  | no | no | no | yes | 0 | 0 | same
151009_M02641_0036_000000000-AFN30 | 151009_M02641_0036_000000000-AFN30 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151012_M02641_0037_000000000-AFNA5 | 151012_M02641_0037_000000000-AFNA5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151019_M02641_0040_000000000-AJDVH | 151019_M02641_0040_000000000-AJDVH | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 151021_M00766_0144_000000000-AG028 | diff | 151021_M00766_0144_000000000-AG028 | no | no | yes | no | no | no | 151021_M00766_0144_000000000-AG028 | 0 | diff | 151021_M00766_0144_000000000-AG028 | no | yes | Cluster density pf gap triggered
0 | 0 | same |  |  |  | no | no | yes | no | 151022_M02641_0042_000000000-AJ5B5 | 151022_M02641_0042_000000000-AJ5B5 | same |  |  |  | Under Q30 threshold triggered
151026_M00766_0148_000000000-AFLVV | 151026_M00766_0148_000000000-AFLVV | same |  |  |  | no | no | no | yes | 0 | 0 | same
151102_M02641_0045_000000000-AFN3H | 151102_M02641_0045_000000000-AFN3H | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 151104_M00766_0151_000000000-AK8EW | diff | 151104_M00766_0151_000000000-AK8EW | no | no | yes | no | no | no | 151104_M00766_0151_000000000-AK8EW | 0 | diff | 151104_M00766_0151_000000000-AK8EW | no | yes | COV cluster density triggered | Cluster density pf gap triggered
151106_M00766_0152_000000000-AJ5W5 | 151106_M00766_0152_000000000-AJ5W5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151111_M02641_0049_000000000-AJD47 | 151111_M02641_0049_000000000-AJD47 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 151116_M00766_0154_000000000-AJF4C | diff | 151116_M00766_0154_000000000-AJF4C | no | no | yes | no | no | no | 151116_M00766_0154_000000000-AJF4C | 0 | diff | 151116_M00766_0154_000000000-AJF4C | no | yes | COV cluster density triggered
151119_M02641_0052_000000000-AJJ6B | 151119_M02641_0052_000000000-AJJ6B | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 151125_M00766_0158_000000000-AJJC1 | diff | 151125_M00766_0158_000000000-AJJC1 | yes | no | yes | no | no | no | 151125_M00766_0158_000000000-AJJC1 | 0 | diff | 151125_M00766_0158_000000000-AJJC1 | no | yes | % reads passing filter triggered | Cluster density pf gap triggered
0 | 0 | same |  |  |  | no | no | yes | no | 151130_M00766_0163_000000000-AJGC9 | 151130_M00766_0163_000000000-AJGC9 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
151204_M02641_0056_000000000-AJDC6 | 151204_M02641_0056_000000000-AJDC6 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151208_M02641_0057_000000000-AJHRF | 151208_M02641_0057_000000000-AJHRF | same |  |  |  | no | no | no | yes | 0 | 0 | same
151210_M00766_0168_000000000-AJDDH | 151210_M00766_0168_000000000-AJDDH | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 151211_M00766_0169_000000000-AK5DW | 151211_M00766_0169_000000000-AK5DW | same |  |  |  | Under Q30 threshold triggered
0 | 0 | same |  |  |  | no | no | yes | no | 151218_M00766_0171_000000000-AK5FL | 151218_M00766_0171_000000000-AK5FL | same |  |  |  | Under Q30 threshold triggered
160104_M02641_0062_000000000-AL603 | 160104_M02641_0062_000000000-AL603 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160105_M00766_0174_000000000-AL4LB | 160105_M00766_0174_000000000-AL4LB | same |  |  |  | no | no | no | yes | 0 | 0 | same
160107_M02641_0063_000000000-AJDC9 | 160107_M02641_0063_000000000-AJDC9 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160108_M02641_0064_000000000-ALNH2 | 160108_M02641_0064_000000000-ALNH2 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 160115_M00766_0178_000000000-ALKJV | diff | 160115_M00766_0178_000000000-ALKJV | no | no | yes | no | no | no | 160115_M00766_0178_000000000-ALKJV | 0 | diff | 160115_M00766_0178_000000000-ALKJV | no | yes | COV cluster density triggered | Cluster density pf gap triggered
160118_M02641_0066_000000000-AJETF | 160118_M02641_0066_000000000-AJETF | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 160119_M02641_0067_000000000-ALN4P | diff | 160119_M02641_0067_000000000-ALN4P | no | no | yes | no | no | no | 160119_M02641_0067_000000000-ALN4P | 0 | diff | 160119_M02641_0067_000000000-ALN4P | no | yes | COV cluster density triggered
160127_M00766_0001_000000000-ALR2L | 160127_M00766_0001_000000000-ALR2L | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 160127_M02641_0069_000000000-AMDH4 | 160127_M02641_0069_000000000-AMDH4 | same |  |  |  | % reads passing filter triggered | COV cluster density triggered | Cluster density pf gap triggered
160128_M00766_0002_000000000-ALAJ1 | 160128_M00766_0002_000000000-ALAJ1 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160204_M00766_0004_000000000-AL4H0 | 160204_M00766_0004_000000000-AL4H0 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160204_M02641_0072_000000000-AK6CE | 160204_M02641_0072_000000000-AK6CE | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 160205_M02641_0073_000000000-ALY9Y | 160205_M02641_0073_000000000-ALY9Y | same |  |  |  | Cluster density pf gap triggered
160208_M00766_0006_000000000-AMF4G | 0 | diff | 160208_M00766_0006_000000000-AMF4G | no | yes | no | yes | no | no | 0 | 160208_M00766_0006_000000000-AMF4G | diff | 160208_M00766_0006_000000000-AMF4G | yes | no
0 | 0 | same |  |  |  | no | no | yes | no | 160208_M02641_0075_000000000-AMF40 | 160208_M02641_0075_000000000-AMF40 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
160209_M02641_0076_000000000-AMF44 | 160209_M02641_0076_000000000-AMF44 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160210_M02641_0077_000000000-AME81 | 160210_M02641_0077_000000000-AME81 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160212_M00766_0008_000000000-AL607 | 160212_M00766_0008_000000000-AL607 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160216_M02641_0079_000000000-AL75M | 160216_M02641_0079_000000000-AL75M | same |  |  |  | no | no | no | yes | 0 | 0 | same
160218_M00766_0010_000000000-AMF48 | 160218_M00766_0010_000000000-AMF48 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160218_M02641_0080_000000000-AMF3Y | 160218_M02641_0080_000000000-AMF3Y | same |  |  |  | no | no | no | yes | 0 | 0 | same
[6 empty] | 12 | 12 | 16 | 82
| same pass | 98 | [8 empty] | same fail | 98
| diff pass | 24 | [8 empty] | diff fail | 24
| false passes by automated | 12 | [8 empty] | false fails by automated | 12
| missed passes by automated | 12 | [8 empty] | missed fails by automated | 12
0.2449
[10 empty] | concordant | 98
[10 empty] | not concordant | 24
[9 empty] | Failing | false +ves | 12
[10 empty] | false -ves | 12
[10 empty] | true +ves | 16
[10 empty] | true -ves | 82
[10 empty] | total | 122
[12 empty] | 95% CI | 95% CI
[10 empty] | sensitivity (true +ve rate) | 57.1429 | 37.18 | 75.54
[10 empty] | specificity (true -ve rate) | 87.234 | 78.76 | 93.23

## CompareManual2Auto
Pass Method 2 (Automated) | Passes | same/diff | run name for lookup | borderline? | auto wrong (to pass)? | false +ves | false -ves | true +ves | true -ves | Fail Method 2 (Automated) | Fails | same/diff | run name for lookup | borderline? | auto wrong? (to fail) | what test(s) triggered?
130417_M00766_0007_000000000-A3R1F | 130417_M00766_0007_000000000-A3R1F | same |  |  |  | no | no | no | yes | 0 | 0 | same
130508_M00766_0013_000000000-A3RB5 | 130508_M00766_0013_000000000-A3RB5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
130510_M00766_0014_000000000-A3PFJ | 130510_M00766_0014_000000000-A3PFJ | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 130613_M00766_0018_000000000-A4FN3 | 130613_M00766_0018_000000000-A4FN3 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered | Under Q30 threshold triggered | Under Q30 threshold triggered R2
0 | 130708_M00766_0023_000000000-A20R8 | diff | 130708_M00766_0023_000000000-A20R8 | no | no | yes | no | no | no | 130708_M00766_0023_000000000-A20R8 | 0 | diff | 130708_M00766_0023_000000000-A20R8 | no | yes | Under Q30 threshold triggered R2 | Read Slope R2 triggered (new test)
130902_M00766_0045_000000000-A5BVD | 130902_M00766_0045_000000000-A5BVD | same |  |  |  | no | no | no | yes | 0 | 0 | same
130909_M00766_0047_000000000-A5PEH | 130909_M00766_0047_000000000-A5PEH | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 140205_M00766_0021_000000000-A7RVR | 140205_M00766_0021_000000000-A7RVR | same |  |  |  | Under Q30 threshold triggered R2
0 | 0 | same |  |  |  | no | no | yes | no | 140603_M00766_0037_000000000-A7WU9 | 140603_M00766_0037_000000000-A7WU9 | same |  |  |  | Under Q30 threshold triggered R2
140620_M02641_0004_000000000-A8R7V | 140620_M02641_0004_000000000-A8R7V | same |  |  |  | no | no | no | yes | 0 | 0 | same
140625_M00766_0041_000000000-A8P78 | 0 | diff | 140625_M00766_0041_000000000-A8P78 | no | yes | no | yes | no | no | 0 | 140625_M00766_0041_000000000-A8P78 | diff | 140625_M00766_0041_000000000-A8P78 | no | no
140715_M02641_0008_000000000-A8P81 | 140715_M02641_0008_000000000-A8P81 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 140804_M02641_0011_000000000-AAD6A | 140804_M02641_0011_000000000-AAD6A | same |  |  |  | Under Q30 threshold triggered R2
0 | 140905_M02641_0017_000000000-AA3FN | diff | 140905_M02641_0017_000000000-AA3FN | no | no | yes | no | no | no | 140905_M02641_0017_000000000-AA3FN | 0 | diff | 140905_M02641_0017_000000000-AA3FN | no | yes | COV cluster density triggered
140917_M02641_0018_000000000-AA3H7 | 140917_M02641_0018_000000000-AA3H7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
141024_M02641_0024_000000000-A8P84 | 141024_M02641_0024_000000000-A8P84 | same |  |  |  | no | no | no | yes | 0 | 0 | same
141113_M02641_0025_000000000-A8RTY | 141113_M02641_0025_000000000-A8RTY | same |  |  |  | no | no | no | yes | 0 | 0 | same
141125_M00766_0062_000000000-A7BKW | 141125_M00766_0062_000000000-A7BKW | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 141222_M00766_0067_000000000-ACCB3 | 141222_M00766_0067_000000000-ACCB3 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered | Under Q30 threshold triggered R2
150112_M00766_0068_000000000-A8PD8 | 150112_M00766_0068_000000000-A8PD8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150114_M00766_0069_000000000-AA68B | 150114_M00766_0069_000000000-AA68B | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
150116_M00766_0070_000000000-ACC3G | 150116_M00766_0070_000000000-ACC3G | same |  |  |  | no | no | no | yes | 0 | 0 | same
150127_M00766_0072_000000000-ABRK7 | 150127_M00766_0072_000000000-ABRK7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150130_M00766_0073_000000000-ACBVB | 150130_M00766_0073_000000000-ACBVB | same |  |  |  | Cluster density pf gap triggered | Under Q30 threshold triggered | Under Q30 threshold triggered R1 | Under Q30 threshold triggered R2
0 | 0 | same |  |  |  | no | no | yes | no | 150205_M02641_0029_000000000-ACCE2 | 150205_M02641_0029_000000000-ACCE2 | same |  |  |  | Under Q30 threshold triggered R2
0 | 0 | same |  |  |  | no | no | yes | no | 150220_M00766_0077_000000000-AA2U8 | 150220_M00766_0077_000000000-AA2U8 | same |  |  |  | % reads passing filter triggered | COV cluster density triggered | Cluster density pf gap triggered | Under Q30 threshold triggered | Under Q30 threshold triggered R1 | Under Q30 threshold triggered R2
150227_M00766_0079_000000000-ACPR9 | 150227_M00766_0079_000000000-ACPR9 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150227_M02641_0032_000000000-ACNF3 | diff | 150227_M02641_0032_000000000-ACNF3 | no | no | yes | no | no | no | 150227_M02641_0032_000000000-ACNF3 | 0 | diff | 150227_M02641_0032_000000000-ACNF3 | no | yes | COV cluster density triggered
150302_M02641_0033_000000000-ACC9K | 150302_M02641_0033_000000000-ACC9K | same |  |  |  | no | no | no | yes | 0 | 0 | same
150306_M00766_0080_000000000-ACNEP | 150306_M00766_0080_000000000-ACNEP | same |  |  |  | no | no | no | yes | 0 | 0 | same
150306_M02641_0034_000000000-ACRUF | 150306_M02641_0034_000000000-ACRUF | same |  |  |  | no | no | no | yes | 0 | 0 | same
150313_M02641_0037_000000000-ACC8R | 150313_M02641_0037_000000000-ACC8R | same |  |  |  | no | no | no | yes | 0 | 0 | same
150325_M00766_0084_000000000-AD8KN | 150325_M00766_0084_000000000-AD8KN | same |  |  |  | no | no | no | yes | 0 | 0 | same
150325_M02641_0038_000000000-AD8KY | 150325_M02641_0038_000000000-AD8KY | same |  |  |  | no | no | no | yes | 0 | 0 | same
150401_M00766_0086_000000000-ACBU5 | 150401_M00766_0086_000000000-ACBU5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150407_M00766_0088_000000000-ACCAT | 150407_M00766_0088_000000000-ACCAT | same |  |  |  | Under Q30 threshold triggered R2
150410_M00766_0089_000000000-ACBU4 | 150410_M00766_0089_000000000-ACBU4 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150423_M02641_0041_000000000-ACN0P | 150423_M02641_0041_000000000-ACN0P | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150429_M02641_0043_000000000-AD8KJ | 150429_M02641_0043_000000000-AD8KJ | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
150430_M02641_0044_000000000-AD8VC | 150430_M02641_0044_000000000-AD8VC | same |  |  |  | no | no | no | yes | 0 | 0 | same
150501_M00766_0093_000000000-AF9MH | 0 | diff | 150501_M00766_0093_000000000-AF9MH | yes | yes | no | yes | no | no | 0 | 150501_M00766_0093_000000000-AF9MH | diff | 150501_M00766_0093_000000000-AF9MH | yes | no
0 | 150501_M02641_0045_000000000-AD6UA | diff | 150501_M02641_0045_000000000-AD6UA | no | no | yes | no | no | no | 150501_M02641_0045_000000000-AD6UA | 0 | diff | 150501_M02641_0045_000000000-AD6UA | no | yes | Cluster density pf gap triggered
150507_M00766_0095_000000000-AF7N3 | 150507_M00766_0095_000000000-AF7N3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150507_M02641_0046_000000000-AF7N5 | 150507_M02641_0046_000000000-AF7N5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150515_M00766_0098_000000000-AF9ND | 150515_M00766_0098_000000000-AF9ND | same |  |  |  | no | no | no | yes | 0 | 0 | same
150518_M00766_0100_000000000-AATJ7 | 150518_M00766_0100_000000000-AATJ7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150521_M00766_0102_000000000-AF9MV | 150521_M00766_0102_000000000-AF9MV | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150526_M02641_0053_000000000-AFHDV | 150526_M02641_0053_000000000-AFHDV | same |  |  |  | % reads passing filter triggered | COV cluster density triggered | Cluster density pf gap triggered | Under Q30 threshold triggered | Under Q30 threshold triggered R1 | Under Q30 threshold triggered R2
150528_M00766_0105_000000000-AF9MJ | 150528_M00766_0105_000000000-AF9MJ | same |  |  |  | no | no | no | yes | 0 | 0 | same
150529_M00766_0106_000000000-AEVP5 | 150529_M00766_0106_000000000-AEVP5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150605_M02641_0058_000000000-AFF15 | 150605_M02641_0058_000000000-AFF15 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150610_M02641_0059_000000000-AFN4H | diff | 150610_M02641_0059_000000000-AFN4H | no | no | yes | no | no | no | 150610_M02641_0059_000000000-AFN4H | 0 | diff | 150610_M02641_0059_000000000-AFN4H | no | yes | Cluster density pf gap triggered
150612_M00766_0112_000000000-AFMW5 | 0 | diff | 150612_M00766_0112_000000000-AFMW5 | yes | yes | no | yes | no | no | 0 | 150612_M00766_0112_000000000-AFMW5 | diff | 150612_M00766_0112_000000000-AFMW5 | yes | no
0 | 150612_M02641_0061_000000000-AFMVT | diff | 150612_M02641_0061_000000000-AFMVT | no | no | yes | no | no | no | 150612_M02641_0061_000000000-AFMVT | 0 | diff | 150612_M02641_0061_000000000-AFMVT | no | yes | % reads passing filter triggered | Cluster density pf gap triggered
0 | 0 | same |  |  |  | no | no | yes | no | 150615_M00766_0113_000000000-AFBH3 | 150615_M00766_0113_000000000-AFBH3 | same |  |  |  | Under Q30 threshold triggered | Under Q30 threshold triggered R2
150619_M00766_0114_000000000-AFMYN | 150619_M00766_0114_000000000-AFMYN | same |  |  |  | no | no | no | yes | 0 | 0 | same
150622_M00766_0115_000000000-AFN2H | 150622_M00766_0115_000000000-AFN2H | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150623_M00766_0116_000000000-AFLEW | diff | 150623_M00766_0116_000000000-AFLEW | no | no | yes | no | no | no | 150623_M00766_0116_000000000-AFLEW | 0 | diff | 150623_M00766_0116_000000000-AFLEW | no | yes | Under Q30 threshold triggered R2
150706_M00766_0118_000000000-AF512 | 150706_M00766_0118_000000000-AF512 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150729_M02641_0008_000000000-AGET9 | 150729_M02641_0008_000000000-AGET9 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150730_M00766_0120_000000000-AFN4E | diff | 150730_M00766_0120_000000000-AFN4E | no | no | yes | no | no | no | 150730_M00766_0120_000000000-AFN4E | 0 | diff | 150730_M00766_0120_000000000-AFN4E | no | yes | COV cluster density triggered
150731_M00766_0121_000000000-AFMYC | 150731_M00766_0121_000000000-AFMYC | same |  |  |  | no | no | no | yes | 0 | 0 | same
150731_M02641_0009_000000000-AFN3M | 150731_M02641_0009_000000000-AFN3M | same |  |  |  | no | no | no | yes | 0 | 0 | same
150807_M02641_0012_000000000-AFMYT | 150807_M02641_0012_000000000-AFMYT | same |  |  |  | no | no | no | yes | 0 | 0 | same
150810_M00766_0123_000000000-AGTUY | 150810_M00766_0123_000000000-AGTUY | same |  |  |  | no | no | no | yes | 0 | 0 | same
150811_M02641_0014_000000000-AGK0F | 150811_M02641_0014_000000000-AGK0F | same |  |  |  | no | no | no | yes | 0 | 0 | same
150819_M02641_0016_000000000-AGJHU | 150819_M02641_0016_000000000-AGJHU | same |  |  |  | no | no | no | yes | 0 | 0 | same
150820_M00766_0126_000000000-AGHUU | 150820_M00766_0126_000000000-AGHUU | same |  |  |  | no | no | no | yes | 0 | 0 | same
150820_M02641_0017_000000000-AGJJA | 150820_M02641_0017_000000000-AGJJA | same |  |  |  | no | no | no | yes | 0 | 0 | same
150821_M02641_0018_000000000-AGJNU | 150821_M02641_0018_000000000-AGJNU | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150902_M00766_0128_000000000-AFLDG | 150902_M00766_0128_000000000-AFLDG | same |  |  |  | Under Q30 threshold triggered R2
150904_M02641_0022_000000000-AH991 | 0 | diff | 150904_M02641_0022_000000000-AH991 | yes | yes | no | yes | no | no | 0 | 150904_M02641_0022_000000000-AH991 | diff | 150904_M02641_0022_000000000-AH991 | yes | no
150910_M02641_0023_000000000-AFLHH | 0 | diff | 150910_M02641_0023_000000000-AFLHH | yes | yes | no | yes | no | no | 0 | 150910_M02641_0023_000000000-AFLHH | diff | 150910_M02641_0023_000000000-AFLHH | yes | no
150911_M00766_0131_000000000-AGLDT | 150911_M00766_0131_000000000-AGLDT | same |  |  |  | no | no | no | yes | 0 | 0 | same
150911_M02641_0024_AF50J | 150911_M02641_0024_AF50J | same |  |  |  | no | no | no | yes | 0 | 0 | same
150914_M00766_0132_000000000-AF41F | 150914_M00766_0132_000000000-AF41F | same |  |  |  | no | no | no | yes | 0 | 0 | same
150925_M02641_0028_000000000-AGK08 | 150925_M02641_0028_000000000-AGK08 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150928_M02641_0029_000000000-AGHV5 | 150928_M02641_0029_000000000-AGHV5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150930_M02641_0031_000000000-AGJGT | 150930_M02641_0031_000000000-AGJGT | same |  |  |  | no | no | no | yes | 0 | 0 | same
151001_M02641_0032_000000000-AGLE0 | 0 | diff | 151001_M02641_0032_000000000-AGLE0 | yes | yes | no | yes | no | no | 0 | 151001_M02641_0032_000000000-AGLE0 | diff | 151001_M02641_0032_000000000-AGLE0 | yes | no
151002_M02641_0033_000000000-AFLDA | 151002_M02641_0033_000000000-AFLDA | same |  |  |  | no | no | no | yes | 0 | 0 | same
151009_M00766_0140_000000000-AGK0B | 151009_M00766_0140_000000000-AGK0B | same |  |  |  | no | no | no | yes | 0 | 0 | same
151009_M02641_0036_000000000-AFN30 | 151009_M02641_0036_000000000-AFN30 | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Read Slope R2 triggered (new test)
151012_M02641_0037_000000000-AFNA5 | 151012_M02641_0037_000000000-AFNA5 | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Read Slope R2 triggered (new test)
151019_M02641_0040_000000000-AJDVH | 151019_M02641_0040_000000000-AJDVH | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 151021_M00766_0144_000000000-AG028 | diff | 151021_M00766_0144_000000000-AG028 | no | no | yes | no | no | no | 151021_M00766_0144_000000000-AG028 | 0 | diff | 151021_M00766_0144_000000000-AG028 | no | yes | Cluster density pf gap triggered
0 | 0 | same |  |  |  | no | no | yes | no | 151022_M02641_0042_000000000-AJ5B5 | 151022_M02641_0042_000000000-AJ5B5 | same |  |  |  | Under Q30 threshold triggered | Under Q30 threshold triggered R2
151026_M00766_0148_000000000-AFLVV | 151026_M00766_0148_000000000-AFLVV | same |  |  |  | no | no | no | yes | 0 | 0 | same
151102_M02641_0045_000000000-AFN3H | 151102_M02641_0045_000000000-AFN3H | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 151104_M00766_0151_000000000-AK8EW | diff | 151104_M00766_0151_000000000-AK8EW | no | no | yes | no | no | no | 151104_M00766_0151_000000000-AK8EW | 0 | diff | 151104_M00766_0151_000000000-AK8EW | no | yes | COV cluster density triggered | Cluster density pf gap triggered
151106_M00766_0152_000000000-AJ5W5 | 151106_M00766_0152_000000000-AJ5W5 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151111_M02641_0049_000000000-AJD47 | 151111_M02641_0049_000000000-AJD47 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 151116_M00766_0154_000000000-AJF4C | diff | 151116_M00766_0154_000000000-AJF4C | no | no | yes | no | no | no | 151116_M00766_0154_000000000-AJF4C | 0 | diff | 151116_M00766_0154_000000000-AJF4C | no | yes | COV cluster density triggered
151119_M02641_0052_000000000-AJJ6B | 151119_M02641_0052_000000000-AJJ6B | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 151125_M00766_0158_000000000-AJJC1 | diff | 151125_M00766_0158_000000000-AJJC1 | no | no | yes | no | no | no | 151125_M00766_0158_000000000-AJJC1 | 0 | diff | 151125_M00766_0158_000000000-AJJC1 | no | yes | % reads passing filter triggered | Cluster density pf gap triggered
0 | 0 | same |  |  |  | no | no | yes | no | 151130_M00766_0163_000000000-AJGC9 | 151130_M00766_0163_000000000-AJGC9 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
151204_M02641_0056_000000000-AJDC6 | 151204_M02641_0056_000000000-AJDC6 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151208_M02641_0057_000000000-AJHRF | 151208_M02641_0057_000000000-AJHRF | same |  |  |  | no | no | no | yes | 0 | 0 | same
151210_M00766_0168_000000000-AJDDH | 151210_M00766_0168_000000000-AJDDH | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 151211_M00766_0169_000000000-AK5DW | 151211_M00766_0169_000000000-AK5DW | same |  |  |  | Under Q30 threshold triggered | Under Q30 threshold triggered R1 | Under Q30 threshold triggered R2
0 | 0 | same |  |  |  | no | no | yes | no | 151218_M00766_0171_000000000-AK5FL | 151218_M00766_0171_000000000-AK5FL | same |  |  |  | Under Q30 threshold triggered | Under Q30 threshold triggered R1 | Under Q30 threshold triggered R2
160104_M02641_0062_000000000-AL603 | 160104_M02641_0062_000000000-AL603 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160105_M00766_0174_000000000-AL4LB | 160105_M00766_0174_000000000-AL4LB | same |  |  |  | no | no | no | yes | 0 | 0 | same
160107_M02641_0063_000000000-AJDC9 | 160107_M02641_0063_000000000-AJDC9 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160108_M02641_0064_000000000-ALNH2 | 160108_M02641_0064_000000000-ALNH2 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 160115_M00766_0178_000000000-ALKJV | diff | 160115_M00766_0178_000000000-ALKJV | no | no | yes | no | no | no | 160115_M00766_0178_000000000-ALKJV | 0 | diff | 160115_M00766_0178_000000000-ALKJV | no | yes | COV cluster density triggered | Cluster density pf gap triggered
160118_M02641_0066_000000000-AJETF | 160118_M02641_0066_000000000-AJETF | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 160119_M02641_0067_000000000-ALN4P | diff | 160119_M02641_0067_000000000-ALN4P | no | no | yes | no | no | no | 160119_M02641_0067_000000000-ALN4P | 0 | diff | 160119_M02641_0067_000000000-ALN4P | no | yes | COV cluster density triggered | Under Q30 threshold triggered R2
160127_M00766_0001_000000000-ALR2L | 160127_M00766_0001_000000000-ALR2L | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 160127_M02641_0069_000000000-AMDH4 | 160127_M02641_0069_000000000-AMDH4 | same |  |  |  | % reads passing filter triggered | COV cluster density triggered | Cluster density pf gap triggered
160128_M00766_0002_000000000-ALAJ1 | 160128_M00766_0002_000000000-ALAJ1 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160204_M00766_0004_000000000-AL4H0 | 160204_M00766_0004_000000000-AL4H0 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160204_M02641_0072_000000000-AK6CE | 160204_M02641_0072_000000000-AK6CE | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 160205_M02641_0073_000000000-ALY9Y | 160205_M02641_0073_000000000-ALY9Y | same |  |  |  | Cluster density pf gap triggered
160208_M00766_0006_000000000-AMF4G | 0 | diff | 160208_M00766_0006_000000000-AMF4G | yes | yes | no | yes | no | no | 0 | 160208_M00766_0006_000000000-AMF4G | diff | 160208_M00766_0006_000000000-AMF4G | yes | no
0 | 0 | same |  |  |  | no | no | yes | no | 160208_M02641_0075_000000000-AMF40 | 160208_M02641_0075_000000000-AMF40 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
160209_M02641_0076_000000000-AMF44 | 160209_M02641_0076_000000000-AMF44 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160210_M02641_0077_000000000-AME81 | 160210_M02641_0077_000000000-AME81 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160212_M00766_0008_000000000-AL607 | 160212_M00766_0008_000000000-AL607 | same |  |  |  | no | no | no | yes | 0 | 0 | same | Read Slope R2 triggered (new test)
160216_M02641_0079_000000000-AL75M | 160216_M02641_0079_000000000-AL75M | same |  |  |  | no | no | no | yes | 0 | 0 | same
160218_M00766_0010_000000000-AMF48 | 160218_M00766_0010_000000000-AMF48 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160218_M02641_0080_000000000-AMF3Y | 160218_M02641_0080_000000000-AMF3Y | same |  |  |  | no | no | no | yes | 0 | 0 | same
[6 empty] | 14 | 7 | 21 | 80
| same pass | 101 | [8 empty] | same fail | 101
| diff pass | 21 | [8 empty] | diff fail | 21
| false passes by automated | 7 | [8 empty] | false fails by automated | 14
| missed passes by automated | 14 | [8 empty] | missed fails by automated | 7
[10 empty] | concordant | 101
[10 empty] | not concordant | 21
[9 empty] | Failing | false +ves | 14
[10 empty] | false -ves | 7
[10 empty] | true +ves | 21
[10 empty] | true -ves | 80
[10 empty] | total | 122
[12 empty] | 95% CI | 95% CI
[10 empty] | sensitivity (true +ve rate) | 75 | 55.13 | 89.31
[10 empty] | specificity (true -ve rate) | 85.1064 | 76.28 | 91.61

## COVClusterDensity
RunID | InstrumentID | RunDate | NumberOfCyclesR1 | NumberofCyclesR2 | %ReadsPassingFilter | %ReadsPassingFilterOutsideThreshold? | NumberofReadsPassingFilter | ReadCountWithinThreshold? | MeanClusterDensityPassingFilter | MeanClusterDensityWithinThreshold | COVClusterDensity | COVClusterDensityExceedsThreshold? | %MedianClusterDensityPassingFilter | GapBetweenClusterDensitiesPassingFilterAndNotExceedsThreshold? | ProportionOverQ30 | IlluminaQ30Threshold | ProportionOverQ30ExceedsThreshold? | ProportionOverQ30Read1 | ProportionOverQ30ExceedsThresholdRead1 | ProportionOverQ30Read2 | ProportionOverQ30ExceedsThresholdRead2 | K-SPValue | K-SPValueSignificant? | K-SUnpackPvalue | M-WUnpackPvalue | NumberOfCyclesExceedingThresholdUnderQ30Read1 | SlopeOfLineOfBestFitRead1 | SlopeOfLineOfBestFitRead1ExceedsThreshold? | NumberOfCyclesExceedingThresholdUnderQ30Read2 | SlopeOfLineOfBestFitRead2 | SlopeOfLineOfBestFitRead2ExceedsThreshold? | Fail? | Fail2? | Fail? No CD | Fail2? No CD | R1vsR2PropQ30 | R1vsR2K-S | NumCycleR1Threshold | NumCycleR2Threshold | R1vsR2CyclesQ30 | R1vsR2bestfit | Pass/Fail
130417_M00766_0007_000000000-A3R1F | M00766 | 2013-04-17 | 151 | 151 | 93.6874 | no | 14918743 | low | 773.0359 | very low | 0.024644 | no | 93.6673 | no | 0.97122 | 0.8 | no | 0.98017 | no | 0.96226 | no | 0.84095 | no | 0.0022684 | 4.2948e-06 | 0 | -0.00013884 | no | 0 | -0.00025711 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
130508_M00766_0013_000000000-A3RB5 | M00766 | 2013-05-08 | 151 | 151 | 89.695 | no | 17060727 | low | 897.222 | low | 0.013058 | no | 89.6111 | no | 0.92341 | 0.8 | no | 0.94982 | no | 0.89785 | no | 0.67794 | no | 1.7122e-13 | 9.1751e-65 | 0 | -0.00060464 | yes | 0 | -0.00096571 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
130510_M00766_0014_000000000-A3PFJ | M00766 | 2013-05-10 | 156 | 151 | 88.5142 | no | 18378480 | low | 995.2515 | low | 0.013484 | no | 88.6611 | no | 0.95249 | 0.75 | no | 0.96867 | no | 0.93502 | no | 0.95413 | no | 3.3216e-08 | nan | 0 | -0.0002029 | no | 0 | -0.00047138 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
130613_M00766_0018_000000000-A4FN3 | M00766 | 2013-06-13 | 251 | 251 | 62.2349 | yes | 14759376 | low | 877.4421 | low | 0.0052548 | no | 61.6486 | yes | 0.68299 | 0.75 | yes | 0.7656 | no | 0.58473 | yes | 0.67794 | no | 5.4172e-57 | 1.5425e-86 | 28 | -0.0020629 | yes | 55 | -0.0033012 | yes | fail | fail | fail | fail | diff | pass | exceeded | exceeded | same | same | Fail
130708_M00766_0023_000000000-A20R8 | M00766 | 2013-07-08 | 151 | 151 | 90.9082 | no | 8144995 | None | 872.0913 | None | 0.015455 | no | 90.7757 | no | 0.86589 | 0.85 | no | 0.91799 | no | 0.81047 | yes | 0.67794 | no | 9.6929e-19 | 6.9078e-114 | 0 | -0.0013571 | yes | 0 | -0.0030958 | yes | fail | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same | Pass
130902_M00766_0045_000000000-A5BVD | M00766 | 2013-09-02 | 156 | 151 | 92.4439 | no | 17107054 | low | 889.9218 | low | 0.020943 | no | 92.3467 | no | 0.95012 | 0.75 | no | 0.96778 | no | 0.93094 | no | 0.84095 | no | 1.4736e-06 | nan | 0 | -0.00028935 | no | 1 | -0.00072347 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
130909_M00766_0047_000000000-A5PEH | M00766 | 2013-09-09 | 151 | 151 | 91.8202 | no | 17818894 | low | 943.064 | low | 0.013531 | no | 91.9043 | no | 0.91659 | 0.8 | no | 0.93844 | no | 0.89323 | no | 0.95413 | no | 2.3926e-11 | 3.6095e-18 | 0 | -0.00083727 | yes | 0 | -0.00095948 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
140205_M00766_0021_000000000-A7RVR | M00766 | 2014-02-05 | 251 | 251 | 96.2303 | no | 10640092 | low | 549.8594 | very low | 0.015526 | no | 96.2427 | no | 0.77554 | 0.75 | no | 0.82459 | no | 0.7232 | yes | 0.84095 | no | 3.5805e-28 | 1.8115e-29 | 1 | -0.0023111 | yes | 65 | -0.0033538 | yes | fail | fail | pass | fail | diff | pass | not exceeded | exceeded | diff | same | Fail
140603_M00766_0037_000000000-A7WU9 | M00766 | 2014-06-03 | 251 | 251 | 80.6392 | no | 20279789 | low | 1121.2903 | OK | 0.020185 | no | 81.2363 | yes | 0.67262 | 0.75 | yes | 0.87096 | no | 0.46478 | yes | 0.24079 | no | 0 | 0 | 0 | -0.0011435 | yes | 73 | -0.0033779 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same | Fail
140620_M02641_0004_000000000-A8R7V | M02641 | 2014-06-20 | 151 | 151 | 95.2324 | no | 15895393 | low | 815.6811 | very low | 0.010793 | no | 95.1913 | no | 0.92041 | 0.8 | no | 0.96394 | no | 0.87468 | no | 0.35842 | no | 5.5248e-61 | nan | 0 | -0.0004185 | no | 0 | -0.00070599 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
140625_M00766_0041_000000000-A8P78 | M00766 | 2014-06-25 | 251 | 251 | 93.6991 | no | 12306226 | low | 638.3155 | very low | 0.015947 | no | 93.5516 | no | 0.80794 | 0.75 | no | 0.83721 | no | 0.77401 | no | 0.95413 | no | 1.446e-13 | nan | 15 | -0.0023266 | yes | 34 | -0.0029216 | yes | fail | fail | pass | pass | same | pass | not exceeded | exceeded | diff | same | Fail
140715_M02641_0008_000000000-A8P81 | M02641 | 2014-07-15 | 151 | 151 | 98.3026 | no | 5446264 | low | 277.542 | very low | 0.020978 | no | 98.4264 | no | 0.94495 | 0.8 | no | 0.95412 | no | 0.93542 | no | 0.99584 | no | 0.593 | 0.077423 | 0 | -0.00077434 | yes | 0 | -0.00075567 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
140804_M02641_0011_000000000-AAD6A | M02641 | 2014-08-04 | 251 | 251 | 92.8278 | no | 18236023 | low | 947.4631 | low | 0.014242 | no | 92.5473 | no | 0.75551 | 0.75 | no | 0.79909 | no | 0.70633 | yes | 0.84095 | no | 4.579e-31 | nan | 19 | -0.0025558 | yes | 41 | -0.0033381 | yes | fail | fail | pass | fail | diff | pass | not exceeded | exceeded | diff | same | Fail
140905_M02641_0017_000000000-AA3FN | M02641 | 2014-09-05 | 75 | 75 | 90.0019 | no | 23949422 | low | 1015.424 | very low | 0.055448 | yes | 89.6438 | no | 0.95625 | 0.85 | no | 0.96176 | no | 0.95019 | no | 0.99584 | no | 0.99831 | nan | 0 | -0.00042771 | no | 0 | -0.00021245 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
140917_M02641_0018_000000000-AA3H7 | M02641 | 2014-09-17 | 75 | 75 | 91.5859 | no | 24528146 | low | 1028.5468 | very low | 0.035153 | no | 91.5237 | no | 0.96229 | 0.85 | no | 0.96999 | no | 0.95452 | no | 0.84095 | no | 0.88947 | nan | 0 | -0.00031376 | no | 0 | -0.00016688 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
141024_M02641_0024_000000000-A8P84 | M02641 | 2014-10-24 | 151 | 151 | 94.9748 | no | 15281418 | low | 794.7029 | very low | 0.031108 | no | 95.0899 | no | 0.94073 | 0.8 | no | 0.95744 | no | 0.92359 | no | 0.84095 | no | 1.9766e-08 | 4.7689e-23 | 0 | -0.00055662 | yes | 0 | -0.00062808 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
141113_M02641_0025_000000000-A8RTY | M02641 | 2014-11-13 | 151 | 151 | 97.3253 | no | 9834883 | low | 502.4441 | very low | 0.049911 | no | 97.4887 | no | 0.91483 | 0.8 | no | 0.92802 | no | 0.90038 | no | 0.95413 | no | 0.00015291 | 0 | 0 | -0.00132 | yes | 0 | -0.0014718 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
141125_M00766_0062_000000000-A7BKW | M00766 | 2014-11-25 | 151 | 151 | 94.983 | no | 14809543 | low | 762.9367 | very low | 0.018589 | no | 95.0964 | no | 0.89379 | 0.8 | no | 0.96312 | no | 0.82977 | no | 0.24079 | no | 2.4117e-86 | nan | 0 | -0.00041911 | no | 4 | -0.00042946 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
141222_M00766_0067_000000000-ACCB3 | M00766 | 2014-12-22 | 151 | 151 | 67.7653 | yes | 14253970 | low | 760.392 | very low | 0.029815 | no | 68.1874 | yes | 0.82139 | 0.8 | no | 0.85456 | no | 0.78426 | yes | 0.95413 | no | 3.8283e-06 | nan | 0 | -0.0021203 | yes | 11 | -0.0013537 | yes | fail | fail | fail | fail | diff | pass | not exceeded | not exceeded | same | same | Fail
150112_M00766_0068_000000000-A8PD8 | M00766 | 2015-01-12 | 151 | 151 | 94.817 | no | 15773684 | low | 797.5388 | very low | 0.028582 | no | 94.9895 | no | 0.92451 | 0.8 | no | 0.93876 | no | 0.90863 | no | 0.99584 | no | 0.00017877 | 1.2387e-09 | 0 | -0.00090845 | yes | 0 | -0.00084942 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150114_M00766_0069_000000000-AA68B | M00766 | 2015-01-14 | 151 | 151 | 77.8262 | yes | 20704375 | low | 1148.0758 | OK | 0.024397 | no | 77.3315 | yes | 0.86381 | 0.8 | no | 0.91949 | no | 0.81333 | no | 0.95413 | no | 2.6748e-38 | nan | 0 | -0.00060681 | yes | 4 | -0.00057587 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Fail
150116_M00766_0070_000000000-ACC3G | M00766 | 2015-01-16 | 151 | 151 | 86.297 | no | 14147272 | low | 752.8015 | very low | 0.040986 | no | 86.2322 | no | 0.93362 | 0.8 | no | 0.95795 | no | 0.91769 | no | 0.95413 | no | 0.027224 | 4.5409e-07 | 0 | -0.00014326 | no | 3 | 0.00022032 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150127_M00766_0072_000000000-ABRK7 | M00766 | 2015-01-27 | 151 | 151 | 94.2571 | no | 17854832 | low | 926.0782 | low | 0.029511 | no | 94.3435 | no | 0.93257 | 0.8 | no | 0.95668 | no | 0.90691 | no | 0.67794 | no | 1.1008e-18 | 1.0313e-37 | 0 | -0.00054076 | yes | 0 | -0.00067649 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150130_M00766_0073_000000000-ACBVB | M00766 | 2015-01-30 | 251 | 251 | 81.2781 | no | 23929178 | low | 1312.9896 | very high | 0.017972 | no | 81.1332 | yes | 0.66236 | 0.75 | yes | 0.70216 | yes | 0.61962 | yes | 0.84095 | no | 6.7473e-31 | nan | 47 | -0.0033611 | yes | 64 | -0.0034493 | yes | fail | fail | fail | fail | same | pass | exceeded | exceeded | same | same | Fail
150205_M02641_0029_000000000-ACCE2 | M02641 | 2015-02-05 | 151 | 151 | 87.5711 | no | 18965301 | low | 1006.645 | OK | 0.019074 | no | 87.6822 | no | 0.8298 | 0.8 | no | 0.8663 | no | 0.79202 | yes | 0.99584 | no | 1.5184e-12 | 8.2211e-115 | 0 | -0.0022136 | yes | 2 | -0.0018972 | yes | pass | pass | pass | fail | diff | pass | not exceeded | not exceeded | same | same | Borderline Fail
150220_M00766_0077_000000000-AA2U8 | M00766 | 2015-02-20 | 26 | 26 | 11.2989 | yes | 3422864 | low | 175.3978 | very low | 0.069387 | yes | 0 | yes | 0.63211 | 0.9 | yes | 0.88444 | yes | 0.47035 | yes | 0.95413 | no | 6.5681e-12 | 7.5094e-13 | 0 | 0 | no | 26 | 0 | no | fail | fail | fail | fail | same | pass | not exceeded | exceeded | diff | same | Fail
150227_M00766_0079_000000000-ACPR9 | M00766 | 2015-02-27 | 75 | 75 | 91.9723 | no | 29954337 | low | 1225.6634 | OK | 0.019386 | no | 91.9816 | no | 0.96464 | 0.85 | no | 0.97372 | no | 0.95482 | no | 0.84095 | no | 0.028842 | 2.0571e-07 | 0 | -0.00030049 | no | 0 | -0.00039375 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150227_M02641_0032_000000000-ACNF3 | M02641 | 2015-02-27 | 75 | 75 | 90.2941 | no | 25317785 | low | 1066.3007 | very low | 0.076375 | yes | 90.5194 | no | 0.95732 | 0.85 | no | 0.9695 | no | 0.94388 | no | 0.84095 | no | 0.0023737 | nan | 0 | -0.00027283 | no | 0 | -0.00038603 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
150302_M02641_0033_000000000-ACC9K | M02641 | 2015-03-02 | 151 | 151 | 86.4847 | no | 20383236 | low | 1091.5584 | OK | 0.0091317 | no | 86.4229 | no | 0.8794 | 0.8 | no | 0.92333 | no | 0.83429 | no | 0.95413 | no | 4.0711e-39 | 8.175e-86 | 0 | -0.00085422 | yes | 0 | -0.00080079 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
150306_M00766_0080_000000000-ACNEP | M00766 | 2015-03-06 | 75 | 75 | 94.106 | no | 26017554 | low | 1059.9167 | very low | 0.02142 | no | 94.169 | no | 0.97008 | 0.85 | no | 0.97849 | no | 0.96173 | no | 0.84095 | no | 0.078515 | 0.00029715 | 0 | -0.00028625 | no | 0 | -0.00023866 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150306_M02641_0034_000000000-ACRUF | M02641 | 2015-03-06 | 75 | 75 | 91.3955 | no | 24975924 | low | 1049.1583 | very low | 0.0067808 | no | 91.297 | no | 0.95124 | 0.85 | no | 0.9701 | no | 0.93153 | no | 0.95413 | no | 3.4327e-07 | nan | 0 | -0.00040766 | no | 0 | -0.00082529 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150313_M02641_0037_000000000-ACC8R | M02641 | 2015-03-13 | 151 | 151 | 94.5096 | no | 17647098 | low | 906.6463 | low | 0.017228 | no | 93.7282 | no | 0.94213 | 0.8 | no | 0.9595 | no | 0.9242 | no | 0.67794 | no | 2.7478e-10 | nan | 0 | -0.00051188 | yes | 0 | -0.00060154 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150325_M00766_0084_000000000-AD8KN | M00766 | 2015-03-25 | 75 | 75 | 93.0443 | no | 26726137 | low | 1096.1198 | very low | 0.025085 | no | 93.5238 | no | 0.96479 | 0.85 | no | 0.97449 | no | 0.95451 | no | 0.84095 | no | 0.047651 | 0.00076142 | 0 | -0.00034315 | no | 0 | -0.00025609 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150325_M02641_0038_000000000-AD8KY | M02641 | 2015-03-25 | 75 | 75 | 92.8928 | no | 28292154 | low | 1150.8931 | low | 0.016396 | no | 92.9934 | no | 0.96863 | 0.85 | no | 0.97441 | no | 0.9626 | no | 0.95413 | no | 0.97963 | nan | 0 | -0.00030202 | no | 0 | -0.00017631 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150401_M00766_0086_000000000-ACBU5 | M00766 | 2015-04-01 | 151 | 151 | 90.4494 | no | 18801773 | low | 986.6817 | low | 0.029182 | no | 90.7591 | no | 0.8401 | 0.8 | no | 0.87195 | no | 0.80659 | no | 0.99584 | no | 6.8105e-11 | 1.2646e-33 | 0 | -0.0022957 | yes | 6 | -0.0022224 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
150407_M00766_0088_000000000-ACCAT | M00766 | 2015-04-07 | 151 | 151 | 86.0053 | no | 20081955 | low | 1075.4261 | OK | 0.031951 | no | 86.1466 | no | 0.8006 | 0.8 | no | 0.8519 | no | 0.74788 | yes | 0.95413 | no | 6.3802e-32 | nan | 0 | -0.002369 | yes | 12 | -0.001854 | yes | pass | pass | pass | fail | diff | pass | not exceeded | not exceeded | same | same | Borderline Fail
150410_M00766_0089_000000000-ACBU4 | M00766 | 2015-04-10 | 151 | 151 | 94.4563 | no | 16861157 | low | 867.0412 | low | 0.021851 | no | 94.4152 | no | 0.93636 | 0.8 | no | 0.9672 | no | 0.90877 | no | 0.50766 | no | 1.738e-17 | nan | 0 | -0.00043159 | no | 0 | -0.001257 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150423_M02641_0041_000000000-ACN0P | M02641 | 2015-04-23 | 75 | 75 | 91.9691 | no | 30009055 | low | 1228.4129 | OK | 0.027957 | no | 92.2905 | no | 0.9501 | 0.85 | no | 0.96725 | no | 0.93306 | no | 0.84095 | no | 2.6066e-06 | 3.4467e-12 | 0 | -0.00047255 | no | 0 | -0.00049099 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150429_M02641_0043_000000000-AD8KJ | M02641 | 2015-04-29 | 75 | 75 | 77.8177 | yes | 32906358 | low | 1415.0412 | high | 0.018771 | no | 78.178 | yes | 0.8983 | 0.85 | no | 0.92026 | no | 0.87594 | no | 0.84095 | no | 0.00033968 | nan | 0 | -0.00089215 | yes | 0 | -0.00067513 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Fail
150430_M02641_0044_000000000-AD8VC | M02641 | 2015-04-30 | 75 | 75 | 86.9368 | no | 27790413 | low | 1196.1245 | low | 0.022911 | no | 87.4623 | no | 0.92905 | 0.85 | no | 0.95387 | no | 0.90368 | no | 0.84095 | no | 1.2394e-10 | 1.7232e-11 | 0 | -0.00060018 | yes | 0 | -0.0005903 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150501_M00766_0093_000000000-AF9MH | M00766 | 2015-05-01 | 75 | 75 | 87.2253 | no | 34467195 | low | 1450.6461 | high | 0.022012 | no | 87.0197 | no | 0.94139 | 0.85 | no | 0.95639 | no | 0.92569 | no | 0.84095 | no | 0.00057945 | 2.9759e-05 | 0 | -0.00056762 | yes | 0 | -0.00052432 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
150501_M02641_0045_000000000-AD6UA | M02641 | 2015-05-01 | 75 | 75 | 80.8519 | no | 25913017 | low | 1131.0451 | low | 0.015126 | no | 81.3041 | yes | 0.93089 | 0.85 | no | 0.92992 | no | 0.93269 | no | 0.99998 | no | 0.013057 | nan | 0 | -0.00054803 | yes | 0 | -0.00041428 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
150507_M00766_0095_000000000-AF7N3 | M00766 | 2015-05-07 | 75 | 75 | 92.6658 | no | 27410104 | low | 1119.7122 | low | 0.036603 | no | 93.5327 | no | 0.96159 | 0.85 | no | 0.9741 | no | 0.94953 | no | 0.84095 | no | 0.00099419 | 8.626e-08 | 0 | -0.00036389 | no | 0 | -0.00042845 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150507_M02641_0046_000000000-AF7N5 | M02641 | 2015-05-07 | 75 | 75 | 92.9775 | no | 27595623 | low | 1124.3595 | low | 0.040298 | no | 93.1424 | no | 0.95227 | 0.85 | no | 0.96981 | no | 0.93556 | no | 0.84095 | no | 9.3853e-07 | 5.8544e-20 | 0 | -0.00053503 | yes | 0 | -0.0001815 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150515_M00766_0098_000000000-AF9ND | M00766 | 2015-05-15 | 75 | 75 | 96.8787 | no | 15942131 | low | 638.6353 | very low | 0.030068 | no | 96.9762 | no | 0.97995 | 0.85 | no | 0.98705 | no | 0.97446 | no | 0.84095 | no | 0.57572 | nan | 0 | -0.00024399 | no | 0 | -0.00016041 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
150518_M00766_0100_000000000-AATJ7 | M00766 | 2015-05-18 | 151 | 151 | 89.6123 | no | 7467630 | low | 351.2212 | very low | 0.03759 | no | 88.7216 | no | 0.94106 | 0.8 | no | 0.95572 | no | 0.92587 | no | 0.84095 | no | 0.55116 | 0.39443 | 0 | -0.00037129 | no | 0 | -0.00084743 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
150521_M00766_0102_000000000-AF9MV | M00766 | 2015-05-21 | 75 | 75 | 89.6458 | no | 30563261 | low | 1273.6592 | OK | 0.043729 | no | 90.9553 | no | 0.95232 | 0.85 | no | 0.96492 | no | 0.93927 | no | 0.84095 | no | 0.0097477 | nan | 0 | -0.00050706 | yes | 0 | -0.00046026 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150526_M02641_0053_000000000-AFHDV | M02641 | 2015-05-26 | 200 | 200 | 48.4208 | yes | 7559009 | low | 310.8307 | very low | 0.22965 | yes | 50.5562 | yes | 0.64122 | 0.7 | yes | 0.51285 | yes | 0.75917 | no | 0.50766 | no | 1.0758e-55 | 2.1338e-299 | 146 | -0.0039621 | yes | 1 | -0.0018861 | yes | fail | fail | fail | fail | diff | pass | exceeded | not exceeded | diff | same | Fail
150528_M00766_0105_000000000-AF9MJ | M00766 | 2015-05-28 | 75 | 75 | 91.4093 | no | 30434248 | low | 1259.146 | OK | 0.033027 | no | 91.4154 | no | 0.9591 | 0.85 | no | 0.96908 | no | 0.94896 | no | 0.84095 | no | 0.011641 | 5.539e-05 | 0 | -0.00043918 | no | 0 | -0.00032915 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150529_M00766_0106_000000000-AEVP5 | M00766 | 2015-05-29 | 151 | 151 | 88.3589 | no | 15502868 | low | 806.7578 | very low | 0.032653 | no | 89.3692 | no | 0.91331 | 0.8 | no | 0.92662 | no | 0.89824 | no | 0.99584 | no | 0.34729 | 5.8531e-07 | 0 | -0.00082214 | yes | 0 | -0.0010282 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150605_M02641_0058_000000000-AFF15 | M02641 | 2015-06-05 | 75 | 75 | 86.9846 | no | 20446753 | low | 822.314 | very low | 0.026851 | no | 87.5712 | no | 0.93738 | 0.85 | no | 0.94849 | no | 0.92669 | no | 0.95413 | no | 0.71981 | nan | 0 | -0.00062227 | yes | 0 | -0.00041516 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
150610_M02641_0059_000000000-AFN4H | M02641 | 2015-06-10 | 75 | 75 | 82.3028 | no | 25297112 | low | 1022.9657 | very low | 0.024105 | no | 83.0805 | yes | 0.91102 | 0.85 | no | 0.93192 | no | 0.89027 | no | 0.84095 | no | 0.00015742 | 7.1708e-18 | 0 | -0.00077999 | yes | 0 | -0.00063882 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
150612_M00766_0112_000000000-AFMW5 | M00766 | 2015-06-12 | 75 | 75 | 85.2766 | no | 34132691 | low | 1448.1255 | high | 0.019585 | no | 85.0541 | no | 0.93156 | 0.85 | no | 0.95081 | no | 0.91105 | no | 0.84095 | no | 3.193e-06 | 6.9549e-35 | 0 | -0.0006103 | yes | 0 | -0.00064344 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
150612_M02641_0061_000000000-AFMVT | M02641 | 2015-06-12 | 75 | 75 | 79.341 | yes | 25070490 | low | 1016.9582 | very low | 0.03704 | no | 80.1022 | yes | 0.90347 | 0.85 | no | 0.92645 | no | 0.88019 | no | 0.84095 | no | 3.0151e-05 | 3.0045e-11 | 0 | -0.00078584 | yes | 0 | -0.00069582 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
150615_M00766_0113_000000000-AFBH3 | M00766 | 2015-06-15 | 251 | 251 | 89.384 | no | 20944521 | low | 1111.4379 | OK | 0.037506 | no | 89.0846 | no | 0.71363 | 0.75 | yes | 0.81632 | no | 0.60535 | yes | 0.50766 | no | 3.3983e-154 | 5.608e-166 | 24 | -0.0023213 | yes | 53 | -0.004186 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same | Fail
150619_M00766_0114_000000000-AFMYN | M00766 | 2015-06-19 | 75 | 75 | 92.1565 | no | 28125244 | low | 1158.4526 | low | 0.034411 | no | 93.2083 | no | 0.95863 | 0.85 | no | 0.97242 | no | 0.94406 | no | 0.84095 | no | 0.00010309 | nan | 0 | -0.00041056 | no | 0 | -0.00058758 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150622_M00766_0115_000000000-AFN2H | M00766 | 2015-06-22 | 75 | 75 | 93.2875 | no | 26234040 | low | 1075.6113 | very low | 0.020578 | no | 93.4073 | no | 0.96472 | 0.85 | no | 0.97473 | no | 0.95557 | no | 0.84095 | no | 0.048507 | 0.0078432 | 0 | -0.00038391 | no | 0 | -0.0002776 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150623_M00766_0116_000000000-AFLEW | M00766 | 2015-06-23 | 151 | 151 | 87.7026 | no | 20733081 | low | 1104.0961 | OK | 0.034923 | no | 86.8671 | no | 0.84466 | 0.8 | no | 0.89266 | no | 0.79387 | yes | 0.84095 | no | 1.473e-43 | 7.7036e-97 | 0 | -0.0019268 | yes | 9 | -0.0023623 | yes | pass | pass | pass | fail | diff | pass | not exceeded | not exceeded | same | same | Pass
150706_M00766_0118_000000000-AF512 | M00766 | 2015-07-06 | 200 | 200 | 94.4229 | no | 16447526 | low | 844.1499 | very low | 0.016368 | no | 94.5523 | no | 0.86041 | 0.75 | no | 0.89416 | no | 0.82397 | no | 0.84095 | no | 3.1972e-28 | nan | 1 | -0.001537 | yes | 1 | -0.0023974 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150729_M02641_0008_000000000-AGET9 | M02641 | 2015-07-29 | 200 | 200 | 91.862 | no | 30439243 | low | 1258.6104 | OK | 0.012085 | no | 92.3982 | no | 0.84428 | 0.7 | no | 0.87341 | no | 0.81136 | no | 0.67794 | no | 1.7873e-34 | 2.8556e-270 | 1 | -0.0016885 | yes | 2 | -0.0024684 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150730_M00766_0120_000000000-AFN4E | M00766 | 2015-07-30 | 75 | 75 | 94.0611 | no | 23945923 | low | 972.2454 | very low | 0.052603 | yes | 94.2795 | no | 0.9647 | 0.85 | no | 0.97668 | no | 0.95322 | no | 0.84095 | no | 0.0076657 | nan | 0 | -0.00038193 | no | 0 | -0.0003459 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
150731_M00766_0121_000000000-AFMYC | M00766 | 2015-07-31 | 75 | 75 | 89.5212 | no | 27660854 | low | 1152.6219 | low | 0.036004 | no | 89.6942 | no | 0.94787 | 0.85 | no | 0.95439 | no | 0.94056 | no | 0.99584 | no | 0.98595 | nan | 0 | -0.00062082 | yes | 0 | -0.00041404 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150731_M02641_0009_000000000-AFN3M | M02641 | 2015-07-31 | 75 | 75 | 90.2753 | no | 32967275 | low | 1364.8961 | OK | 0.020875 | no | 90.6855 | no | 0.9485 | 0.85 | no | 0.96322 | no | 0.93353 | no | 0.84095 | no | 0.00029106 | 1.2672e-06 | 0 | -0.00052951 | yes | 0 | -0.00044329 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150807_M02641_0012_000000000-AFMYT | M02641 | 2015-08-07 | 75 | 75 | 96.8172 | no | 16466109 | low | 661.2988 | very low | 0.043831 | no | 97.2034 | no | 0.97599 | 0.85 | no | 0.98258 | no | 0.97147 | no | 0.95413 | no | 0.81203 | 0.23825 | 0 | -0.00028401 | no | 0 | -0.0001526 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
150810_M00766_0123_000000000-AGTUY | M00766 | 2015-08-10 | 151 | 151 | 89.2945 | no | 22283310 | low | 1188.9462 | OK | 0.016194 | no | 88.9362 | no | 0.92177 | 0.8 | no | 0.95497 | no | 0.88475 | no | 0.84095 | no | 2.7093e-32 | nan | 0 | -0.00035243 | no | 0 | -0.00083315 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150811_M02641_0014_000000000-AGK0F | M02641 | 2015-08-11 | 151 | 151 | 92.4253 | no | 22253985 | low | 1160.0949 | OK | 0.016889 | no | 92.3445 | no | 0.90106 | 0.8 | no | 0.9283 | no | 0.87132 | no | 0.95413 | no | 7.8772e-25 | nan | 0 | -0.0009934 | yes | 0 | -0.0015021 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150819_M02641_0016_000000000-AGJHU | M02641 | 2015-08-19 | 75 | 75 | 94.1214 | no | 26775268 | low | 1083.954 | very low | 0.032934 | no | 94.4666 | no | 0.96462 | 0.85 | no | 0.97519 | no | 0.95489 | no | 0.84095 | no | 0.013871 | 2.5966e-09 | 0 | -0.00041278 | no | 0 | -0.00025665 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150820_M00766_0126_000000000-AGHUU | M00766 | 2015-08-20 | 75 | 75 | 92.7672 | no | 28026234 | low | 1148.9642 | low | 0.020853 | no | 93.2904 | no | 0.96152 | 0.85 | no | 0.97405 | no | 0.94902 | no | 0.84095 | no | 0.00031602 | nan | 0 | -0.00033273 | no | 0 | -0.0004084 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150820_M02641_0017_000000000-AGJJA | M02641 | 2015-08-20 | 75 | 75 | 93.4116 | no | 26592471 | low | 1089.8361 | very low | 0.015153 | no | 93.7457 | no | 0.96481 | 0.85 | no | 0.97351 | no | 0.95717 | no | 0.84095 | no | 0.072701 | 0.016306 | 0 | -0.00030311 | no | 0 | -0.00031655 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150821_M02641_0018_000000000-AGJNU | M02641 | 2015-08-21 | 151 | 151 | 96.3232 | no | 11702340 | low | 597.5554 | very low | 0.030041 | no | 96.5218 | no | 0.96101 | 0.8 | no | 0.97032 | no | 0.95086 | no | 0.95413 | no | 0.0070316 | nan | 0 | -0.00034988 | no | 0 | -0.000607 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
150902_M00766_0128_000000000-AFLDG | M00766 | 2015-09-02 | 151 | 151 | 89.3025 | no | 21077655 | low | 1129.9527 | OK | 0.015246 | no | 89.1889 | no | 0.83273 | 0.8 | no | 0.88185 | no | 0.78 | yes | 0.67794 | no | 7.5008e-47 | 3.9097e-271 | 2 | -0.0022787 | yes | 13 | -0.0028073 | yes | pass | pass | pass | fail | diff | pass | not exceeded | not exceeded | same | same | Borderline Fail
150904_M02641_0022_000000000-AH991 | M02641 | 2015-09-04 | 200 | 200 | 90.7156 | no | 32586889 | low | 1348.7337 | OK | 0.0084128 | no | 90.8839 | no | 0.80842 | 0.7 | no | 0.83792 | no | 0.774 | no | 0.50766 | no | 6.5299e-46 | 4.3193e-60 | 33 | -0.002385 | yes | 45 | -0.0029542 | yes | pass | pass | pass | pass | same | pass | exceeded | exceeded | same | same | Borderline Fail
150910_M02641_0023_000000000-AFLHH | M02641 | 2015-09-10 | 151 | 151 | 89.5554 | no | 20794433 | low | 1100.488 | OK | 0.02384 | no | 88.8672 | no | 0.85359 | 0.8 | no | 0.88668 | no | 0.81835 | no | 0.95413 | no | 2.3039e-18 | nan | 0 | -0.0021309 | yes | 9 | -0.0024176 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
150911_M00766_0131_000000000-AGLDT | M00766 | 2015-09-11 | 75 | 75 | 91.8669 | no | 29894066 | low | 1233.1842 | OK | 0.017257 | no | 92.0012 | no | 0.95954 | 0.85 | no | 0.97065 | no | 0.94872 | no | 0.84095 | no | 0.0036196 | 0.00029275 | 0 | -0.0004291 | no | 0 | -0.00031438 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150911_M02641_0024_AF50J | M02641 | 2015-09-11 | 151 | 151 | 91.3886 | no | 20193785 | low | 1043.6462 | OK | 0.026601 | no | 91.1381 | no | 0.85976 | 0.8 | no | 0.89102 | no | 0.82694 | no | 0.84095 | no | 4.1997e-18 | nan | 0 | -0.0020335 | yes | 5 | -0.0021409 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
150914_M00766_0132_000000000-AF41F | M00766 | 2015-09-14 | 151 | 151 | 86.8892 | no | 3991745 | low | 184.4777 | very low | 0.012611 | no | 87.677 | no | 0.95124 | 0.8 | no | 0.96221 | no | 0.94056 | no | 0.95413 | no | 0.17881 | 0.0092674 | 0 | -0.00035999 | no | 0 | -0.00086178 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
150925_M02641_0028_000000000-AGK08 | M02641 | 2015-09-25 | 75 | 75 | 95.0618 | no | 24906667 | low | 1010.387 | very low | 0.01775 | no | 95.4096 | no | 0.96734 | 0.85 | no | 0.97861 | no | 0.95784 | no | 0.84095 | no | 0.0070482 | 1.5449e-09 | 0 | -0.00035963 | no | 0 | -0.00022766 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150928_M02641_0029_000000000-AGHV5 | M02641 | 2015-09-28 | 75 | 75 | 93.773 | no | 28292454 | low | 1144.9595 | low | 0.021423 | no | 93.5295 | no | 0.96266 | 0.85 | no | 0.97493 | no | 0.95191 | no | 0.84095 | no | 0.012643 | nan | 0 | -0.00041356 | no | 0 | -0.00026611 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150930_M02641_0031_000000000-AGJGT | M02641 | 2015-09-30 | 75 | 75 | 93.3234 | no | 29303547 | low | 1206.8802 | OK | 0.049805 | no | 93.5156 | no | 0.9573 | 0.85 | no | 0.9715 | no | 0.94372 | no | 0.84095 | no | 7.6841e-05 | nan | 0 | -0.00045184 | no | 0 | -0.00030011 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
151001_M02641_0032_000000000-AGLE0 | M02641 | 2015-10-01 | 75 | 75 | 87.3795 | no | 35725439 | low | 1504.6951 | very high | 0.017939 | no | 86.8081 | no | 0.9359 | 0.85 | no | 0.95659 | no | 0.91522 | no | 0.84095 | no | 5.6512e-08 | nan | 0 | -0.00058153 | yes | 0 | -0.00045814 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Fail
151002_M02641_0033_000000000-AFLDA | M02641 | 2015-10-02 | 151 | 151 | 88.0983 | no | 23364962 | low | 1229.9955 | high | 0.042476 | no | 87.9251 | no | 0.90553 | 0.8 | no | 0.937 | no | 0.87138 | no | 0.84095 | no | 1.4773e-23 | 1.9239e-34 | 0 | -0.00058099 | yes | 0 | -0.0010877 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
151009_M00766_0140_000000000-AGK0B | M00766 | 2015-10-09 | 75 | 75 | 87.3055 | no | 24687917 | low | 1048.2614 | very low | 0.024799 | no | 87.0808 | no | 0.95594 | 0.85 | no | 0.96314 | no | 0.94902 | no | 0.99584 | no | 0.95838 | 0.16034 | 0 | -0.00045147 | no | 0 | -0.00036923 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
151009_M02641_0036_000000000-AFN30 | M02641 | 2015-10-09 | 151 | 151 | 91.4518 | no | 21458630 | low | 1123.3539 | OK | 0.024973 | no | 91.2766 | no | 0.86713 | 0.8 | no | 0.89737 | no | 0.83419 | no | 0.95413 | no | 2.2346e-20 | 0 | 0 | -0.0020801 | yes | 10 | -0.0026428 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
151012_M02641_0037_000000000-AFNA5 | M02641 | 2015-10-12 | 151 | 151 | 93.0707 | no | 19894718 | low | 1031.818 | OK | 0.019111 | no | 93.4892 | no | 0.88194 | 0.8 | no | 0.90915 | no | 0.85233 | no | 0.95413 | no | 1.2493e-20 | 1.0755e-37 | 0 | -0.0019675 | yes | 8 | -0.0026129 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
151019_M02641_0040_000000000-AJDVH | M02641 | 2015-10-19 | 151 | 151 | 90.6324 | no | 24425659 | OK | 1287.4209 | very high | 0.019011 | no | 90.7816 | no | 0.92696 | 0.8 | no | 0.94837 | no | 0.90505 | no | 0.95413 | no | 6.5136e-18 | nan | 0 | -0.00048911 | no | 0 | -0.00086419 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
151021_M00766_0144_000000000-AG028 | M00766 | 2015-10-21 | 75 | 75 | 85.0537 | no | 24774347 | low | 1085.7652 | very low | 0.013905 | no | 84.6897 | yes | 0.94129 | 0.85 | no | 0.95646 | no | 0.92591 | no | 0.84095 | no | 0.0030374 | 4.0367e-11 | 0 | -0.00055587 | yes | 0 | -0.00055374 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
151022_M02641_0042_000000000-AJ5B5 | M02641 | 2015-10-22 | 200 | 200 | 89.6462 | no | 28861426 | low | 1217.8503 | OK | 0.010571 | no | 89.8396 | no | 0.67585 | 0.7 | yes | 0.7507 | no | 0.59122 | yes | 0.67794 | no | 4.6203e-88 | 7.0766e-116 | 12 | -0.0033975 | yes | 75 | -0.0048318 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same | Fail
151026_M00766_0148_000000000-AFLVV | M00766 | 2015-10-26 | 151 | 151 | 88.3189 | no | 18895886 | low | 1013.3256 | OK | 0.032209 | no | 89.0093 | no | 0.93493 | 0.8 | no | 0.95661 | no | 0.91268 | no | 0.84095 | no | 9.4968e-15 | 5.2193e-112 | 0 | -0.00031813 | no | 0 | -0.00073043 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
151102_M02641_0045_000000000-AFN3H | M02641 | 2015-11-02 | 151 | 151 | 94.4283 | no | 15063309 | low | 779.0627 | very low | 0.048767 | no | 94.8292 | no | 0.95449 | 0.8 | no | 0.96887 | no | 0.93956 | no | 0.84095 | no | 2.4447e-05 | 2.8037e-38 | 0 | -0.00027595 | no | 0 | -0.00058311 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
151104_M00766_0151_000000000-AK8EW | M00766 | 2015-11-04 | 151 | 151 | 82.696 | no | 16143415 | low | 883.0439 | low | 0.069792 | yes | 83.2962 | yes | 0.91934 | 0.8 | no | 0.94177 | no | 0.89677 | no | 0.95413 | no | 1.0753e-07 | 5.9859e-42 | 0 | -0.00035735 | no | 0 | -0.00067331 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff | Pass
151106_M00766_0152_000000000-AJ5W5 | M00766 | 2015-11-06 | 151 | 151 | 89.6104 | no | 16168604 | low | 863.0096 | low | 0.024136 | no | 90.7829 | no | 0.9317 | 0.8 | no | 0.95041 | no | 0.91232 | no | 0.95413 | no | 3.0502e-07 | 1.2764e-25 | 0 | -0.00049715 | no | 0 | -0.00054569 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
151111_M02641_0049_000000000-AJD47 | M02641 | 2015-11-11 | 200 | 200 | 92.8535 | no | 30208881 | low | 1238.2172 | OK | 0.021332 | no | 93.316 | no | 0.83991 | 0.7 | no | 0.87018 | no | 0.80592 | no | 0.50766 | no | 8.1049e-40 | 6.164e-109 | 14 | -0.0019638 | yes | 19 | -0.0027647 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
151116_M00766_0154_000000000-AJF4C | M00766 | 2015-11-16 | 151 | 151 | 87.9376 | no | 13835502 | low | 734.2783 | very low | 0.054577 | yes | 88.5491 | no | 0.9313 | 0.8 | no | 0.95109 | no | 0.90966 | no | 0.95413 | no | 5.1553e-08 | nan | 0 | -0.00033626 | no | 0 | -0.00061836 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff | Pass
151119_M02641_0052_000000000-AJJ6B | M02641 | 2015-11-19 | 75 | 75 | 94.5733 | no | 24895626 | low | 1016.266 | very low | 0.029336 | no | 94.9966 | no | 0.96439 | 0.85 | no | 0.97782 | no | 0.95208 | no | 0.84095 | no | 0.00044318 | 1.1267e-06 | 0 | -0.00028501 | no | 0 | -0.00037122 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
151125_M00766_0158_000000000-AJJC1 | M00766 | 2015-11-25 | 151 | 151 | 78.1687 | yes | 3563033 | low | 150.7739 | very low | 0.031274 | no | 78.1863 | yes | 0.9558 | 0.8 | no | 0.96712 | no | 0.94724 | no | 0.95413 | no | 0.17297 | 0.010728 | 0 | -0.00022353 | no | 0 | -0.00043897 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
151130_M00766_0163_000000000-AJGC9 | M00766 | 2015-11-30 | 151 | 151 | 79.0357 | yes | 21816702 | low | 1234.7264 | high | 0.043221 | no | 79.6269 | yes | 0.91879 | 0.8 | no | 0.94269 | no | 0.8946 | no | 0.95413 | no | 4.5777e-17 | 3.5682e-28 | 0 | -0.0002497 | no | 0 | -0.00046641 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
151204_M02641_0056_000000000-AJDC6 | M02641 | 2015-12-04 | 200 | 200 | 93.7017 | no | 25575432 | low | 1042.2805 | very low | 0.036678 | no | 94.2736 | no | 0.85421 | 0.7 | no | 0.88062 | no | 0.82404 | no | 0.50766 | no | 2.3201e-29 | nan | 13 | -0.0018496 | yes | 19 | -0.0025865 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
151208_M02641_0057_000000000-AJHRF | M02641 | 2015-12-08 | 151 | 151 | 97.1211 | no | 4496775 | low | 228.8877 | very low | 0.02523 | no | 97.1074 | no | 0.94501 | 0.8 | no | 0.95817 | no | 0.93055 | no | 0.95413 | no | 0.083233 | 0.003379 | 0 | -0.0007036 | yes | 0 | -0.00058399 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
151210_M00766_0168_000000000-AJDDH | M00766 | 2015-12-10 | 75 | 75 | 91.376 | no | 26145512 | low | 1086.9972 | very low | 0.021546 | no | 90.8931 | no | 0.956 | 0.85 | no | 0.96684 | no | 0.94501 | no | 0.84095 | no | 0.10224 | 0.014367 | 0 | -0.00051081 | yes | 0 | -0.0004594 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
151211_M00766_0169_000000000-AK5DW | M00766 | 2015-12-11 | 300 | 300 | 93.0016 | no | 21872195 | low | 891.8405 | very low | 0.023515 | no | 93.6566 | no | 0.54826 | 0.7 | yes | 0.59108 | yes | 0.48586 | yes | 0.24079 | no | 3.3919e-41 | 2.4542e-59 | 169 | -0.0035153 | yes | 181 | -0.004229 | yes | fail | fail | fail | fail | same | pass | exceeded | exceeded | same | same | Fail
151218_M00766_0171_000000000-AK5FL | M00766 | 2015-12-18 | 226 | 226 | 92.0182 | no | 20535429 | low | 846.6865 | very low | 0.020936 | no | 91.9926 | no | 0.57401 | 0.7 | yes | 0.58212 | yes | 0.54171 | yes | 0.67794 | no | 3.7819e-14 | 8.2184e-29 | 24 | -0.005182 | yes | 97 | -0.0055294 | yes | fail | fail | fail | fail | same | pass | exceeded | exceeded | same | same | Fail
160104_M02641_0062_000000000-AL603 | M02641 | 2016-01-04 | 101 | 101 | 93.4734 | no | 19858211 | low | 1030.1337 | OK | 0.01816 | no | 93.4478 | no | 0.95548 | 0.8 | no | 0.97027 | no | 0.94346 | no | 0.84095 | no | 1.6436e-05 | 1.1366e-31 | 0 | -0.00025903 | no | 0 | -0.00019241 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160105_M00766_0174_000000000-AL4LB | M00766 | 2016-01-05 | 151 | 151 | 91.2681 | no | 14218489 | low | 749.2227 | very low | 0.036601 | no | 91.4658 | no | 0.93032 | 0.8 | no | 0.95273 | no | 0.9068 | no | 0.95413 | no | 1.8718e-11 | nan | 0 | -0.00053469 | yes | 0 | -0.00071583 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160107_M02641_0063_000000000-AJDC9 | M02641 | 2016-01-07 | 75 | 75 | 94.1624 | no | 24963706 | low | 1022.5057 | very low | 0.017799 | no | 94.5508 | no | 0.9694 | 0.85 | no | 0.97637 | no | 0.96261 | no | 0.84095 | no | 0.22798 | nan | 0 | -0.00029644 | no | 0 | -0.00032503 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160108_M02641_0064_000000000-ALNH2 | M02641 | 2016-01-08 | 200 | 200 | 91.9702 | no | 21242610 | low | 1109.3897 | OK | 0.022049 | no | 92.3443 | no | 0.8579 | 0.75 | no | 0.90283 | no | 0.81012 | no | 0.67794 | no | 5.5542e-59 | nan | 0 | -0.0013823 | yes | 1 | -0.0025063 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160115_M00766_0178_000000000-ALKJV | M00766 | 2016-01-15 | 151 | 151 | 83.9557 | no | 15235349 | low | 831.5431 | very low | 0.066527 | yes | 83.4762 | yes | 0.91226 | 0.8 | no | 0.93609 | no | 0.8844 | no | 0.95413 | no | 5.102e-09 | nan | 0 | -0.00057903 | yes | 0 | -0.0011478 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
160118_M02641_0066_000000000-AJETF | M02641 | 2016-01-18 | 75 | 75 | 98.1898 | no | 9321517 | low | 373.3045 | very low | 0.028008 | no | 98.3733 | no | 0.97492 | 0.85 | no | 0.97945 | no | 0.97382 | no | 0.99998 | no | 1 | 0.5283 | 0 | -0.00022747 | no | 0 | -0.000157 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
160119_M02641_0067_000000000-ALN4P | M02641 | 2016-01-19 | 151 | 151 | 87.2125 | no | 18949772 | low | 995.7574 | low | 0.091472 | yes | 86.1573 | no | 0.9035 | 0.8 | no | 0.92594 | no | 0.87789 | no | 0.95413 | no | 2.2467e-08 | nan | 0 | -0.00078266 | yes | 0 | -0.0012195 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
160127_M00766_0001_000000000-ALR2L | M00766 | 2016-01-27 | 151 | 151 | 94.4219 | no | 16728171 | low | 869.4499 | low | 0.024493 | no | 94.8131 | no | 0.94892 | 0.8 | no | 0.95934 | no | 0.93843 | no | 0.95413 | no | 5.3787e-06 | nan | 0 | -0.00045839 | no | 0 | -0.00066696 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
160127_M02641_0069_000000000-AMDH4 | M02641 | 2016-01-27 | 75 | 75 | 76.69 | yes | 29747487 | low | 1328.6387 | OK | 0.096919 | yes | 78.2455 | yes | 0.91416 | 0.85 | no | 0.93371 | no | 0.89378 | no | 0.84095 | no | 0.00062682 | nan | 0 | -0.00074637 | yes | 0 | -0.00037367 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff | Borderline Fail
160128_M00766_0002_000000000-ALAJ1 | M00766 | 2016-01-28 | 75 | 75 | 88.2213 | no | 28975202 | low | 1238.6174 | OK | 0.0080531 | no | 88.3898 | no | 0.95207 | 0.85 | no | 0.96129 | no | 0.94181 | no | 0.84095 | no | 0.078659 | 0.012869 | 0 | -0.00042109 | no | 0 | -0.00041041 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160204_M00766_0004_000000000-AL4H0 | M00766 | 2016-02-04 | 151 | 151 | 95.4825 | no | 14166224 | low | 726.9693 | very low | 0.021398 | no | 95.3691 | no | 0.95549 | 0.8 | no | 0.97197 | no | 0.93852 | no | 0.84095 | no | 3.1187e-14 | 2.8941e-31 | 0 | -0.00029231 | no | 0 | -0.00059973 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
160204_M02641_0072_000000000-AK6CE | M02641 | 2016-02-04 | 75 | 75 | 91.6672 | no | 25918178 | low | 1060.0101 | very low | 0.015091 | no | 91.5669 | no | 0.95246 | 0.85 | no | 0.96475 | no | 0.93912 | no | 0.84095 | no | 0.00023857 | 2.281e-41 | 0 | -0.00053931 | yes | 0 | -0.00055384 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160205_M02641_0073_000000000-ALY9Y | M02641 | 2016-02-05 | 75 | 75 | 81.6149 | no | 30119640 | low | 1331.3992 | OK | 0.025391 | no | 80.9054 | yes | 0.92844 | 0.85 | no | 0.94565 | no | 0.90982 | no | 0.84095 | no | 0.023743 | nan | 0 | -0.0004287 | no | 0 | -0.00027127 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
160208_M00766_0006_000000000-AMF4G | M00766 | 2016-02-08 | 75 | 75 | 85.7346 | no | 32902642 | low | 1404.3574 | high | 0.015626 | no | 85.8175 | no | 0.93494 | 0.85 | no | 0.95198 | no | 0.91641 | no | 0.84095 | no | 1.9762e-05 | 5.3645e-08 | 0 | -0.00062874 | yes | 0 | -0.00051979 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
160208_M02641_0075_000000000-AMF40 | M02641 | 2016-02-08 | 75 | 75 | 79.3999 | yes | 30335883 | low | 1362.3448 | OK | 0.022595 | no | 80.0194 | yes | 0.9085 | 0.85 | no | 0.93381 | no | 0.88187 | no | 0.84095 | no | 9.8314e-08 | 2.2541e-295 | 0 | -0.00078696 | yes | 0 | -0.00037771 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff | Borderline Fail
160209_M02641_0076_000000000-AMF44 | M02641 | 2016-02-09 | 75 | 75 | 89.416 | no | 23928513 | low | 1023.4734 | very low | 0.02047 | no | 89.4341 | no | 0.95438 | 0.85 | no | 0.96439 | no | 0.94402 | no | 0.84095 | no | 0.15811 | 4.2223e-80 | 0 | -0.00036997 | no | 0 | -0.00030272 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160210_M02641_0077_000000000-AME81 | M02641 | 2016-02-10 | 75 | 75 | 93.9536 | no | 19770059 | low | 818.1394 | very low | 0.024181 | no | 94.3148 | no | 0.97229 | 0.85 | no | 0.97917 | no | 0.96582 | no | 0.95413 | no | 0.58184 | 6.9811e-05 | 0 | -0.00021377 | no | 0 | -0.00022485 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
160212_M00766_0008_000000000-AL607 | M00766 | 2016-02-12 | 151 | 151 | 87.5244 | no | 20878061 | low | 1123.9732 | OK | 0.026805 | no | 87.1194 | no | 0.85079 | 0.8 | no | 0.8926 | no | 0.80667 | no | 0.84095 | no | 5.1098e-39 | nan | 0 | -0.002166 | yes | 10 | -0.0030426 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
160216_M02641_0079_000000000-AL75M | M02641 | 2016-02-16 | 151 | 151 | 92.5527 | no | 17259426 | low | 905.0069 | low | 0.019188 | no | 93.4136 | no | 0.9397 | 0.8 | no | 0.96218 | no | 0.91691 | no | 0.67794 | no | 4.7356e-16 | nan | 0 | -0.00030351 | no | 0 | -0.00073347 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
160218_M00766_0010_000000000-AMF48 | M00766 | 2016-02-18 | 75 | 75 | 90.3938 | no | 26783167 | low | 1115.6764 | low | 0.019331 | no | 90.2226 | no | 0.95211 | 0.85 | no | 0.96432 | no | 0.93856 | no | 0.84095 | no | 0.0028637 | 0.00027102 | 0 | -0.00054184 | yes | 0 | -0.00048552 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
160218_M02641_0080_000000000-AMF3Y | M02641 | 2016-02-18 | 75 | 75 | 92.7164 | no | 22919923 | low | 955.1644 | very low | 0.010077 | no | 92.7435 | no | 0.95689 | 0.85 | no | 0.97361 | no | 0.93955 | no | 0.84095 | no | 0.00019591 | nan | 0 | -0.00031985 | no | 0 | -0.00022652 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
| 63 |  |  | NUM FAILS |  | 11 |  | 120 |  | 89 |  | 10 |  | 19 |  |  | 9 |  | 5 |  | 16 |  | 0 |  |  | 16 |  | 66 | 33 |  | 73 | 96 | 96 | 28 | 35 | 13 | 0 |  |  | 8 | 29
|  |  |  | % FAILS |  | 9.0164 |  | 98.3607 |  | 72.9508 |  | 8.1967 |  | 15.5738 |  |  | 7.377 |  | 4.0984 |  | 13.1148 |  | 0 |  |  | 13.1148 |  | 54.0984 | 27.0492 |  | 59.8361 | 26 | 26 | 94 | 87
Exclude this one as useless
[31 empty] | Overall | Method 1 | Method 2 | Method 1 (no CD) | Method 2 (no CD)
[31 empty] | Pass | 26 | 26 | 94 | 87
[31 empty] | Fail | 96 | 96 | 28 | 35

## ClusterDensityPF
RunID | InstrumentID | RunDate | NumberOfCyclesR1 | NumberofCyclesR2 | %ReadsPassingFilter | %ReadsPassingFilterOutsideThreshold? | NumberofReadsPassingFilter | ReadCountWithinThreshold? | MeanClusterDensityPassingFilter | MeanClusterDensityWithinThreshold | COVClusterDensity | COVClusterDensityExceedsThreshold? | %MedianClusterDensityPassingFilter | GapBetweenClusterDensitiesPassingFilterAndNotExceedsThreshold? | ProportionOverQ30 | IlluminaQ30Threshold | ProportionOverQ30ExceedsThreshold? | ProportionOverQ30Read1 | ProportionOverQ30ExceedsThresholdRead1 | ProportionOverQ30Read2 | ProportionOverQ30ExceedsThresholdRead2 | K-SPValue | K-SPValueSignificant? | K-SUnpackPvalue | M-WUnpackPvalue | NumberOfCyclesExceedingThresholdUnderQ30Read1 | SlopeOfLineOfBestFitRead1 | SlopeOfLineOfBestFitRead1ExceedsThreshold? | NumberOfCyclesExceedingThresholdUnderQ30Read2 | SlopeOfLineOfBestFitRead2 | SlopeOfLineOfBestFitRead2ExceedsThreshold? | Fail? | Fail2? | Fail? No CD | Fail2? No CD | R1vsR2PropQ30 | R1vsR2K-S | NumCycleR1Threshold | NumCycleR2Threshold | R1vsR2CyclesQ30 | R1vsR2bestfit | Pass/Fail
130417_M00766_0007_000000000-A3R1F | M00766 | 2013-04-17 | 151 | 151 | 93.6874 | no | 14918743 | low | 773.0359 | very low | 0.024644 | no | 93.6673 | no | 0.97122 | 0.8 | no | 0.98017 | no | 0.96226 | no | 0.84095 | no | 0.0022684 | 4.2948e-06 | 0 | -0.00013884 | no | 0 | -0.00025711 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
130508_M00766_0013_000000000-A3RB5 | M00766 | 2013-05-08 | 151 | 151 | 89.695 | no | 17060727 | low | 897.222 | low | 0.013058 | no | 89.6111 | no | 0.92341 | 0.8 | no | 0.94982 | no | 0.89785 | no | 0.67794 | no | 1.7122e-13 | 9.1751e-65 | 0 | -0.00060464 | yes | 0 | -0.00096571 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
130510_M00766_0014_000000000-A3PFJ | M00766 | 2013-05-10 | 156 | 151 | 88.5142 | no | 18378480 | low | 995.2515 | low | 0.013484 | no | 88.6611 | no | 0.95249 | 0.75 | no | 0.96867 | no | 0.93502 | no | 0.95413 | no | 3.3216e-08 | nan | 0 | -0.0002029 | no | 0 | -0.00047138 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
130613_M00766_0018_000000000-A4FN3 | M00766 | 2013-06-13 | 251 | 251 | 62.2349 | yes | 14759376 | low | 877.4421 | low | 0.0052548 | no | 61.6486 | yes | 0.68299 | 0.75 | yes | 0.7656 | no | 0.58473 | yes | 0.67794 | no | 5.4172e-57 | 1.5425e-86 | 28 | -0.0020629 | yes | 55 | -0.0033012 | yes | fail | fail | fail | fail | diff | pass | exceeded | exceeded | same | same | Fail
130708_M00766_0023_000000000-A20R8 | M00766 | 2013-07-08 | 151 | 151 | 90.9082 | no | 8144995 | None | 872.0913 | None | 0.015455 | no | 90.7757 | no | 0.86589 | 0.85 | no | 0.91799 | no | 0.81047 | yes | 0.67794 | no | 9.6929e-19 | 6.9078e-114 | 0 | -0.0013571 | yes | 0 | -0.0030958 | yes | fail | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same | Pass
130902_M00766_0045_000000000-A5BVD | M00766 | 2013-09-02 | 156 | 151 | 92.4439 | no | 17107054 | low | 889.9218 | low | 0.020943 | no | 92.3467 | no | 0.95012 | 0.75 | no | 0.96778 | no | 0.93094 | no | 0.84095 | no | 1.4736e-06 | nan | 0 | -0.00028935 | no | 1 | -0.00072347 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
130909_M00766_0047_000000000-A5PEH | M00766 | 2013-09-09 | 151 | 151 | 91.8202 | no | 17818894 | low | 943.064 | low | 0.013531 | no | 91.9043 | no | 0.91659 | 0.8 | no | 0.93844 | no | 0.89323 | no | 0.95413 | no | 2.3926e-11 | 3.6095e-18 | 0 | -0.00083727 | yes | 0 | -0.00095948 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
140205_M00766_0021_000000000-A7RVR | M00766 | 2014-02-05 | 251 | 251 | 96.2303 | no | 10640092 | low | 549.8594 | very low | 0.015526 | no | 96.2427 | no | 0.77554 | 0.75 | no | 0.82459 | no | 0.7232 | yes | 0.84095 | no | 3.5805e-28 | 1.8115e-29 | 1 | -0.0023111 | yes | 65 | -0.0033538 | yes | fail | fail | pass | fail | diff | pass | not exceeded | exceeded | diff | same | Fail
140603_M00766_0037_000000000-A7WU9 | M00766 | 2014-06-03 | 251 | 251 | 80.6392 | no | 20279789 | low | 1121.2903 | OK | 0.020185 | no | 81.2363 | yes | 0.67262 | 0.75 | yes | 0.87096 | no | 0.46478 | yes | 0.24079 | no | 0 | 0 | 0 | -0.0011435 | yes | 73 | -0.0033779 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same | Fail
140620_M02641_0004_000000000-A8R7V | M02641 | 2014-06-20 | 151 | 151 | 95.2324 | no | 15895393 | low | 815.6811 | very low | 0.010793 | no | 95.1913 | no | 0.92041 | 0.8 | no | 0.96394 | no | 0.87468 | no | 0.35842 | no | 5.5248e-61 | nan | 0 | -0.0004185 | no | 0 | -0.00070599 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
140625_M00766_0041_000000000-A8P78 | M00766 | 2014-06-25 | 251 | 251 | 93.6991 | no | 12306226 | low | 638.3155 | very low | 0.015947 | no | 93.5516 | no | 0.80794 | 0.75 | no | 0.83721 | no | 0.77401 | no | 0.95413 | no | 1.446e-13 | nan | 15 | -0.0023266 | yes | 34 | -0.0029216 | yes | fail | fail | pass | pass | same | pass | not exceeded | exceeded | diff | same | Fail
140715_M02641_0008_000000000-A8P81 | M02641 | 2014-07-15 | 151 | 151 | 98.3026 | no | 5446264 | low | 277.542 | very low | 0.020978 | no | 98.4264 | no | 0.94495 | 0.8 | no | 0.95412 | no | 0.93542 | no | 0.99584 | no | 0.593 | 0.077423 | 0 | -0.00077434 | yes | 0 | -0.00075567 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
140804_M02641_0011_000000000-AAD6A | M02641 | 2014-08-04 | 251 | 251 | 92.8278 | no | 18236023 | low | 947.4631 | low | 0.014242 | no | 92.5473 | no | 0.75551 | 0.75 | no | 0.79909 | no | 0.70633 | yes | 0.84095 | no | 4.579e-31 | nan | 19 | -0.0025558 | yes | 41 | -0.0033381 | yes | fail | fail | pass | fail | diff | pass | not exceeded | exceeded | diff | same | Fail
140905_M02641_0017_000000000-AA3FN | M02641 | 2014-09-05 | 75 | 75 | 90.0019 | no | 23949422 | low | 1015.424 | very low | 0.055448 | yes | 89.6438 | no | 0.95625 | 0.85 | no | 0.96176 | no | 0.95019 | no | 0.99584 | no | 0.99831 | nan | 0 | -0.00042771 | no | 0 | -0.00021245 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
140917_M02641_0018_000000000-AA3H7 | M02641 | 2014-09-17 | 75 | 75 | 91.5859 | no | 24528146 | low | 1028.5468 | very low | 0.035153 | no | 91.5237 | no | 0.96229 | 0.85 | no | 0.96999 | no | 0.95452 | no | 0.84095 | no | 0.88947 | nan | 0 | -0.00031376 | no | 0 | -0.00016688 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
141024_M02641_0024_000000000-A8P84 | M02641 | 2014-10-24 | 151 | 151 | 94.9748 | no | 15281418 | low | 794.7029 | very low | 0.031108 | no | 95.0899 | no | 0.94073 | 0.8 | no | 0.95744 | no | 0.92359 | no | 0.84095 | no | 1.9766e-08 | 4.7689e-23 | 0 | -0.00055662 | yes | 0 | -0.00062808 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
141113_M02641_0025_000000000-A8RTY | M02641 | 2014-11-13 | 151 | 151 | 97.3253 | no | 9834883 | low | 502.4441 | very low | 0.049911 | no | 97.4887 | no | 0.91483 | 0.8 | no | 0.92802 | no | 0.90038 | no | 0.95413 | no | 0.00015291 | 0 | 0 | -0.00132 | yes | 0 | -0.0014718 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
141125_M00766_0062_000000000-A7BKW | M00766 | 2014-11-25 | 151 | 151 | 94.983 | no | 14809543 | low | 762.9367 | very low | 0.018589 | no | 95.0964 | no | 0.89379 | 0.8 | no | 0.96312 | no | 0.82977 | no | 0.24079 | no | 2.4117e-86 | nan | 0 | -0.00041911 | no | 4 | -0.00042946 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
141222_M00766_0067_000000000-ACCB3 | M00766 | 2014-12-22 | 151 | 151 | 67.7653 | yes | 14253970 | low | 760.392 | very low | 0.029815 | no | 68.1874 | yes | 0.82139 | 0.8 | no | 0.85456 | no | 0.78426 | yes | 0.95413 | no | 3.8283e-06 | nan | 0 | -0.0021203 | yes | 11 | -0.0013537 | yes | fail | fail | fail | fail | diff | pass | not exceeded | not exceeded | same | same | Fail
150112_M00766_0068_000000000-A8PD8 | M00766 | 2015-01-12 | 151 | 151 | 94.817 | no | 15773684 | low | 797.5388 | very low | 0.028582 | no | 94.9895 | no | 0.92451 | 0.8 | no | 0.93876 | no | 0.90863 | no | 0.99584 | no | 0.00017877 | 1.2387e-09 | 0 | -0.00090845 | yes | 0 | -0.00084942 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150114_M00766_0069_000000000-AA68B | M00766 | 2015-01-14 | 151 | 151 | 77.8262 | yes | 20704375 | low | 1148.0758 | OK | 0.024397 | no | 77.3315 | yes | 0.86381 | 0.8 | no | 0.91949 | no | 0.81333 | no | 0.95413 | no | 2.6748e-38 | nan | 0 | -0.00060681 | yes | 4 | -0.00057587 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Fail
150116_M00766_0070_000000000-ACC3G | M00766 | 2015-01-16 | 151 | 151 | 86.297 | no | 14147272 | low | 752.8015 | very low | 0.040986 | no | 86.2322 | no | 0.93362 | 0.8 | no | 0.95795 | no | 0.91769 | no | 0.95413 | no | 0.027224 | 4.5409e-07 | 0 | -0.00014326 | no | 3 | 0.00022032 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150127_M00766_0072_000000000-ABRK7 | M00766 | 2015-01-27 | 151 | 151 | 94.2571 | no | 17854832 | low | 926.0782 | low | 0.029511 | no | 94.3435 | no | 0.93257 | 0.8 | no | 0.95668 | no | 0.90691 | no | 0.67794 | no | 1.1008e-18 | 1.0313e-37 | 0 | -0.00054076 | yes | 0 | -0.00067649 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150130_M00766_0073_000000000-ACBVB | M00766 | 2015-01-30 | 251 | 251 | 81.2781 | no | 23929178 | low | 1312.9896 | very high | 0.017972 | no | 81.1332 | yes | 0.66236 | 0.75 | yes | 0.70216 | yes | 0.61962 | yes | 0.84095 | no | 6.7473e-31 | nan | 47 | -0.0033611 | yes | 64 | -0.0034493 | yes | fail | fail | fail | fail | same | pass | exceeded | exceeded | same | same | Fail
150205_M02641_0029_000000000-ACCE2 | M02641 | 2015-02-05 | 151 | 151 | 87.5711 | no | 18965301 | low | 1006.645 | OK | 0.019074 | no | 87.6822 | no | 0.8298 | 0.8 | no | 0.8663 | no | 0.79202 | yes | 0.99584 | no | 1.5184e-12 | 8.2211e-115 | 0 | -0.0022136 | yes | 2 | -0.0018972 | yes | pass | pass | pass | fail | diff | pass | not exceeded | not exceeded | same | same | Borderline Fail
150220_M00766_0077_000000000-AA2U8 | M00766 | 2015-02-20 | 26 | 26 | 11.2989 | yes | 3422864 | low | 175.3978 | very low | 0.069387 | yes | 0 | yes | 0.63211 | 0.9 | yes | 0.88444 | yes | 0.47035 | yes | 0.95413 | no | 6.5681e-12 | 7.5094e-13 | 0 | 0 | no | 26 | 0 | no | fail | fail | fail | fail | same | pass | not exceeded | exceeded | diff | same | Fail
150227_M00766_0079_000000000-ACPR9 | M00766 | 2015-02-27 | 75 | 75 | 91.9723 | no | 29954337 | low | 1225.6634 | OK | 0.019386 | no | 91.9816 | no | 0.96464 | 0.85 | no | 0.97372 | no | 0.95482 | no | 0.84095 | no | 0.028842 | 2.0571e-07 | 0 | -0.00030049 | no | 0 | -0.00039375 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150227_M02641_0032_000000000-ACNF3 | M02641 | 2015-02-27 | 75 | 75 | 90.2941 | no | 25317785 | low | 1066.3007 | very low | 0.076375 | yes | 90.5194 | no | 0.95732 | 0.85 | no | 0.9695 | no | 0.94388 | no | 0.84095 | no | 0.0023737 | nan | 0 | -0.00027283 | no | 0 | -0.00038603 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
150302_M02641_0033_000000000-ACC9K | M02641 | 2015-03-02 | 151 | 151 | 86.4847 | no | 20383236 | low | 1091.5584 | OK | 0.0091317 | no | 86.4229 | no | 0.8794 | 0.8 | no | 0.92333 | no | 0.83429 | no | 0.95413 | no | 4.0711e-39 | 8.175e-86 | 0 | -0.00085422 | yes | 0 | -0.00080079 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
150306_M00766_0080_000000000-ACNEP | M00766 | 2015-03-06 | 75 | 75 | 94.106 | no | 26017554 | low | 1059.9167 | very low | 0.02142 | no | 94.169 | no | 0.97008 | 0.85 | no | 0.97849 | no | 0.96173 | no | 0.84095 | no | 0.078515 | 0.00029715 | 0 | -0.00028625 | no | 0 | -0.00023866 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150306_M02641_0034_000000000-ACRUF | M02641 | 2015-03-06 | 75 | 75 | 91.3955 | no | 24975924 | low | 1049.1583 | very low | 0.0067808 | no | 91.297 | no | 0.95124 | 0.85 | no | 0.9701 | no | 0.93153 | no | 0.95413 | no | 3.4327e-07 | nan | 0 | -0.00040766 | no | 0 | -0.00082529 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150313_M02641_0037_000000000-ACC8R | M02641 | 2015-03-13 | 151 | 151 | 94.5096 | no | 17647098 | low | 906.6463 | low | 0.017228 | no | 93.7282 | no | 0.94213 | 0.8 | no | 0.9595 | no | 0.9242 | no | 0.67794 | no | 2.7478e-10 | nan | 0 | -0.00051188 | yes | 0 | -0.00060154 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150325_M00766_0084_000000000-AD8KN | M00766 | 2015-03-25 | 75 | 75 | 93.0443 | no | 26726137 | low | 1096.1198 | very low | 0.025085 | no | 93.5238 | no | 0.96479 | 0.85 | no | 0.97449 | no | 0.95451 | no | 0.84095 | no | 0.047651 | 0.00076142 | 0 | -0.00034315 | no | 0 | -0.00025609 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150325_M02641_0038_000000000-AD8KY | M02641 | 2015-03-25 | 75 | 75 | 92.8928 | no | 28292154 | low | 1150.8931 | low | 0.016396 | no | 92.9934 | no | 0.96863 | 0.85 | no | 0.97441 | no | 0.9626 | no | 0.95413 | no | 0.97963 | nan | 0 | -0.00030202 | no | 0 | -0.00017631 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150401_M00766_0086_000000000-ACBU5 | M00766 | 2015-04-01 | 151 | 151 | 90.4494 | no | 18801773 | low | 986.6817 | low | 0.029182 | no | 90.7591 | no | 0.8401 | 0.8 | no | 0.87195 | no | 0.80659 | no | 0.99584 | no | 6.8105e-11 | 1.2646e-33 | 0 | -0.0022957 | yes | 6 | -0.0022224 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
150407_M00766_0088_000000000-ACCAT | M00766 | 2015-04-07 | 151 | 151 | 86.0053 | no | 20081955 | low | 1075.4261 | OK | 0.031951 | no | 86.1466 | no | 0.8006 | 0.8 | no | 0.8519 | no | 0.74788 | yes | 0.95413 | no | 6.3802e-32 | nan | 0 | -0.002369 | yes | 12 | -0.001854 | yes | pass | pass | pass | fail | diff | pass | not exceeded | not exceeded | same | same | Borderline Fail
150410_M00766_0089_000000000-ACBU4 | M00766 | 2015-04-10 | 151 | 151 | 94.4563 | no | 16861157 | low | 867.0412 | low | 0.021851 | no | 94.4152 | no | 0.93636 | 0.8 | no | 0.9672 | no | 0.90877 | no | 0.50766 | no | 1.738e-17 | nan | 0 | -0.00043159 | no | 0 | -0.001257 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150423_M02641_0041_000000000-ACN0P | M02641 | 2015-04-23 | 75 | 75 | 91.9691 | no | 30009055 | low | 1228.4129 | OK | 0.027957 | no | 92.2905 | no | 0.9501 | 0.85 | no | 0.96725 | no | 0.93306 | no | 0.84095 | no | 2.6066e-06 | 3.4467e-12 | 0 | -0.00047255 | no | 0 | -0.00049099 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150429_M02641_0043_000000000-AD8KJ | M02641 | 2015-04-29 | 75 | 75 | 77.8177 | yes | 32906358 | low | 1415.0412 | high | 0.018771 | no | 78.178 | yes | 0.8983 | 0.85 | no | 0.92026 | no | 0.87594 | no | 0.84095 | no | 0.00033968 | nan | 0 | -0.00089215 | yes | 0 | -0.00067513 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Fail
150430_M02641_0044_000000000-AD8VC | M02641 | 2015-04-30 | 75 | 75 | 86.9368 | no | 27790413 | low | 1196.1245 | low | 0.022911 | no | 87.4623 | no | 0.92905 | 0.85 | no | 0.95387 | no | 0.90368 | no | 0.84095 | no | 1.2394e-10 | 1.7232e-11 | 0 | -0.00060018 | yes | 0 | -0.0005903 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150501_M00766_0093_000000000-AF9MH | M00766 | 2015-05-01 | 75 | 75 | 87.2253 | no | 34467195 | low | 1450.6461 | high | 0.022012 | no | 87.0197 | no | 0.94139 | 0.85 | no | 0.95639 | no | 0.92569 | no | 0.84095 | no | 0.00057945 | 2.9759e-05 | 0 | -0.00056762 | yes | 0 | -0.00052432 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
150501_M02641_0045_000000000-AD6UA | M02641 | 2015-05-01 | 75 | 75 | 80.8519 | no | 25913017 | low | 1131.0451 | low | 0.015126 | no | 81.3041 | yes | 0.93089 | 0.85 | no | 0.92992 | no | 0.93269 | no | 0.99998 | no | 0.013057 | nan | 0 | -0.00054803 | yes | 0 | -0.00041428 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
150507_M00766_0095_000000000-AF7N3 | M00766 | 2015-05-07 | 75 | 75 | 92.6658 | no | 27410104 | low | 1119.7122 | low | 0.036603 | no | 93.5327 | no | 0.96159 | 0.85 | no | 0.9741 | no | 0.94953 | no | 0.84095 | no | 0.00099419 | 8.626e-08 | 0 | -0.00036389 | no | 0 | -0.00042845 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150507_M02641_0046_000000000-AF7N5 | M02641 | 2015-05-07 | 75 | 75 | 92.9775 | no | 27595623 | low | 1124.3595 | low | 0.040298 | no | 93.1424 | no | 0.95227 | 0.85 | no | 0.96981 | no | 0.93556 | no | 0.84095 | no | 9.3853e-07 | 5.8544e-20 | 0 | -0.00053503 | yes | 0 | -0.0001815 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150515_M00766_0098_000000000-AF9ND | M00766 | 2015-05-15 | 75 | 75 | 96.8787 | no | 15942131 | low | 638.6353 | very low | 0.030068 | no | 96.9762 | no | 0.97995 | 0.85 | no | 0.98705 | no | 0.97446 | no | 0.84095 | no | 0.57572 | nan | 0 | -0.00024399 | no | 0 | -0.00016041 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
150518_M00766_0100_000000000-AATJ7 | M00766 | 2015-05-18 | 151 | 151 | 89.6123 | no | 7467630 | low | 351.2212 | very low | 0.03759 | no | 88.7216 | no | 0.94106 | 0.8 | no | 0.95572 | no | 0.92587 | no | 0.84095 | no | 0.55116 | 0.39443 | 0 | -0.00037129 | no | 0 | -0.00084743 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
150521_M00766_0102_000000000-AF9MV | M00766 | 2015-05-21 | 75 | 75 | 89.6458 | no | 30563261 | low | 1273.6592 | OK | 0.043729 | no | 90.9553 | no | 0.95232 | 0.85 | no | 0.96492 | no | 0.93927 | no | 0.84095 | no | 0.0097477 | nan | 0 | -0.00050706 | yes | 0 | -0.00046026 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150526_M02641_0053_000000000-AFHDV | M02641 | 2015-05-26 | 200 | 200 | 48.4208 | yes | 7559009 | low | 310.8307 | very low | 0.22965 | yes | 50.5562 | yes | 0.64122 | 0.7 | yes | 0.51285 | yes | 0.75917 | no | 0.50766 | no | 1.0758e-55 | 2.1338e-299 | 146 | -0.0039621 | yes | 1 | -0.0018861 | yes | fail | fail | fail | fail | diff | pass | exceeded | not exceeded | diff | same | Fail
150528_M00766_0105_000000000-AF9MJ | M00766 | 2015-05-28 | 75 | 75 | 91.4093 | no | 30434248 | low | 1259.146 | OK | 0.033027 | no | 91.4154 | no | 0.9591 | 0.85 | no | 0.96908 | no | 0.94896 | no | 0.84095 | no | 0.011641 | 5.539e-05 | 0 | -0.00043918 | no | 0 | -0.00032915 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150529_M00766_0106_000000000-AEVP5 | M00766 | 2015-05-29 | 151 | 151 | 88.3589 | no | 15502868 | low | 806.7578 | very low | 0.032653 | no | 89.3692 | no | 0.91331 | 0.8 | no | 0.92662 | no | 0.89824 | no | 0.99584 | no | 0.34729 | 5.8531e-07 | 0 | -0.00082214 | yes | 0 | -0.0010282 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150605_M02641_0058_000000000-AFF15 | M02641 | 2015-06-05 | 75 | 75 | 86.9846 | no | 20446753 | low | 822.314 | very low | 0.026851 | no | 87.5712 | no | 0.93738 | 0.85 | no | 0.94849 | no | 0.92669 | no | 0.95413 | no | 0.71981 | nan | 0 | -0.00062227 | yes | 0 | -0.00041516 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
150610_M02641_0059_000000000-AFN4H | M02641 | 2015-06-10 | 75 | 75 | 82.3028 | no | 25297112 | low | 1022.9657 | very low | 0.024105 | no | 83.0805 | yes | 0.91102 | 0.85 | no | 0.93192 | no | 0.89027 | no | 0.84095 | no | 0.00015742 | 7.1708e-18 | 0 | -0.00077999 | yes | 0 | -0.00063882 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
150612_M00766_0112_000000000-AFMW5 | M00766 | 2015-06-12 | 75 | 75 | 85.2766 | no | 34132691 | low | 1448.1255 | high | 0.019585 | no | 85.0541 | no | 0.93156 | 0.85 | no | 0.95081 | no | 0.91105 | no | 0.84095 | no | 3.193e-06 | 6.9549e-35 | 0 | -0.0006103 | yes | 0 | -0.00064344 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
150612_M02641_0061_000000000-AFMVT | M02641 | 2015-06-12 | 75 | 75 | 79.341 | yes | 25070490 | low | 1016.9582 | very low | 0.03704 | no | 80.1022 | yes | 0.90347 | 0.85 | no | 0.92645 | no | 0.88019 | no | 0.84095 | no | 3.0151e-05 | 3.0045e-11 | 0 | -0.00078584 | yes | 0 | -0.00069582 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
150615_M00766_0113_000000000-AFBH3 | M00766 | 2015-06-15 | 251 | 251 | 89.384 | no | 20944521 | low | 1111.4379 | OK | 0.037506 | no | 89.0846 | no | 0.71363 | 0.75 | yes | 0.81632 | no | 0.60535 | yes | 0.50766 | no | 3.3983e-154 | 5.608e-166 | 24 | -0.0023213 | yes | 53 | -0.004186 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same | Fail
150619_M00766_0114_000000000-AFMYN | M00766 | 2015-06-19 | 75 | 75 | 92.1565 | no | 28125244 | low | 1158.4526 | low | 0.034411 | no | 93.2083 | no | 0.95863 | 0.85 | no | 0.97242 | no | 0.94406 | no | 0.84095 | no | 0.00010309 | nan | 0 | -0.00041056 | no | 0 | -0.00058758 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150622_M00766_0115_000000000-AFN2H | M00766 | 2015-06-22 | 75 | 75 | 93.2875 | no | 26234040 | low | 1075.6113 | very low | 0.020578 | no | 93.4073 | no | 0.96472 | 0.85 | no | 0.97473 | no | 0.95557 | no | 0.84095 | no | 0.048507 | 0.0078432 | 0 | -0.00038391 | no | 0 | -0.0002776 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150623_M00766_0116_000000000-AFLEW | M00766 | 2015-06-23 | 151 | 151 | 87.7026 | no | 20733081 | low | 1104.0961 | OK | 0.034923 | no | 86.8671 | no | 0.84466 | 0.8 | no | 0.89266 | no | 0.79387 | yes | 0.84095 | no | 1.473e-43 | 7.7036e-97 | 0 | -0.0019268 | yes | 9 | -0.0023623 | yes | pass | pass | pass | fail | diff | pass | not exceeded | not exceeded | same | same | Pass
150706_M00766_0118_000000000-AF512 | M00766 | 2015-07-06 | 200 | 200 | 94.4229 | no | 16447526 | low | 844.1499 | very low | 0.016368 | no | 94.5523 | no | 0.86041 | 0.75 | no | 0.89416 | no | 0.82397 | no | 0.84095 | no | 3.1972e-28 | nan | 1 | -0.001537 | yes | 1 | -0.0023974 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150729_M02641_0008_000000000-AGET9 | M02641 | 2015-07-29 | 200 | 200 | 91.862 | no | 30439243 | low | 1258.6104 | OK | 0.012085 | no | 92.3982 | no | 0.84428 | 0.7 | no | 0.87341 | no | 0.81136 | no | 0.67794 | no | 1.7873e-34 | 2.8556e-270 | 1 | -0.0016885 | yes | 2 | -0.0024684 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150730_M00766_0120_000000000-AFN4E | M00766 | 2015-07-30 | 75 | 75 | 94.0611 | no | 23945923 | low | 972.2454 | very low | 0.052603 | yes | 94.2795 | no | 0.9647 | 0.85 | no | 0.97668 | no | 0.95322 | no | 0.84095 | no | 0.0076657 | nan | 0 | -0.00038193 | no | 0 | -0.0003459 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
150731_M00766_0121_000000000-AFMYC | M00766 | 2015-07-31 | 75 | 75 | 89.5212 | no | 27660854 | low | 1152.6219 | low | 0.036004 | no | 89.6942 | no | 0.94787 | 0.85 | no | 0.95439 | no | 0.94056 | no | 0.99584 | no | 0.98595 | nan | 0 | -0.00062082 | yes | 0 | -0.00041404 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150731_M02641_0009_000000000-AFN3M | M02641 | 2015-07-31 | 75 | 75 | 90.2753 | no | 32967275 | low | 1364.8961 | OK | 0.020875 | no | 90.6855 | no | 0.9485 | 0.85 | no | 0.96322 | no | 0.93353 | no | 0.84095 | no | 0.00029106 | 1.2672e-06 | 0 | -0.00052951 | yes | 0 | -0.00044329 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150807_M02641_0012_000000000-AFMYT | M02641 | 2015-08-07 | 75 | 75 | 96.8172 | no | 16466109 | low | 661.2988 | very low | 0.043831 | no | 97.2034 | no | 0.97599 | 0.85 | no | 0.98258 | no | 0.97147 | no | 0.95413 | no | 0.81203 | 0.23825 | 0 | -0.00028401 | no | 0 | -0.0001526 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
150810_M00766_0123_000000000-AGTUY | M00766 | 2015-08-10 | 151 | 151 | 89.2945 | no | 22283310 | low | 1188.9462 | OK | 0.016194 | no | 88.9362 | no | 0.92177 | 0.8 | no | 0.95497 | no | 0.88475 | no | 0.84095 | no | 2.7093e-32 | nan | 0 | -0.00035243 | no | 0 | -0.00083315 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
150811_M02641_0014_000000000-AGK0F | M02641 | 2015-08-11 | 151 | 151 | 92.4253 | no | 22253985 | low | 1160.0949 | OK | 0.016889 | no | 92.3445 | no | 0.90106 | 0.8 | no | 0.9283 | no | 0.87132 | no | 0.95413 | no | 7.8772e-25 | nan | 0 | -0.0009934 | yes | 0 | -0.0015021 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150819_M02641_0016_000000000-AGJHU | M02641 | 2015-08-19 | 75 | 75 | 94.1214 | no | 26775268 | low | 1083.954 | very low | 0.032934 | no | 94.4666 | no | 0.96462 | 0.85 | no | 0.97519 | no | 0.95489 | no | 0.84095 | no | 0.013871 | 2.5966e-09 | 0 | -0.00041278 | no | 0 | -0.00025665 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150820_M00766_0126_000000000-AGHUU | M00766 | 2015-08-20 | 75 | 75 | 92.7672 | no | 28026234 | low | 1148.9642 | low | 0.020853 | no | 93.2904 | no | 0.96152 | 0.85 | no | 0.97405 | no | 0.94902 | no | 0.84095 | no | 0.00031602 | nan | 0 | -0.00033273 | no | 0 | -0.0004084 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150820_M02641_0017_000000000-AGJJA | M02641 | 2015-08-20 | 75 | 75 | 93.4116 | no | 26592471 | low | 1089.8361 | very low | 0.015153 | no | 93.7457 | no | 0.96481 | 0.85 | no | 0.97351 | no | 0.95717 | no | 0.84095 | no | 0.072701 | 0.016306 | 0 | -0.00030311 | no | 0 | -0.00031655 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150821_M02641_0018_000000000-AGJNU | M02641 | 2015-08-21 | 151 | 151 | 96.3232 | no | 11702340 | low | 597.5554 | very low | 0.030041 | no | 96.5218 | no | 0.96101 | 0.8 | no | 0.97032 | no | 0.95086 | no | 0.95413 | no | 0.0070316 | nan | 0 | -0.00034988 | no | 0 | -0.000607 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
150902_M00766_0128_000000000-AFLDG | M00766 | 2015-09-02 | 151 | 151 | 89.3025 | no | 21077655 | low | 1129.9527 | OK | 0.015246 | no | 89.1889 | no | 0.83273 | 0.8 | no | 0.88185 | no | 0.78 | yes | 0.67794 | no | 7.5008e-47 | 3.9097e-271 | 2 | -0.0022787 | yes | 13 | -0.0028073 | yes | pass | pass | pass | fail | diff | pass | not exceeded | not exceeded | same | same | Borderline Fail
150904_M02641_0022_000000000-AH991 | M02641 | 2015-09-04 | 200 | 200 | 90.7156 | no | 32586889 | low | 1348.7337 | OK | 0.0084128 | no | 90.8839 | no | 0.80842 | 0.7 | no | 0.83792 | no | 0.774 | no | 0.50766 | no | 6.5299e-46 | 4.3193e-60 | 33 | -0.002385 | yes | 45 | -0.0029542 | yes | pass | pass | pass | pass | same | pass | exceeded | exceeded | same | same | Borderline Fail
150910_M02641_0023_000000000-AFLHH | M02641 | 2015-09-10 | 151 | 151 | 89.5554 | no | 20794433 | low | 1100.488 | OK | 0.02384 | no | 88.8672 | no | 0.85359 | 0.8 | no | 0.88668 | no | 0.81835 | no | 0.95413 | no | 2.3039e-18 | nan | 0 | -0.0021309 | yes | 9 | -0.0024176 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
150911_M00766_0131_000000000-AGLDT | M00766 | 2015-09-11 | 75 | 75 | 91.8669 | no | 29894066 | low | 1233.1842 | OK | 0.017257 | no | 92.0012 | no | 0.95954 | 0.85 | no | 0.97065 | no | 0.94872 | no | 0.84095 | no | 0.0036196 | 0.00029275 | 0 | -0.0004291 | no | 0 | -0.00031438 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150911_M02641_0024_AF50J | M02641 | 2015-09-11 | 151 | 151 | 91.3886 | no | 20193785 | low | 1043.6462 | OK | 0.026601 | no | 91.1381 | no | 0.85976 | 0.8 | no | 0.89102 | no | 0.82694 | no | 0.84095 | no | 4.1997e-18 | nan | 0 | -0.0020335 | yes | 5 | -0.0021409 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
150914_M00766_0132_000000000-AF41F | M00766 | 2015-09-14 | 151 | 151 | 86.8892 | no | 3991745 | low | 184.4777 | very low | 0.012611 | no | 87.677 | no | 0.95124 | 0.8 | no | 0.96221 | no | 0.94056 | no | 0.95413 | no | 0.17881 | 0.0092674 | 0 | -0.00035999 | no | 0 | -0.00086178 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
150925_M02641_0028_000000000-AGK08 | M02641 | 2015-09-25 | 75 | 75 | 95.0618 | no | 24906667 | low | 1010.387 | very low | 0.01775 | no | 95.4096 | no | 0.96734 | 0.85 | no | 0.97861 | no | 0.95784 | no | 0.84095 | no | 0.0070482 | 1.5449e-09 | 0 | -0.00035963 | no | 0 | -0.00022766 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150928_M02641_0029_000000000-AGHV5 | M02641 | 2015-09-28 | 75 | 75 | 93.773 | no | 28292454 | low | 1144.9595 | low | 0.021423 | no | 93.5295 | no | 0.96266 | 0.85 | no | 0.97493 | no | 0.95191 | no | 0.84095 | no | 0.012643 | nan | 0 | -0.00041356 | no | 0 | -0.00026611 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
150930_M02641_0031_000000000-AGJGT | M02641 | 2015-09-30 | 75 | 75 | 93.3234 | no | 29303547 | low | 1206.8802 | OK | 0.049805 | no | 93.5156 | no | 0.9573 | 0.85 | no | 0.9715 | no | 0.94372 | no | 0.84095 | no | 7.6841e-05 | nan | 0 | -0.00045184 | no | 0 | -0.00030011 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
151001_M02641_0032_000000000-AGLE0 | M02641 | 2015-10-01 | 75 | 75 | 87.3795 | no | 35725439 | low | 1504.6951 | very high | 0.017939 | no | 86.8081 | no | 0.9359 | 0.85 | no | 0.95659 | no | 0.91522 | no | 0.84095 | no | 5.6512e-08 | nan | 0 | -0.00058153 | yes | 0 | -0.00045814 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Fail
151002_M02641_0033_000000000-AFLDA | M02641 | 2015-10-02 | 151 | 151 | 88.0983 | no | 23364962 | low | 1229.9955 | high | 0.042476 | no | 87.9251 | no | 0.90553 | 0.8 | no | 0.937 | no | 0.87138 | no | 0.84095 | no | 1.4773e-23 | 1.9239e-34 | 0 | -0.00058099 | yes | 0 | -0.0010877 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
151009_M00766_0140_000000000-AGK0B | M00766 | 2015-10-09 | 75 | 75 | 87.3055 | no | 24687917 | low | 1048.2614 | very low | 0.024799 | no | 87.0808 | no | 0.95594 | 0.85 | no | 0.96314 | no | 0.94902 | no | 0.99584 | no | 0.95838 | 0.16034 | 0 | -0.00045147 | no | 0 | -0.00036923 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
151009_M02641_0036_000000000-AFN30 | M02641 | 2015-10-09 | 151 | 151 | 91.4518 | no | 21458630 | low | 1123.3539 | OK | 0.024973 | no | 91.2766 | no | 0.86713 | 0.8 | no | 0.89737 | no | 0.83419 | no | 0.95413 | no | 2.2346e-20 | 0 | 0 | -0.0020801 | yes | 10 | -0.0026428 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
151012_M02641_0037_000000000-AFNA5 | M02641 | 2015-10-12 | 151 | 151 | 93.0707 | no | 19894718 | low | 1031.818 | OK | 0.019111 | no | 93.4892 | no | 0.88194 | 0.8 | no | 0.90915 | no | 0.85233 | no | 0.95413 | no | 1.2493e-20 | 1.0755e-37 | 0 | -0.0019675 | yes | 8 | -0.0026129 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
151019_M02641_0040_000000000-AJDVH | M02641 | 2015-10-19 | 151 | 151 | 90.6324 | no | 24425659 | OK | 1287.4209 | very high | 0.019011 | no | 90.7816 | no | 0.92696 | 0.8 | no | 0.94837 | no | 0.90505 | no | 0.95413 | no | 6.5136e-18 | nan | 0 | -0.00048911 | no | 0 | -0.00086419 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Borderline Pass
151021_M00766_0144_000000000-AG028 | M00766 | 2015-10-21 | 75 | 75 | 85.0537 | no | 24774347 | low | 1085.7652 | very low | 0.013905 | no | 84.6897 | yes | 0.94129 | 0.85 | no | 0.95646 | no | 0.92591 | no | 0.84095 | no | 0.0030374 | 4.0367e-11 | 0 | -0.00055587 | yes | 0 | -0.00055374 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
151022_M02641_0042_000000000-AJ5B5 | M02641 | 2015-10-22 | 200 | 200 | 89.6462 | no | 28861426 | low | 1217.8503 | OK | 0.010571 | no | 89.8396 | no | 0.67585 | 0.7 | yes | 0.7507 | no | 0.59122 | yes | 0.67794 | no | 4.6203e-88 | 7.0766e-116 | 12 | -0.0033975 | yes | 75 | -0.0048318 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same | Fail
151026_M00766_0148_000000000-AFLVV | M00766 | 2015-10-26 | 151 | 151 | 88.3189 | no | 18895886 | low | 1013.3256 | OK | 0.032209 | no | 89.0093 | no | 0.93493 | 0.8 | no | 0.95661 | no | 0.91268 | no | 0.84095 | no | 9.4968e-15 | 5.2193e-112 | 0 | -0.00031813 | no | 0 | -0.00073043 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
151102_M02641_0045_000000000-AFN3H | M02641 | 2015-11-02 | 151 | 151 | 94.4283 | no | 15063309 | low | 779.0627 | very low | 0.048767 | no | 94.8292 | no | 0.95449 | 0.8 | no | 0.96887 | no | 0.93956 | no | 0.84095 | no | 2.4447e-05 | 2.8037e-38 | 0 | -0.00027595 | no | 0 | -0.00058311 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
151104_M00766_0151_000000000-AK8EW | M00766 | 2015-11-04 | 151 | 151 | 82.696 | no | 16143415 | low | 883.0439 | low | 0.069792 | yes | 83.2962 | yes | 0.91934 | 0.8 | no | 0.94177 | no | 0.89677 | no | 0.95413 | no | 1.0753e-07 | 5.9859e-42 | 0 | -0.00035735 | no | 0 | -0.00067331 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff | Pass
151106_M00766_0152_000000000-AJ5W5 | M00766 | 2015-11-06 | 151 | 151 | 89.6104 | no | 16168604 | low | 863.0096 | low | 0.024136 | no | 90.7829 | no | 0.9317 | 0.8 | no | 0.95041 | no | 0.91232 | no | 0.95413 | no | 3.0502e-07 | 1.2764e-25 | 0 | -0.00049715 | no | 0 | -0.00054569 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
151111_M02641_0049_000000000-AJD47 | M02641 | 2015-11-11 | 200 | 200 | 92.8535 | no | 30208881 | low | 1238.2172 | OK | 0.021332 | no | 93.316 | no | 0.83991 | 0.7 | no | 0.87018 | no | 0.80592 | no | 0.50766 | no | 8.1049e-40 | 6.164e-109 | 14 | -0.0019638 | yes | 19 | -0.0027647 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
151116_M00766_0154_000000000-AJF4C | M00766 | 2015-11-16 | 151 | 151 | 87.9376 | no | 13835502 | low | 734.2783 | very low | 0.054577 | yes | 88.5491 | no | 0.9313 | 0.8 | no | 0.95109 | no | 0.90966 | no | 0.95413 | no | 5.1553e-08 | nan | 0 | -0.00033626 | no | 0 | -0.00061836 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff | Pass
151119_M02641_0052_000000000-AJJ6B | M02641 | 2015-11-19 | 75 | 75 | 94.5733 | no | 24895626 | low | 1016.266 | very low | 0.029336 | no | 94.9966 | no | 0.96439 | 0.85 | no | 0.97782 | no | 0.95208 | no | 0.84095 | no | 0.00044318 | 1.1267e-06 | 0 | -0.00028501 | no | 0 | -0.00037122 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
151125_M00766_0158_000000000-AJJC1 | M00766 | 2015-11-25 | 151 | 151 | 78.1687 | yes | 3563033 | low | 150.7739 | very low | 0.031274 | no | 78.1863 | yes | 0.9558 | 0.8 | no | 0.96712 | no | 0.94724 | no | 0.95413 | no | 0.17297 | 0.010728 | 0 | -0.00022353 | no | 0 | -0.00043897 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
151130_M00766_0163_000000000-AJGC9 | M00766 | 2015-11-30 | 151 | 151 | 79.0357 | yes | 21816702 | low | 1234.7264 | high | 0.043221 | no | 79.6269 | yes | 0.91879 | 0.8 | no | 0.94269 | no | 0.8946 | no | 0.95413 | no | 4.5777e-17 | 3.5682e-28 | 0 | -0.0002497 | no | 0 | -0.00046641 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
151204_M02641_0056_000000000-AJDC6 | M02641 | 2015-12-04 | 200 | 200 | 93.7017 | no | 25575432 | low | 1042.2805 | very low | 0.036678 | no | 94.2736 | no | 0.85421 | 0.7 | no | 0.88062 | no | 0.82404 | no | 0.50766 | no | 2.3201e-29 | nan | 13 | -0.0018496 | yes | 19 | -0.0025865 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
151208_M02641_0057_000000000-AJHRF | M02641 | 2015-12-08 | 151 | 151 | 97.1211 | no | 4496775 | low | 228.8877 | very low | 0.02523 | no | 97.1074 | no | 0.94501 | 0.8 | no | 0.95817 | no | 0.93055 | no | 0.95413 | no | 0.083233 | 0.003379 | 0 | -0.0007036 | yes | 0 | -0.00058399 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
151210_M00766_0168_000000000-AJDDH | M00766 | 2015-12-10 | 75 | 75 | 91.376 | no | 26145512 | low | 1086.9972 | very low | 0.021546 | no | 90.8931 | no | 0.956 | 0.85 | no | 0.96684 | no | 0.94501 | no | 0.84095 | no | 0.10224 | 0.014367 | 0 | -0.00051081 | yes | 0 | -0.0004594 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
151211_M00766_0169_000000000-AK5DW | M00766 | 2015-12-11 | 300 | 300 | 93.0016 | no | 21872195 | low | 891.8405 | very low | 0.023515 | no | 93.6566 | no | 0.54826 | 0.7 | yes | 0.59108 | yes | 0.48586 | yes | 0.24079 | no | 3.3919e-41 | 2.4542e-59 | 169 | -0.0035153 | yes | 181 | -0.004229 | yes | fail | fail | fail | fail | same | pass | exceeded | exceeded | same | same | Fail
151218_M00766_0171_000000000-AK5FL | M00766 | 2015-12-18 | 226 | 226 | 92.0182 | no | 20535429 | low | 846.6865 | very low | 0.020936 | no | 91.9926 | no | 0.57401 | 0.7 | yes | 0.58212 | yes | 0.54171 | yes | 0.67794 | no | 3.7819e-14 | 8.2184e-29 | 24 | -0.005182 | yes | 97 | -0.0055294 | yes | fail | fail | fail | fail | same | pass | exceeded | exceeded | same | same | Fail
160104_M02641_0062_000000000-AL603 | M02641 | 2016-01-04 | 101 | 101 | 93.4734 | no | 19858211 | low | 1030.1337 | OK | 0.01816 | no | 93.4478 | no | 0.95548 | 0.8 | no | 0.97027 | no | 0.94346 | no | 0.84095 | no | 1.6436e-05 | 1.1366e-31 | 0 | -0.00025903 | no | 0 | -0.00019241 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160105_M00766_0174_000000000-AL4LB | M00766 | 2016-01-05 | 151 | 151 | 91.2681 | no | 14218489 | low | 749.2227 | very low | 0.036601 | no | 91.4658 | no | 0.93032 | 0.8 | no | 0.95273 | no | 0.9068 | no | 0.95413 | no | 1.8718e-11 | nan | 0 | -0.00053469 | yes | 0 | -0.00071583 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160107_M02641_0063_000000000-AJDC9 | M02641 | 2016-01-07 | 75 | 75 | 94.1624 | no | 24963706 | low | 1022.5057 | very low | 0.017799 | no | 94.5508 | no | 0.9694 | 0.85 | no | 0.97637 | no | 0.96261 | no | 0.84095 | no | 0.22798 | nan | 0 | -0.00029644 | no | 0 | -0.00032503 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160108_M02641_0064_000000000-ALNH2 | M02641 | 2016-01-08 | 200 | 200 | 91.9702 | no | 21242610 | low | 1109.3897 | OK | 0.022049 | no | 92.3443 | no | 0.8579 | 0.75 | no | 0.90283 | no | 0.81012 | no | 0.67794 | no | 5.5542e-59 | nan | 0 | -0.0013823 | yes | 1 | -0.0025063 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160115_M00766_0178_000000000-ALKJV | M00766 | 2016-01-15 | 151 | 151 | 83.9557 | no | 15235349 | low | 831.5431 | very low | 0.066527 | yes | 83.4762 | yes | 0.91226 | 0.8 | no | 0.93609 | no | 0.8844 | no | 0.95413 | no | 5.102e-09 | nan | 0 | -0.00057903 | yes | 0 | -0.0011478 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
160118_M02641_0066_000000000-AJETF | M02641 | 2016-01-18 | 75 | 75 | 98.1898 | no | 9321517 | low | 373.3045 | very low | 0.028008 | no | 98.3733 | no | 0.97492 | 0.85 | no | 0.97945 | no | 0.97382 | no | 0.99998 | no | 1 | 0.5283 | 0 | -0.00022747 | no | 0 | -0.000157 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
160119_M02641_0067_000000000-ALN4P | M02641 | 2016-01-19 | 151 | 151 | 87.2125 | no | 18949772 | low | 995.7574 | low | 0.091472 | yes | 86.1573 | no | 0.9035 | 0.8 | no | 0.92594 | no | 0.87789 | no | 0.95413 | no | 2.2467e-08 | nan | 0 | -0.00078266 | yes | 0 | -0.0012195 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Pass
160127_M00766_0001_000000000-ALR2L | M00766 | 2016-01-27 | 151 | 151 | 94.4219 | no | 16728171 | low | 869.4499 | low | 0.024493 | no | 94.8131 | no | 0.94892 | 0.8 | no | 0.95934 | no | 0.93843 | no | 0.95413 | no | 5.3787e-06 | nan | 0 | -0.00045839 | no | 0 | -0.00066696 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
160127_M02641_0069_000000000-AMDH4 | M02641 | 2016-01-27 | 75 | 75 | 76.69 | yes | 29747487 | low | 1328.6387 | OK | 0.096919 | yes | 78.2455 | yes | 0.91416 | 0.85 | no | 0.93371 | no | 0.89378 | no | 0.84095 | no | 0.00062682 | nan | 0 | -0.00074637 | yes | 0 | -0.00037367 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff | Borderline Fail
160128_M00766_0002_000000000-ALAJ1 | M00766 | 2016-01-28 | 75 | 75 | 88.2213 | no | 28975202 | low | 1238.6174 | OK | 0.0080531 | no | 88.3898 | no | 0.95207 | 0.85 | no | 0.96129 | no | 0.94181 | no | 0.84095 | no | 0.078659 | 0.012869 | 0 | -0.00042109 | no | 0 | -0.00041041 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160204_M00766_0004_000000000-AL4H0 | M00766 | 2016-02-04 | 151 | 151 | 95.4825 | no | 14166224 | low | 726.9693 | very low | 0.021398 | no | 95.3691 | no | 0.95549 | 0.8 | no | 0.97197 | no | 0.93852 | no | 0.84095 | no | 3.1187e-14 | 2.8941e-31 | 0 | -0.00029231 | no | 0 | -0.00059973 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
160204_M02641_0072_000000000-AK6CE | M02641 | 2016-02-04 | 75 | 75 | 91.6672 | no | 25918178 | low | 1060.0101 | very low | 0.015091 | no | 91.5669 | no | 0.95246 | 0.85 | no | 0.96475 | no | 0.93912 | no | 0.84095 | no | 0.00023857 | 2.281e-41 | 0 | -0.00053931 | yes | 0 | -0.00055384 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160205_M02641_0073_000000000-ALY9Y | M02641 | 2016-02-05 | 75 | 75 | 81.6149 | no | 30119640 | low | 1331.3992 | OK | 0.025391 | no | 80.9054 | yes | 0.92844 | 0.85 | no | 0.94565 | no | 0.90982 | no | 0.84095 | no | 0.023743 | nan | 0 | -0.0004287 | no | 0 | -0.00027127 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
160208_M00766_0006_000000000-AMF4G | M00766 | 2016-02-08 | 75 | 75 | 85.7346 | no | 32902642 | low | 1404.3574 | high | 0.015626 | no | 85.8175 | no | 0.93494 | 0.85 | no | 0.95198 | no | 0.91641 | no | 0.84095 | no | 1.9762e-05 | 5.3645e-08 | 0 | -0.00062874 | yes | 0 | -0.00051979 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Fail
160208_M02641_0075_000000000-AMF40 | M02641 | 2016-02-08 | 75 | 75 | 79.3999 | yes | 30335883 | low | 1362.3448 | OK | 0.022595 | no | 80.0194 | yes | 0.9085 | 0.85 | no | 0.93381 | no | 0.88187 | no | 0.84095 | no | 9.8314e-08 | 2.2541e-295 | 0 | -0.00078696 | yes | 0 | -0.00037771 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff | Borderline Fail
160209_M02641_0076_000000000-AMF44 | M02641 | 2016-02-09 | 75 | 75 | 89.416 | no | 23928513 | low | 1023.4734 | very low | 0.02047 | no | 89.4341 | no | 0.95438 | 0.85 | no | 0.96439 | no | 0.94402 | no | 0.84095 | no | 0.15811 | 4.2223e-80 | 0 | -0.00036997 | no | 0 | -0.00030272 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
160210_M02641_0077_000000000-AME81 | M02641 | 2016-02-10 | 75 | 75 | 93.9536 | no | 19770059 | low | 818.1394 | very low | 0.024181 | no | 94.3148 | no | 0.97229 | 0.85 | no | 0.97917 | no | 0.96582 | no | 0.95413 | no | 0.58184 | 6.9811e-05 | 0 | -0.00021377 | no | 0 | -0.00022485 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
160212_M00766_0008_000000000-AL607 | M00766 | 2016-02-12 | 151 | 151 | 87.5244 | no | 20878061 | low | 1123.9732 | OK | 0.026805 | no | 87.1194 | no | 0.85079 | 0.8 | no | 0.8926 | no | 0.80667 | no | 0.84095 | no | 5.1098e-39 | nan | 0 | -0.002166 | yes | 10 | -0.0030426 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same | Borderline Pass
160216_M02641_0079_000000000-AL75M | M02641 | 2016-02-16 | 151 | 151 | 92.5527 | no | 17259426 | low | 905.0069 | low | 0.019188 | no | 93.4136 | no | 0.9397 | 0.8 | no | 0.96218 | no | 0.91691 | no | 0.67794 | no | 4.7356e-16 | nan | 0 | -0.00030351 | no | 0 | -0.00073347 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
160218_M00766_0010_000000000-AMF48 | M00766 | 2016-02-18 | 75 | 75 | 90.3938 | no | 26783167 | low | 1115.6764 | low | 0.019331 | no | 90.2226 | no | 0.95211 | 0.85 | no | 0.96432 | no | 0.93856 | no | 0.84095 | no | 0.0028637 | 0.00027102 | 0 | -0.00054184 | yes | 0 | -0.00048552 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff | Pass
160218_M02641_0080_000000000-AMF3Y | M02641 | 2016-02-18 | 75 | 75 | 92.7164 | no | 22919923 | low | 955.1644 | very low | 0.010077 | no | 92.7435 | no | 0.95689 | 0.85 | no | 0.97361 | no | 0.93955 | no | 0.84095 | no | 0.00019591 | nan | 0 | -0.00031985 | no | 0 | -0.00022652 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same | Pass
| 63 |  |  | NUM FAILS |  | 11 |  | 120 |  | 89 |  | 10 |  | 19 |  |  | 9 |  | 5 |  | 16 |  | 0 |  |  | 16 |  | 66 | 33 |  | 73 | 96 | 96 | 28 | 35 | 13 | 0 |  |  | 8 | 29
|  |  |  | % FAILS |  | 9.0164 |  | 98.3607 |  | 72.9508 |  | 8.1967 |  | 15.5738 |  |  | 7.377 |  | 4.0984 |  | 13.1148 |  | 0 |  |  | 13.1148 |  | 54.0984 | 27.0492 |  | 59.8361 | 26 | 26 | 94 | 87
Exclude this one as useless
[31 empty] | Overall | Method 1 | Method 2 | Method 1 (no CD) | Method 2 (no CD)
[31 empty] | Pass | 26 | 26 | 94 | 87
[31 empty] | Fail | 96 | 96 | 28 | 35