## AnalysisDataTest2OldParams
RunID | InstrumentID | RunDate | NumberOfCyclesR1 | NumberofCyclesR2 | FlowCellInDate? | PR2BottleInDate? | ReagentKitInDate? | %ReadsPassingFilter | %ReadsPassingFilterOutsideThreshold? | NumberofReadsPassingFilter | ReadCountWithinThreshold? | MeanClusterDensityPassingFilter | MeanClusterDensityWithinThreshold | COVClusterDensity | COVClusterDensityExceedsThreshold? | %MedianClusterDensityPassingFilter | GapBetweenClusterDensitiesPassingFilterAndNotExceedsThreshold? | ProportionOverQ30 | IlluminaQ30Threshold | ProportionOverQ30ExceedsThreshold? | ProportionOverQ30Read1 | ProportionOverQ30ExceedsThresholdRead1 | ProportionOverQ30Read2 | ProportionOverQ30ExceedsThresholdRead2 | K-SPValue | K-SPValueSignificant? | K-SUnpackPvalue | M-WUnpackPvalue | NumberOfCyclesExceedingThresholdUnderQ30Read1 | SlopeOfLineOfBestFitRead1 | SlopeOfLineOfBestFitRead1ExceedsThreshold? | NumberOfCyclesExceedingThresholdUnderQ30Read2 | SlopeOfLineOfBestFitRead2 | SlopeOfLineOfBestFitRead2ExceedsThreshold? | ProportionOfUndeterminedReads? | COVIndexes | COVIndexesExceedsThreshold? | OutlyingIndexes | OutlyingIndexes(Low) | OutlyingIndexes(High) | StateOfNTC
121009_M00766_0002_000000000-A1U6P | M00766 | 2012-10-09 | 151 | 151 | in date | in date | in date | 45.385 | yes | 3138050 | low | 341.3139 | very low | 0.34653 | yes | 0 | yes | 0.87575 | 0.8 | no | 0.87642 | no | 0.87507 | no | 0.99998 | no | 0.99997 | 0.80608 | 1 | 0 | no | 0 | 0 | no | No index information for 121009_M00766_0002_000000000-A1U6P
130206_M00766_0002_000000000-A23JM | M00766 | 2013-02-06 | 250 | 250 | in date | in date | in date | 2.0076 | yes | 160923 | low | 10.2917 | very low | 0.26427 | yes | 1.9741 | yes | 0.27654 | 0.75 | yes | 0.091188 | yes | 0.46189 | yes | 0.35842 | no | 0.0011204 | 0.012662 | 193 | -0.0010965 | yes | 65 | -0.0032351 | yes | No index information for 130206_M00766_0002_000000000-A23JM
130618_M00766_0019_000000000-A4FEU | M00766 | 2013-06-18 | 251 | 251 | in date | in date | in date | 93.703 | no | 10147864 | low | 556.7555 | very low | 0.009214 | no | 93.9257 | no | 0.79545 | 0.75 | no | 0.90383 | no | 0.68024 | yes | 0.35842 | no | 1.6395e-83 | 9.0394e-306 | 0 | -0.0010825 | yes | 47 | -0.0037713 | yes | High | 0.1969 | yes | no outlying indexes |  |  | No NTC on run
130624_M00766_0021_000000000-A53PT | M00766 | 2013-06-24 | 251 | 251 | in date | in date | in date | 88.7381 | no | 17725761 | low | 940.8335 | low | 0.011629 | no | 89.0017 | no | 0.84736 | 0.75 | no | 0.90275 | no | 0.78576 | no | 0.67794 | no | 1.0912e-52 | 1.2366e-63 | 0 | -0.0009554 | yes | 4 | -0.0019889 | yes | OK | 0.04822 | no | no outlying indexes |  |  | No NTC on run
130712_M00766_0024_000000000-A5AVV | M00766 | 2013-07-12 | 151 | 151 | in date | in date | in date | 82.4401 | no | 13180142 | low | 710.2954 | very low | 0.09422 | yes | 82.1308 | yes | 0.92195 | 0.8 | no | 0.92377 | no | 0.92292 | no | 0.99998 | no | 0.0096055 | 4.6381e-11 | 0 | -0.00059232 | yes | 0 | -0.00062141 | yes | High | 0.20831 | yes | no outlying indexes |  |  | No NTC on run
130817_M00766_0037_000000000-A4EDD | M00766 | 2013-08-17 | 176 | 176 | in date | in date | in date | 93.613 | no | 14493580 | low | 748.7721 | very low | 0.020101 | no | 93.6127 | no | 0.94897 | 0.75 | no | 0.96312 | no | 0.93388 | no | 0.84095 | no | 4.6236e-05 | nan | 0 | -0.00039473 | no | 1 | -0.0006017 | yes | OK | 0.4371 | yes | ['50005-d' '60003-d' '112001-f' '50005-f'] | ['50005-d' '50005-f'] | ['60003-d' '112001-f'] | No NTC on run
130913_M00766_0049_000000000-A5B0E | M00766 | 2013-09-13 | 251 | 251 | in date | in date | in date | 89.8636 | no | 17433281 | low | 913.7303 | low | 0.024696 | no | 89.1779 | no | 0.83393 | 0.75 | no | 0.88573 | no | 0.77648 | no | 0.84095 | no | 3.8371e-42 | 5.6819e-66 | 0 | -0.0011656 | yes | 4 | -0.0019759 | yes | High | 0.020958 | no | no outlying indexes |  |  | No NTC on run
130916_M00766_0050_000000000-A5M1U | M00766 | 2013-09-16 | 151 | 151 | in date | in date | in date | 86.0618 | no | 22033351 | low | 1185.4173 | OK | 0.028468 | no | 86.0122 | no | 0.9138 | 0.8 | no | 0.93477 | no | 0.89056 | no | 0.84095 | no | 4.3159e-13 | 1.1035e-23 | 0 | -0.00070235 | yes | 0 | -0.0009045 | yes | OK | 0.18994 | yes | ['12M04425'] | ['12M04425'] | [] | No NTC on run
130924_M00766_0052_000000000-A5BE6 | M00766 | 2013-09-24 | 251 | 251 | in date | in date | in date | 67.1952 | yes | 14006895 | low | 805.1388 | very low | 0.092176 | yes | 66.8443 | yes | 0.71517 | 0.75 | yes | 0.77069 | no | 0.66348 | yes | 0.99584 | no | 6.3637e-20 | nan | 16 | -0.0019789 | yes | 32 | -0.002186 | yes | High | 0.49178 | yes | no outlying indexes |  |  | No NTC on run
131007_M00766_0055_000000000-A59JT | M00766 | 2013-10-07 | 251 | 251 | in date | in date | in date | 88.9869 | no | 7066429 | low | 464.0561 | very low | 0.11395 | yes | 90.7266 | no | 0.86005 | 0.75 | no | 0.86612 | no | 0.85247 | no | 1 | no | 0.071573 | nan | 1 | -0.0014022 | yes | 1 | -0.0015446 | yes | High | 0.021021 | no | no outlying indexes |  |  | No NTC on run
131025_M00766_0059_000000000-A5P9A | M00766 | 2013-10-25 | 151 | 151 | in date | in date | in date | 90.6341 | no | 17754371 | low | 930.3591 | low | 0.030836 | no | 91.4036 | no | 0.92242 | 0.8 | no | 0.93549 | no | 0.90546 | no | 0.99584 | no | 1.9427e-05 | nan | 0 | -0.00094972 | yes | 0 | -0.0010274 | yes | OK | 0.51483 | yes | no outlying indexes |  |  | No NTC on run
140207_M00766_0022_000000000-A7PH8 | M00766 | 2014-02-07 | 251 | 251 | in date | in date | in date | 88.3301 | no | 10785103 | low | 558.1003 | very low | 0.053446 | yes | 89.8657 | no | 0.85142 | 0.75 | no | 0.8548 | no | 0.84338 | no | 1 | no | 0.42818 | nan | 1 | -0.0014636 | yes | 1 | -0.0016776 | yes | High | 0.3944 | yes | no outlying indexes |  |  | No NTC on run
140424_M00766_0033_000000000-A7BNA | M00766 | 2014-04-24 | 151 | 151 | in date | in date | in date | 97.0072 | no | 5094671 | low | 251.0711 | very low | 0.231 | yes | 98.5469 | no | 0.93672 | 0.8 | no | 0.94645 | no | 0.92545 | no | 0.95413 | no | 0.083102 | 0.0055314 | 0 | -0.00096088 | yes | 0 | -0.0010359 | yes | OK | 0.49796 | yes | ['10M1732'] | [] | ['10M1732'] | NTC OK
140514_M00766_0036_000000000-A7BRK | M00766 | 2014-05-14 | 151 | 151 | in date | in date | in date | 81.4684 | no | 20334208 | low | 1124.4827 | OK | 0.017101 | no | 81.2844 | yes | 0.79665 | 0.8 | yes | 0.83677 | no | 0.75502 | yes | 0.99584 | no | 4.1492e-14 | nan | 0 | -0.0024166 | yes | 7 | -0.0018547 | yes | High | 0.51173 | yes | no outlying indexes |  |  | No NTC on run
140612_M00766_0039_000000000-A7BNL | M00766 | 2014-06-12 | 251 | 251 | in date | in date | in date | 91.373 | no | 20643048 | low | 1088.1751 | OK | 0.014529 | no | 90.3695 | no | 0.78766 | 0.75 | no | 0.82619 | no | 0.74426 | yes | 0.84095 | no | 1.1316e-41 | nan | 19 | -0.0022137 | yes | 39 | -0.0026621 | yes | OK | 0.11915 | yes | no outlying indexes |  |  | NTC OK
140616_M00766_0040_000000000-A78V9 | M00766 | 2014-06-16 | 251 | 251 | in date | in date | in date | 97.3311 | no | 10710545 | low | 534.5911 | very low | 0.0065097 | no | 97.505 | no | 0.81498 | 0.75 | no | 0.8406 | no | 0.7849 | no | 0.95413 | no | 7.2918e-10 | nan | 1 | -0.0024062 | yes | 30 | -0.0029712 | yes | OK | 0.087674 | yes | no outlying indexes |  |  | NTC OK
140813_M00766_0047_000000000-A8PJL | M00766 | 2014-08-13 | 151 | 151 | in date | in date | in date | 91.9352 | no | 17936069 | low | 943.905 | low | 0.026213 | no | 91.9798 | no | 0.92465 | 0.8 | no | 0.94352 | no | 0.90463 | no | 0.95413 | no | 3.3998e-09 | 3.0143e-185 | 0 | -0.00068224 | yes | 0 | -0.00075951 | yes | OK | 0.35465 | yes | ['11M08823'] | ['11M08823'] | [] | NTC OK
140901_M00766_0048_000000000-AA63M | M00766 | 2014-09-01 | 151 | 151 | in date | in date | in date | 94.0673 | no | 17640397 | low | 909.8289 | low | 0.02007 | no | 93.8944 | no | 0.93862 | 0.8 | no | 0.96072 | no | 0.91538 | no | 0.84095 | no | 4.2142e-16 | 6.6658e-224 | 0 | -0.00044115 | no | 0 | -0.00052041 | yes | OK | 0.22643 | yes | ['14M05256'] | [] | ['14M05256'] | NTC OK
141014_M00766_0053_000000000-A8R6M | M00766 | 2014-10-14 | 151 | 151 | in date | in date | in date | 89.6387 | no | 19084469 | low | 1011.4321 | OK | 0.014006 | no | 90.0342 | no | 0.89321 | 0.8 | no | 0.91865 | no | 0.86561 | no | 0.95413 | no | 8.0153e-17 | 2.8333e-171 | 0 | -0.0010561 | yes | 0 | -0.0011248 | yes | OK | 0.10441 | yes | no outlying indexes |  |  | NTC OK
141017_M02641_0022_000000000-AA66H | M02641 | 2014-10-17 | 151 | 151 | in date | in date | in date | 79.2948 | yes | 22430099 | low | 1210.1492 | high | 0.017994 | no | 78.9705 | yes | 0.87741 | 0.8 | no | 0.91601 | no | 0.83648 | no | 0.99584 | no | 1.1016e-30 | nan | 0 | -0.00063613 | yes | 0 | -0.00060429 | yes | OK | 0.10826 | yes | ['14M02681'] | ['14M02681'] | [] | NTC OK
141024_M00766_0056_000000000-A8PC5 | M00766 | 2014-10-24 | 151 | 151 | in date | in date | in date | 86.6935 | no | 22517755 | low | 1216.7323 | high | 0.017442 | no | 86.7307 | no | 0.88873 | 0.8 | no | 0.92881 | no | 0.84641 | no | 0.84095 | no | 1.3784e-37 | 2.2642e-64 | 0 | -0.00069708 | yes | 0 | -0.00081119 | yes | High | 0.51078 | yes | no outlying indexes |  |  | NTC OK
141031_M00766_0058_000000000-AA3GN | M00766 | 2014-10-31 | 75 | 75 | in date | in date | in date | 84.8049 | no | 22830617 | low | 1020.36 | very low | 0.066963 | yes | 85.4852 | no | 0.94188 | 0.85 | no | 0.94357 | no | 0.93798 | no | 0.99584 | no | 0.16727 | nan | 0 | -0.00072598 | yes | 0 | -0.0003319 | no | OK | 0.42485 | yes | no outlying indexes |  |  | NTC OK
141107_M00766_0060_000000000-AA8PM | M00766 | 2014-11-07 | 151 | 151 | in date | in date | in date | 90.0798 | no | 11262793 | low | 579.4485 | very low | 0.034966 | no | 90.2121 | no | 0.96028 | 0.8 | no | 0.96678 | no | 0.9539 | no | 0.99584 | no | 0.99203 | 0.8978 | 0 | -0.00013806 | no | 0 | -0.00011747 | no | OK | 0.22736 | yes | ['14M08238' '14M09145' '14M09855'] | ['14M08238' '14M09145' '14M09855'] | [] | NTC OK
141118_M00766_0061_000000000-A8P8J | M00766 | 2014-11-18 | 151 | 151 | in date | in date | in date | 92.7322 | no | 19181201 | low | 1009.9074 | OK | 0.027961 | no | 92.312 | no | 0.93572 | 0.8 | no | 0.94773 | no | 0.92218 | no | 0.95413 | no | 1.3904e-05 | nan | 0 | -0.0006231 | yes | 0 | -0.00072415 | yes | OK | 0.13191 | yes | ['14M07799'] | ['14M07799'] | [] | NTC OK
141208_M00766_0063_000000000-A8P7C | M00766 | 2014-12-08 | 151 | 151 | in date | in date | in date | 90.051 | no | 21795042 | low | 1153.0041 | OK | 0.020941 | no | 89.941 | no | 0.89826 | 0.8 | no | 0.94464 | no | 0.85869 | no | 0.67794 | no | 1.8648e-37 | nan | 0 | -0.00060048 | yes | 3 | -0.00051121 | yes | OK | 0.32891 | yes | no outlying indexes |  |  | NTC OK
141208_M02641_0026_000000000-A8R55 | M02641 | 2014-12-08 | 151 | 151 | in date | in date | in date | 94.9912 | no | 13800416 | low | 717.5691 | very low | 0.027596 | no | 95.1078 | no | 0.84608 | 0.8 | no | 0.94004 | no | 0.74763 | yes | 0.35842 | no | 7.7537e-147 | 6.3912e-157 | 0 | -0.00078235 | yes | 5 | -0.0011709 | yes | OK | 0.52447 | yes | no outlying indexes |  |  | NTC OK
141212_M00766_0064_000000000-ACCEB | M00766 | 2014-12-12 | 151 | 151 | in date | in date | in date | 81.1933 | no | 20826514 | low | 1127.546 | OK | 0.021875 | no | 81.0162 | yes | 0.74058 | 0.8 | yes | 0.80816 | no | 0.67993 | yes | 0.99584 | no | 3.8905e-34 | nan | 0 | -0.0028592 | yes | 24 | -0.0015004 | yes | High | 0.28308 | yes | ['6'] | ['6'] | [] | NTC OK
141216_M00766_0065_000000000-ACCDT | M00766 | 2014-12-16 | 151 | 151 | in date | in date | in date | 79.0858 | yes | 20548879 | low | 1120.7265 | OK | 0.02252 | no | 79.0848 | yes | 0.77398 | 0.8 | yes | 0.84162 | no | 0.70405 | yes | 0.84095 | no | 1.4032e-49 | 7.9356e-94 | 0 | -0.0021868 | yes | 16 | -0.00079947 | yes | OK | 0.55632 | yes | ['13M11854'] | [] | ['13M11854'] | NTC OK
141219_M00766_0066_000000000-ACCB1 | M00766 | 2014-12-19 | 151 | 151 | in date | in date | in date | 84.1075 | no | 16554828 | low | 882.4454 | low | 0.030188 | no | 83.544 | yes | 0.71631 | 0.8 | yes | 0.77794 | yes | 0.66061 | yes | 0.99998 | no | 9.0168e-19 | 1.7886e-30 | 0 | -0.0034797 | yes | 18 | -0.0018779 | yes | High | 0.8782 | yes | ['13M12119'] | [] | ['13M12119'] | NTC OK
150120_M00766_0071_000000000-AA63K | M00766 | 2015-01-20 | 151 | 151 | in date | in date | in date | 85.9694 | no | 14794255 | low | 793.1929 | very low | 0.05599 | yes | 87.4592 | no | 0.90779 | 0.8 | no | 0.93487 | no | 0.89285 | no | 0.95413 | no | 0.01591 | nan | 0 | -0.00060834 | yes | 3 | -0.00064423 | yes | OK | 0.44895 | yes | ['9M4318'] | [] | ['9M4318'] | No NTC on run
150127_M02641_0027_000000000-AA65J | M02641 | 2015-01-27 | 151 | 151 | in date | in date | out of date | 94.4613 | no | 15926233 | low | 820.6824 | very low | 0.035198 | no | 94.2959 | no | 0.93014 | 0.8 | no | 0.95953 | no | 0.9008 | no | 0.67794 | no | 1.0379e-16 | 3.674e-28 | 0 | -0.0004646 | no | 0 | -0.00058277 | yes | OK | 0.41652 | yes | no outlying indexes |  |  | NTC OK
150203_M00766_0074_000000000-AAUMH | M00766 | 2015-02-03 | 151 | 151 | in date | in date | in date | 81.6068 | no | 16649356 | low | 884.3212 | low | 0.052428 | yes | 81.2451 | yes | 0.90342 | 0.8 | no | 0.92886 | no | 0.88527 | no | 0.95413 | no | 0.0090932 | 0.028862 | 0 | -0.00047881 | no | 3 | -0.00020141 | no | OK | 0.3567 | yes | ['13M70724' '13M70411' '13M70386' '13M70429' '13M70467'] | ['13M70724' '13M70386'] | ['13M70411' '13M70429' '13M70467'] | NTC OK
150203_M02641_0028_000000000-ACBYG | M02641 | 2015-02-03 | 151 | 151 | in date | in date | in date | 89.643 | no | 18845203 | low | 988.6001 | low | 0.014743 | no | 90.0779 | no | 0.82841 | 0.8 | no | 0.86795 | no | 0.78849 | yes | 0.99584 | no | 1.2512e-16 | 0 | 0 | -0.0022305 | yes | 3 | -0.0018175 | yes | OK | 0.3176 | yes | ['13M12250' '15M00579'] | ['15M00579'] | ['13M12250'] | NTC OK
150205_M00766_0075_000000000-ACC43 | M00766 | 2015-02-05 | 151 | 151 | in date | in date | in date | 93.254 | no | 17676890 | low | 914.0169 | low | 0.038617 | no | 93.3074 | no | 0.8595 | 0.8 | no | 0.89452 | no | 0.83021 | no | 0.95413 | no | 5.2726e-15 | 3.8357e-55 | 0 | -0.0019394 | yes | 4 | -0.0017318 | yes | High | 0.58276 | yes | no outlying indexes |  |  | NTC OK
150224_M00766_0078_000000000-ACMP5 | M00766 | 2015-02-24 | 75 | 75 | in date | in date | in date | 82.4896 | no | 36080266 | low | 1546.85 | very high | 0.024283 | no | 82.282 | yes | 0.93168 | 0.85 | no | 0.94868 | no | 0.91268 | no | 0.84095 | no | 1.1086e-05 | 0 | 0 | -0.00047772 | no | 0 | -0.00047646 | no | OK | 0.22071 | yes | ['14M09093'] | [] | ['14M09093'] | NTC OK
150309_M02641_0035_000000000-ACC3J | M02641 | 2015-03-09 | 151 | 151 | in date | in date | in date | 88.8745 | no | 22186253 | low | 1157.1795 | OK | 0.017088 | no | 89.0446 | no | 0.79385 | 0.8 | yes | 0.83769 | no | 0.74848 | yes | 0.99584 | no | 2.3483e-15 | nan | 0 | -0.002745 | yes | 7 | -0.0021184 | yes | High | 0.30141 | yes | ['15M00579'] | ['15M00579'] | [] | NTC OK
150311_M00766_0081_000000000-ACN0W | M00766 | 2015-03-11 | 75 | 75 | in date | in date | in date | 92.0795 | no | 28700881 | low | 1177.6318 | low | 0.016158 | no | 92.5321 | no | 0.96191 | 0.85 | no | 0.9725 | no | 0.95082 | no | 0.84095 | no | 0.010916 | nan | 0 | -0.00034381 | no | 0 | -0.00028909 | no | OK | 0.30879 | yes | no outlying indexes |  |  | NTC OK
150311_M02641_0036_000000000-AD8VB | M02641 | 2015-03-11 | 75 | 75 | in date | in date | in date | 94.0269 | no | 26337766 | low | 1071.9868 | very low | 0.027519 | no | 93.3836 | no | 0.95837 | 0.85 | no | 0.9654 | no | 0.95151 | no | 0.95413 | no | 0.49156 | nan | 0 | -0.0006745 | yes | 0 | -0.00014209 | no | OK | 0.38297 | yes | ['14M07894'] | [] | ['14M07894'] | NTC OK
150402_M00766_0087_000000000-ACCAY | M00766 | 2015-04-02 | 151 | 151 | in date | in date | in date | 88.6379 | no | 17224632 | low | 896.3347 | low | 0.068342 | yes | 91.2709 | no | 0.85166 | 0.8 | no | 0.88173 | no | 0.81998 | no | 0.95413 | no | 2.2205e-10 | nan | 0 | -0.0022033 | yes | 5 | -0.0021469 | yes | OK | 0.4939 | yes | no outlying indexes |  |  | NTC OK
150414_M00766_0090_000000000-AAU39 | M00766 | 2015-04-14 | 151 | 151 | in date | in date | in date | 86.0407 | no | 14841878 | low | 771.5068 | very low | 0.03988 | no | 87.1861 | no | 0.93627 | 0.8 | no | 0.95058 | no | 0.9214 | no | 0.95413 | no | 0.27827 | 0.21791 | 0 | -0.00025771 | no | 0 | -0.00046336 | no | OK | 0.16948 | yes | ['13M70646'] | ['13M70646'] | [] | NTC OK
150427_M02641_0042_000000000-AD8LB | M02641 | 2015-04-27 | 75 | 75 | in date | in date | in date | 97.4766 | no | 4116334 | low | 164.4147 | very low | 0.057123 | yes | 97.3182 | no | 0.95813 | 0.85 | no | 0.97251 | no | 0.95269 | no | 0.95413 | no | 1 | 0.43484 | 0 | -0.00032162 | no | 0 | -0.00011879 | no | High | 0.47354 | yes | no outlying indexes |  |  | NTC OK
150430_M00766_0092_000000000-AD77P | M00766 | 2015-04-30 | 75 | 75 | in date | in date | in date | 86.4782 | no | 33656164 | low | 1420.3657 | high | 0.029215 | no | 86.0606 | no | 0.94127 | 0.85 | no | 0.95639 | no | 0.92495 | no | 0.95413 | no | 2.1517e-05 | 1.5334e-07 | 0 | -0.00053145 | yes | 0 | -0.00060527 | yes | OK | 0.32903 | yes | no outlying indexes |  |  | NTC OK
150506_M00766_0094_000000000-AEVP8 | M00766 | 2015-05-06 | 151 | 151 | in date | in date | in date | 89.4984 | no | 20688072 | low | 1090.2056 | OK | 0.029888 | no | 89.7217 | no | 0.92458 | 0.8 | no | 0.94568 | no | 0.90199 | no | 0.84095 | no | 4.8432e-13 | nan | 0 | -0.00051215 | yes | 0 | -0.00090332 | yes | OK | 0.23251 | yes | no outlying indexes |  |  | NTC OK
150508_M00766_0096_000000000-AF9MW | M00766 | 2015-05-08 | 75 | 75 | in date | in date | in date | 91.4565 | no | 29237117 | low | 1200.3221 | OK | 0.049362 | no | 92.6001 | no | 0.95684 | 0.85 | no | 0.97022 | no | 0.94333 | no | 0.84095 | no | 0.0058945 | nan | 0 | -0.00039203 | no | 0 | -0.00040344 | no | OK | 0.37127 | yes | no outlying indexes |  |  | NTC OK
150508_M02641_0047_000000000-AF7NG | M02641 | 2015-05-08 | 75 | 75 | in date | in date | in date | 89.2921 | no | 33452889 | low | 1376.6665 | OK | 0.032357 | no | 90.4595 | no | 0.94491 | 0.85 | no | 0.96109 | no | 0.92874 | no | 0.84095 | no | 5.3189e-05 | nan | 0 | -0.00048487 | no | 0 | -0.00044495 | no | OK | 0.37915 | yes | no outlying indexes |  |  | NTC OK
150511_M02641_0048_000000000-AEUGT | M02641 | 2015-05-11 | 250 | 250 | in date | in date | in date | 25.8223 | yes | 3140165 | low | 125.9118 | very low | 0.11338 | yes | 14.4297 | yes | 0.53707 | 0.7 | yes | 0.38185 | yes | 0.6802 | yes | 0.35842 | no | 2.3353e-41 | 4.6079e-40 | 205 | -0.0023104 | yes | 25 | -0.0033967 | yes | OK | 0.17177 | yes | ['14M12235'] | [] | ['14M12235'] | NTC OK
150516_M00766_0099_000000000-AF81F | M00766 | 2015-05-16 | 75 | 75 | in date | in date | in date | 92.0597 | no | 29948630 | low | 1238.4292 | OK | 0.017269 | no | 91.8305 | no | 0.96082 | 0.85 | no | 0.97249 | no | 0.94905 | no | 0.84095 | no | 0.0018358 | 5.1652e-05 | 0 | -0.00036272 | no | 0 | -0.00037743 | no | OK | 0.30556 | yes | no outlying indexes |  |  | NTC OK
150521_M02641_0051_000000000-AF9JH | M02641 | 2015-05-21 | 75 | 75 | in date | in date | in date | 85.4273 | no | 20774323 | low | 847.1236 | very low | 0.093549 | yes | 85.9442 | no | 0.92414 | 0.85 | no | 0.93171 | no | 0.91686 | no | 0.99584 | no | 0.99485 | 0.61684 | 0 | -0.00098979 | yes | 0 | -0.00048486 | no | OK | 0.40967 | yes | no outlying indexes |  |  | NTC OK
150522_M00766_0103_000000000-AF9M3 | M00766 | 2015-05-22 | 75 | 75 | in date | in date | in date | 88.2546 | no | 33587117 | low | 1405.6196 | high | 0.026305 | no | 88.4186 | no | 0.94466 | 0.85 | no | 0.95853 | no | 0.92962 | no | 0.84095 | no | 0.0015346 | nan | 0 | -0.00056188 | yes | 0 | -0.00050146 | yes | OK | 0.39895 | yes | no outlying indexes |  |  | NTC OK
150522_M02641_0052_000000000-AF82F | M02641 | 2015-05-22 | 75 | 75 | in date | in date | in date | 77.7302 | yes | 26557402 | low | 1103.8322 | low | 0.045966 | no | 76.9724 | yes | 0.90712 | 0.85 | no | 0.92279 | no | 0.89107 | no | 0.99584 | no | 0.060099 | 5.5163e-09 | 0 | -0.00079849 | yes | 0 | -0.00064071 | yes | OK | 0.36661 | yes | ['15M05218'] | [] | ['15M05218'] | NTC OK
150601_M00766_0107_000000000-AF9N8 | M00766 | 2015-06-01 | 75 | 75 | in date | in date | in date | 96.709 | no | 17833767 | low | 714.0694 | very low | 0.036487 | no | 97.1935 | no | 0.97916 | 0.85 | no | 0.98619 | no | 0.97361 | no | 0.84095 | no | 0.47307 | nan | 0 | -0.00024948 | no | 0 | -0.00016826 | no | OK | 0.37843 | yes | ['15M05723'] | ['15M05723'] | [] | NTC OK
150602_M00766_0108_000000000-AFMB8 | M00766 | 2015-06-02 | 200 | 200 | in date | in date | in date | 87.6065 | no | 33648503 | low | 1419.0921 | high | 0.017208 | no | 87.2266 | no | 0.8224 | 0.7 | no | 0.84731 | no | 0.7928 | no | 0.84095 | no | 3.1434e-29 | 0 | 19 | -0.0019091 | yes | 33 | -0.0023749 | yes | OK | 0.36723 | yes | ['14M12202'] | [] | ['14M12202'] | NTC OK
150603_M02641_0056_000000000-AEY8R | M02641 | 2015-06-03 | 151 | 151 | in date | in date | in date | 75.779 | yes | 10182405 | low | 516.5408 | very low | 0.052494 | yes | 76.2721 | yes | 0.92128 | 0.8 | no | 0.93202 | no | 0.91084 | no | 0.99584 | no | 0.72835 | nan | 0 | -0.00018443 | no | 0 | -0.00018378 | no | OK | 0.059979 | yes | ['15M05343' '15M05476-2'] | ['15M05476-2'] | ['15M05343'] | NTC OK
150604_M00766_0109_000000000-AFJ62 | M00766 | 2015-06-04 | 75 | 75 | in date | in date | in date | 90.2876 | no | 30654786 | low | 1283.025 | OK | 0.019566 | no | 90.5316 | no | 0.94854 | 0.85 | no | 0.96577 | no | 0.93018 | no | 0.84095 | no | 6.5656e-06 | nan | 0 | -0.0004699 | no | 0 | -0.00054877 | yes | OK | 0.42556 | yes | no outlying indexes |  |  | NTC OK
150604_M02641_0057_000000000-AFL7N | M02641 | 2015-06-04 | 75 | 75 | in date | in date | in date | 83.6851 | no | 20989024 | low | 846.0443 | very low | 0.089565 | yes | 87.1228 | no | 0.92879 | 0.85 | no | 0.94412 | no | 0.91514 | no | 0.84095 | no | 0.077245 | nan | 0 | -0.00064783 | yes | 0 | -0.00048886 | no | OK | 0.76195 | yes | no outlying indexes |  |  | NTC OK
150605_M00766_0110_000000000-AF80F | M00766 | 2015-06-05 | 75 | 75 | in date | in date | in date | 83.8601 | no | 32100609 | low | 1379.9375 | OK | 0.019606 | no | 84.3276 | yes | 0.92414 | 0.85 | no | 0.94606 | no | 0.89846 | no | 0.84095 | no | 2.0225e-07 | nan | 0 | -0.00071066 | yes | 0 | -0.00077961 | yes | OK | 1.153 | yes | ['15M01355'] | [] | ['15M01355'] | NTC OK
150610_M00766_0111_000000000-AFMWF | M00766 | 2015-06-10 | 75 | 75 | in date | in date | in date | 89.8899 | no | 32008549 | low | 1328.1731 | OK | 0.015813 | no | 89.8802 | no | 0.94912 | 0.85 | no | 0.96373 | no | 0.93447 | no | 0.84095 | no | 0.0020072 | nan | 0 | -0.00052315 | yes | 0 | -0.00050342 | yes | OK | 0.42926 | yes | ['14M10378'] | [] | ['14M10378'] | NTC OK
150611_M02641_0060_000000000-AFN2Y | M02641 | 2015-06-11 | 75 | 75 | in date | in date | in date | 76.6898 | yes | 23799396 | low | 947.2643 | very low | 0.054797 | yes | 75.8247 | yes | 0.88682 | 0.85 | no | 0.90551 | no | 0.86738 | no | 0.99584 | no | 0.046537 | 2.6959e-66 | 0 | -0.001158 | yes | 0 | -0.00077535 | yes | OK | 0.39558 | yes | no outlying indexes |  |  | NTC OK
150615_M02641_0062_000000000-AFH7C | M02641 | 2015-06-15 | 75 | 75 | in date | in date | in date | 76.7102 | yes | 21607117 | low | 847.9618 | very low | 0.051033 | yes | 78.099 | yes | 0.89692 | 0.85 | no | 0.91029 | no | 0.88432 | no | 0.99584 | no | 0.37812 | 0.020859 | 0 | -0.00092944 | yes | 0 | -0.00061072 | yes | OK | 0.23706 | yes | ['15M04473'] | [] | ['15M04473'] | NTC OK
150619_M02641_0063_000000000-AFN2W | M02641 | 2015-06-19 | 75 | 75 | in date | in date | in date | 86.313 | no | 22890237 | low | 921.1523 | very low | 0.031622 | no | 86.5093 | no | 0.93054 | 0.85 | no | 0.94446 | no | 0.91735 | no | 0.95413 | no | 0.013174 | nan | 0 | -0.00066355 | yes | 0 | -0.0006793 | yes | OK | 0.37302 | yes | no outlying indexes |  |  | NTC OK
150625_M02641_0065_000000000-AEY2A | M02641 | 2015-06-25 | 151 | 151 | in date | in date | in date | 67.3086 | yes | 7123794 | low | 354.4358 | very low | 0.027403 | no | 65.5091 | yes | 0.90035 | 0.8 | no | 0.9059 | no | 0.89467 | no | 0.99998 | no | 0.37708 | 0.0069801 | 0 | -0.00018719 | no | 0 | -0.00016475 | no | OK | 0.2852 | yes | ['15M06230-2' '15M06230' '15M06510-2'] | ['15M06230-2' '15M06230' '15M06510-2'] | [] | NTC OK
150703_M00766_0117_000000000-AFN2F | M00766 | 2015-07-03 | 75 | 75 | in date | in date | in date | 89.9328 | no | 28300840 | low | 1173.6259 | low | 0.036091 | no | 90.4158 | no | 0.94919 | 0.85 | no | 0.95814 | no | 0.93984 | no | 0.84095 | no | 0.81385 | nan | 0 | -0.00062428 | yes | 0 | -0.00041123 | no | OK | 0.27389 | yes | ['15M06941'] | [] | ['15M06941'] | NTC OK
150703_M02641_0001_000000000-AFMW2 | M02641 | 2015-07-03 | 75 | 75 | in date | in date | in date | 92.3482 | no | 29421610 | low | 1208.9627 | OK | 0.019251 | no | 92.6367 | no | 0.95661 | 0.85 | no | 0.96958 | no | 0.94384 | no | 0.84095 | no | 0.001929 | 2.1798e-15 | 0 | -0.0004435 | no | 0 | -0.00035874 | no | OK | 0.64377 | yes | no outlying indexes |  |  | NTC OK
150709_M02641_0004_000000000-AFMPC | M02641 | 2015-07-09 | 75 | 75 | in date | in date | in date | 94.2018 | no | 27443669 | low | 1115.769 | low | 0.036555 | no | 95.4991 | no | 0.96502 | 0.85 | no | 0.9763 | no | 0.95448 | no | 0.84095 | no | 0.016984 | nan | 0 | -0.00039431 | no | 0 | -0.00024497 | no | OK | 0.38492 | yes | no outlying indexes |  |  | NTC OK
150727_M00766_0119_000000000-AF3TM | M00766 | 2015-07-27 | 151 | 151 | in date | in date | in date | 88.8616 | no | 15391937 | low | 825.3634 | very low | 0.055595 | yes | 88.285 | no | 0.9196 | 0.8 | no | 0.9501 | no | 0.88668 | no | 0.84095 | no | 1.2664e-16 | nan | 0 | -0.00038471 | no | 0 | -0.00086889 | yes | OK | 0.60581 | yes | no outlying indexes |  |  | NTC OK
150727_M02641_0007_000000000-AGEW7 | M02641 | 2015-07-27 | 230 | 230 | in date | in date | in date | 91.5639 | no | 31830282 | low | 1309.0338 | OK | 0.015312 | no | 91.8906 | no | 0.76141 | 0.7 | no | 0.80125 | no | 0.71555 | no | 0.35842 | no | 5.1559e-54 | 2.0554e-121 | 29 | -0.0026602 | yes | 56 | -0.0035322 | yes | OK | 0.33395 | yes | ['15M06635'] | [] | ['15M06635'] | NTC OK
150807_M00766_0122_000000000-AFMW7 | M00766 | 2015-08-07 | 75 | 75 | in date | in date | in date | 95.8884 | no | 19129362 | low | 771.607 | very low | 0.031184 | no | 96.3208 | no | 0.9746 | 0.85 | no | 0.98266 | no | 0.96788 | no | 0.84095 | no | 0.16299 | nan | 0 | -0.00031347 | no | 0 | -0.0002468 | no | OK | 0.36606 | yes | no outlying indexes |  |  | NTC OK
150813_M00766_0124_000000000-AFYEH | M00766 | 2015-08-13 | 75 | 75 | in date | in date | in date | 94.3227 | no | 25145640 | low | 1025.335 | very low | 0.021357 | no | 94.2549 | no | 0.96769 | 0.85 | no | 0.97825 | no | 0.95812 | no | 0.84095 | no | 0.0084968 | nan | 0 | -0.00033121 | no | 0 | -0.00034652 | no | OK | 0.23439 | yes | no outlying indexes |  |  | NTC OK
150818_M00766_0125_000000000-AE8BP | M00766 | 2015-08-18 | 151 | 151 | in date | in date | in date | 91.3886 | no | 15351123 | low | 804.6945 | very low | 0.058805 | yes | 90.8408 | no | 0.92459 | 0.8 | no | 0.94957 | no | 0.89589 | no | 0.84095 | no | 6.1155e-16 | 3.3622e-25 | 0 | -0.00052265 | yes | 0 | -0.0011087 | yes | OK | 0.31851 | yes | ['15M08233'] | [] | ['15M08233'] | NTC OK
150826_M00766_0127_000000000-AGKLT | M00766 | 2015-08-26 | 200 | 200 | in date | in date | in date | 93.1048 | no | 24799829 | low | 1017.4002 | very low | 0.016826 | no | 93.586 | no | 0.88995 | 0.7 | no | 0.91037 | no | 0.86811 | no | 0.67794 | no | 4.0958e-20 | nan | 0 | -0.0013198 | yes | 0 | -0.0017885 | yes | OK | 0.094928 | yes | no outlying indexes |  |  | NTC OK
150904_M00766_0129_000000000-AFNA2 | M00766 | 2015-09-04 | 151 | 151 | in date | in date | in date | 84.0182 | no | 18908531 | low | 1029.9385 | OK | 0.024332 | no | 83.7698 | yes | 0.81255 | 0.8 | no | 0.86788 | no | 0.75533 | yes | 0.95413 | no | 5.0649e-39 | nan | 0 | -0.0022068 | yes | 9 | -0.0021487 | yes | OK | 0.37398 | yes | ['15M04668' '15M04664'] | ['15M04664'] | ['15M04668'] | NTC OK
150910_M00766_0130_000000000-AGJGG | M00766 | 2015-09-10 | 75 | 75 | in date | in date | in date | 88.355 | no | 24463097 | low | 1059.3697 | very low | 0.0095777 | no | 88.2815 | no | 0.94136 | 0.85 | no | 0.95905 | no | 0.92225 | no | 0.95413 | no | 1.5264e-06 | 1.3678e-23 | 0 | -0.00056717 | yes | 0 | -0.00075898 | yes | OK | 0.28101 | yes | no outlying indexes |  |  | NTC OK
150916_M02641_0025_000000000-AFN33 | M02641 | 2015-09-16 | 151 | 151 | in date | in date | in date | 91.2782 | no | 21293722 | low | 1112.3198 | OK | 0.023575 | no | 90.9013 | no | 0.8777 | 0.8 | no | 0.90422 | no | 0.84924 | no | 0.95413 | no | 3.9616e-16 | nan | 0 | -0.0018497 | yes | 5 | -0.0021883 | yes | OK | 0.10958 | yes | ['15M04682'] | ['15M04682'] | [] | NTC OK
150917_M02641_0026_000000000-AFN0A | M02641 | 2015-09-17 | 151 | 151 | in date | in date | in date | 91.5269 | no | 20122988 | low | 1048.5468 | OK | 0.02524 | no | 91.3116 | no | 0.86192 | 0.8 | no | 0.89213 | no | 0.82928 | no | 0.84095 | no | 2.1049e-19 | nan | 1 | -0.0022155 | yes | 10 | -0.0026282 | yes | OK | 0.18217 | yes | ['15M08778'] | ['15M08778'] | [] | NTC OK
150918_M00766_0133_000000000-AH985 | M00766 | 2015-09-18 | 200 | 200 | in date | in date | in date | 91.3077 | no | 30062610 | low | 1253.19 | OK | 0.031977 | no | 91.5655 | no | 0.84979 | 0.7 | no | 0.88478 | no | 0.81133 | no | 0.67794 | no | 4.0326e-70 | nan | 11 | -0.0017023 | yes | 24 | -0.0021771 | yes | OK | 0.15671 | yes | ['15M06627'] | [] | ['15M06627'] | NTC OK
150923_M00766_0134_000000000-AFN32 | M00766 | 2015-09-23 | 151 | 151 | in date | in date | in date | 89.0495 | no | 20113067 | low | 1072.9998 | OK | 0.033799 | no | 88.906 | no | 0.86242 | 0.8 | no | 0.89294 | no | 0.82952 | no | 0.95413 | no | 1.619e-18 | nan | 0 | -0.002005 | yes | 9 | -0.0025255 | yes | OK | 0.35061 | yes | ['15M08810' '15M08808'] | ['15M08810'] | ['15M08808'] | NTC OK
150924_M02641_0027_000000000-AGJGM | M02641 | 2015-09-24 | 75 | 75 | in date | in date | in date | 95.0401 | no | 24554073 | low | 997.0288 | very low | 0.0085367 | no | 95.0918 | no | 0.97292 | 0.85 | no | 0.98196 | no | 0.96511 | no | 0.84095 | no | 0.019748 | nan | 0 | -0.00019383 | no | 0 | -0.00024233 | no | OK | 0.30211 | yes | ['15M10058'] | [] | ['15M10058'] | NTC OK
150925_M00766_0135_000000000-AFL75 | M00766 | 2015-09-25 | 151 | 151 | in date | in date | in date | 97.708 | no | 9080878 | low | 460.5892 | very low | 0.039145 | no | 97.9623 | no | 0.94702 | 0.8 | no | 0.95724 | no | 0.93649 | no | 0.95413 | no | 0.0051842 | nan | 0 | -0.00069048 | yes | 0 | -0.0008509 | yes | OK | 0.46239 | yes | no outlying indexes |  |  | NTC OK
150929_M02641_0030_000000000-AFWJ3 | M02641 | 2015-09-29 | 151 | 151 | in date | in date | in date | 95.2167 | no | 14824546 | low | 765.6339 | very low | 0.037672 | no | 95.287 | no | 0.94667 | 0.8 | no | 0.96411 | no | 0.92657 | no | 0.84095 | no | 2.2694e-08 | nan | 0 | -0.00044841 | no | 0 | -0.00094835 | yes | OK | 0.48248 | yes | ['15M10676' '15M09021'] | ['15M10676' '15M09021'] | [] | NTC OK
150930_M00766_0138_000000000-AH4BR | M00766 | 2015-09-30 | 75 | 75 | in date | in date | in date | 90.8515 | no | 29081266 | low | 1203.4983 | OK | 0.018375 | no | 90.911 | no | 0.93455 | 0.85 | no | 0.91651 | no | 0.95034 | no | 0.99584 | no | 5.9019e-07 | nan | 0 | -0.0013505 | yes | 0 | -0.00035323 | no | OK | 0.37912 | yes | no outlying indexes |  |  | NTC OK
151005_M02641_0034_000000000-AH9ML | M02641 | 2015-10-05 | 200 | 200 | in date | in date | in date | 91.5287 | no | 33902064 | low | 1389.8402 | OK | 0.014042 | no | 91.6743 | no | 0.83382 | 0.7 | no | 0.86254 | no | 0.80114 | no | 0.50766 | no | 1.3625e-44 | nan | 23 | -0.0021037 | yes | 30 | -0.0027427 | yes | OK | 0.17597 | yes | ['15M06635'] | [] | ['15M06635'] | NTC OK
151008_M02641_0035_000000000-AFL73 | M02641 | 2015-10-08 | 151 | 151 | in date | in date | in date | 92.9471 | no | 20624078 | low | 1070.1471 | OK | 0.02054 | no | 93.2815 | no | 0.88534 | 0.8 | no | 0.91095 | no | 0.85745 | no | 0.95413 | no | 1.7378e-17 | nan | 0 | -0.0018668 | yes | 5 | -0.0023798 | yes | OK | 1.2158 | yes | ['15M09860' '15M09850' '15M09844'] | [] | ['15M09860' '15M09850' '15M09844'] | NTC OK
151013_M02641_0038_000000000-AFLW0 | M02641 | 2015-10-13 | 151 | 151 | in date | in date | in date | 92.2835 | no | 19868115 | low | 1036.6449 | OK | 0.015868 | no | 92.3144 | no | 0.85852 | 0.8 | no | 0.89862 | no | 0.81584 | no | 0.84095 | no | 9.4078e-32 | 1.0701e-96 | 0 | -0.0019454 | yes | 7 | -0.0029862 | yes | OK | 0.1712 | yes | no outlying indexes |  |  | NTC OK
151014_M02641_0039_000000000-AFLFV | M02641 | 2015-10-14 | 151 | 151 | in date | in date | in date | 93.1322 | no | 20043314 | low | 1036.9175 | OK | 0.024404 | no | 93.1321 | no | 0.86964 | 0.8 | no | 0.90574 | no | 0.83081 | no | 0.84095 | no | 1.4818e-31 | nan | 0 | -0.0019202 | yes | 8 | -0.0027745 | yes | OK | 0.27897 | yes | ['15M10686'] | ['15M10686'] | [] | NTC OK
151015_M00766_0142_000000000-AJD8B | M00766 | 2015-10-15 | 151 | 151 | in date | in date | in date | 88.82 | no | 16091719 | low | 856.7867 | low | 0.030984 | no | 88.9965 | no | 0.9206 | 0.8 | no | 0.95153 | no | 0.88855 | no | 0.67794 | no | 9.7576e-25 | nan | 0 | -0.0004536 | no | 1 | -0.0013313 | yes | OK | 0.2848 | yes | ['15M10752' '15M10508'] | ['15M10508'] | ['15M10752'] | NTC OK
151021_M02641_0041_000000000-AGHAA | M02641 | 2015-10-21 | 75 | 75 | in date | in date | in date | 93.3815 | no | 29221047 | low | 1189.4531 | low | 0.018033 | no | 93.5119 | no | 0.96054 | 0.85 | no | 0.97325 | no | 0.9489 | no | 0.84095 | no | 0.001648 | 0.00019262 | 0 | -0.00043517 | no | 0 | -0.00032505 | no | OK | 0.42919 | yes | ['15M11968'] | [] | ['15M11968'] | NTC OK
151022_M00766_0146_000000000-AFN3B | M00766 | 2015-10-22 | 151 | 151 | in date | in date | in date | 81.4565 | no | 19862739 | low | 1107.5448 | OK | 0.031521 | no | 80.9673 | yes | 0.90849 | 0.8 | no | 0.93308 | no | 0.88275 | no | 0.95413 | no | 6.0372e-16 | 1.7055e-184 | 0 | -0.00050949 | yes | 0 | -0.00097367 | yes | OK | 0.21273 | yes | no outlying indexes |  |  | NTC OK
151023_M00766_0147_000000000-AFN37 | M00766 | 2015-10-23 | 151 | 151 | in date | in date | in date | 84.6509 | no | 20693511 | low | 1134.0694 | OK | 0.039199 | no | 85.3046 | no | 0.92707 | 0.8 | no | 0.95123 | no | 0.90247 | no | 0.95413 | no | 6.8289e-18 | nan | 0 | -0.00030046 | no | 0 | -0.00056736 | yes | OK | 0.30167 | yes | ['15M10999' '15M10996'] | ['15M10996'] | ['15M10999'] | NTC OK
151028_M02641_0043_000000000-AFL74 | M02641 | 2015-10-28 | 151 | 151 | in date | in date | in date | 96.8353 | no | 15543815 | low | 805.7896 | very low | 0.035546 | no | 96.7917 | no | 0.96903 | 0.8 | no | 0.97972 | no | 0.95847 | no | 0.67794 | no | 8.1444e-08 | 1.4081e-14 | 0 | -0.00020911 | no | 0 | -0.0004479 | no | OK | 0.15867 | yes | no outlying indexes |  |  | NTC OK
151030_M00766_0150_000000000-AFLDT | M00766 | 2015-10-30 | 151 | 151 | in date | in date | in date | 89.1258 | no | 19498823 | low | 1043.7728 | OK | 0.022316 | no | 89.1735 | no | 0.94282 | 0.8 | no | 0.95917 | no | 0.92637 | no | 0.95413 | no | 6.2812e-11 | 1.3707e-27 | 0 | -0.00027197 | no | 0 | -0.00043246 | no | OK | 0.50206 | yes | ['15M11131'] | [] | ['15M11131'] | NTC OK
151030_M02641_0044_000000000-AJHLJ | M02641 | 2015-10-30 | 75 | 75 | in date | in date | in date | 89.5684 | no | 33168068 | low | 1381.3797 | OK | 0.015021 | no | 89.5367 | no | 0.94181 | 0.85 | no | 0.96027 | no | 0.92327 | no | 0.84095 | no | 1.9212e-09 | 1.7277e-37 | 0 | -0.00053639 | yes | 0 | -0.00046394 | no | OK | 0.50917 | yes | ['15M12049'] | [] | ['15M12049'] | NTC OK
151103_M02641_0046_000000000-AJ5Y7 | M02641 | 2015-11-03 | 151 | 151 | in date | in date | in date | 96.1111 | no | 17192399 | low | 883.2398 | low | 0.021428 | no | 96.0917 | no | 0.96652 | 0.8 | no | 0.97476 | no | 0.95884 | no | 0.95413 | no | 0.00063959 | nan | 0 | -0.00027215 | no | 0 | -0.00034016 | no | OK | 0.41562 | yes | ['15M11227'] | [] | ['15M11227'] | NTC OK
151105_M02641_0047_000000000-AJF4E | M02641 | 2015-11-05 | 151 | 151 | in date | in date | in date | 96.0812 | no | 17316625 | low | 888.5653 | low | 0.024517 | no | 96.1753 | no | 0.96613 | 0.8 | no | 0.97668 | no | 0.95675 | no | 0.84095 | no | 7.3533e-09 | 2.5766e-25 | 0 | -0.00020337 | no | 0 | -0.00027243 | no | OK | 1.7725 | yes | ['15M11323'] | [] | ['15M11323'] | NTC OK
151116_M02641_0050_000000000-AJ6L3 | M02641 | 2015-11-16 | 151 | 151 | in date | in date | in date | 95.1303 | no | 13144452 | low | 672.6944 | very low | 0.074365 | yes | 94.8898 | no | 0.95687 | 0.8 | no | 0.96857 | no | 0.94432 | no | 0.95413 | no | 0.0016197 | 1.4817e-06 | 0 | -0.00029535 | no | 0 | -0.00053054 | yes | OK | 0.17711 | yes | no outlying indexes |  |  | NTC OK
151117_M00766_0155_000000000-AJ6LL | M00766 | 2015-11-17 | 151 | 151 | in date | in date | in date | 87.0005 | no | 16509092 | low | 884.2256 | low | 0.024563 | no | 86.7919 | no | 0.86377 | 0.8 | no | 0.9042 | no | 0.821 | no | 0.84095 | no | 8.628e-16 | nan | 0 | -0.0012235 | yes | 1 | -0.00063026 | yes | OK | 0.51279 | yes | ['15M11867'] | [] | ['15M11867'] | NTC OK
151117_M02641_0051_000000000-AJF4A | M02641 | 2015-11-17 | 151 | 151 | in date | in date | in date | 94.2241 | no | 17785498 | low | 905.3067 | low | 0.03067 | no | 94.2545 | no | 0.88313 | 0.8 | no | 0.91617 | no | 0.84833 | no | 0.84095 | no | 7.9592e-12 | nan | 0 | -0.0012532 | yes | 0 | -0.00072138 | yes | OK | 0.56161 | yes | ['15M12552'] | [] | ['15M12552'] | NTC OK
151125_M02641_0053_000000000-AJJ3G | M02641 | 2015-11-25 | 151 | 151 | in date | in date | in date | 96.6504 | no | 3965190 | low | 201.2481 | very low | 0.02812 | no | 96.9381 | no | 0.97423 | 0.8 | no | 0.98297 | no | 0.96916 | no | 0.67794 | no | 0.069974 | 0.0030109 | 0 | -0.00019201 | no | 0 | -0.00032396 | no | OK | 1.7978 | yes | ['15M12551'] | [] | ['15M12551'] | NTC OK
151126_M00766_0159_000000000-AJD7L | M00766 | 2015-11-26 | 75 | 75 | in date | in date | in date | 88.0477 | no | 24133209 | low | 1032.2091 | very low | 0.017289 | no | 87.5838 | no | 0.95346 | 0.85 | no | 0.96296 | no | 0.94323 | no | 0.84095 | no | 0.093829 | 1.6996e-15 | 0 | -0.0004886 | no | 0 | -0.00047968 | no | OK | 0.27264 | yes | no outlying indexes |  |  | NTC OK
151126_M02641_0054_000000000-AJD7T | M02641 | 2015-11-26 | 75 | 75 | in date | in date | in date | 94.0217 | no | 24949166 | low | 1025.4829 | very low | 0.029656 | no | 93.8083 | no | 0.96822 | 0.85 | no | 0.97532 | no | 0.96185 | no | 0.95413 | no | 0.1922 | nan | 0 | -0.00030117 | no | 0 | -0.00033189 | no | OK | 0.27535 | yes | ['15M13666'] | [] | ['15M13666'] | NTC OK
151127_M00766_0161_000000000-AJERP | M00766 | 2015-11-27 | 75 | 75 | in date | in date | in date | 90.539 | no | 24073552 | low | 1002.8265 | very low | 0.022959 | no | 90.0889 | no | 0.95458 | 0.85 | no | 0.96258 | no | 0.94625 | no | 0.95413 | no | 0.1243 | nan | 0 | -0.00057536 | yes | 0 | -0.00051622 | yes | OK | 0.56739 | yes | no outlying indexes |  |  | NTC OK
151201_M02641_0055_000000000-AJ6B2 | M02641 | 2015-12-01 | 151 | 151 | in date | in date | in date | 93.147 | no | 20354768 | low | 1062.1404 | OK | 0.019929 | no | 93.6277 | no | 0.95729 | 0.8 | no | 0.96997 | no | 0.94485 | no | 0.95413 | no | 2.0042e-09 | nan | 0 | -0.00022088 | no | 0 | -0.0003646 | no | OK | 1.8796 | yes | ['15M12551'] | [] | ['15M12551'] | NTC OK
151203_M00766_0165_000000000-AJFAA | M00766 | 2015-12-03 | 151 | 151 | in date | in date | in date | 86.1622 | no | 16734072 | low | 890.0657 | low | 0.048269 | no | 84.2361 | yes | 0.92678 | 0.8 | no | 0.94416 | no | 0.90681 | no | 0.95413 | no | 1.2134e-06 | nan | 0 | -0.0003637 | no | 0 | -0.00065096 | yes | OK | 0.47912 | yes | ['15M13694'] | ['15M13694'] | [] | NTC OK
151208_M00766_0167_000000000-AJ5JR | M00766 | 2015-12-08 | 151 | 151 | in date | in date | in date | 97.4409 | no | 5807543 | low | 293.2064 | very low | 0.036817 | no | 97.4486 | no | 0.96505 | 0.8 | no | 0.973 | no | 0.95786 | no | 0.95413 | no | 0.12937 | nan | 0 | -0.00036542 | no | 0 | -0.00043496 | no | OK | 0.48404 | yes | no outlying indexes |  |  | NTC OK
151210_M02641_0058_000000000-AJ5YJ | M02641 | 2015-12-10 | 151 | 151 | in date | in date | in date | 91.7526 | no | 18225198 | low | 951.6353 | low | 0.049608 | no | 91.4934 | no | 0.93902 | 0.8 | no | 0.95565 | no | 0.92066 | no | 0.95413 | no | 6.5736e-05 | nan | 0 | -0.00040902 | no | 0 | -0.00072147 | yes | OK | 0.29821 | yes | no outlying indexes |  |  | NTC OK
151216_M00766_0170_000000000-ALEAF | M00766 | 2015-12-16 | 151 | 151 | in date | in date | in date | 93.7279 | no | 17003866 | low | 880.7103 | low | 0.036745 | no | 93.9642 | no | 0.94991 | 0.8 | no | 0.96257 | no | 0.937 | no | 0.95413 | no | 4.5633e-06 | nan | 0 | -0.00035203 | no | 0 | -0.00057532 | yes | OK | 0.58674 | yes | ['15M13059'] | [] | ['15M13059'] | NTC OK
151217_M02641_0059_000000000-AJFAD | M02641 | 2015-12-17 | 151 | 151 | in date | in date | in date | 58.3542 | yes | 4092271 | low | 204.0891 | very low | 0.28007 | yes | 55.6594 | yes | 0.92478 | 0.8 | no | 0.94271 | no | 0.91061 | no | 0.84095 | no | 0.016952 | 0.00029975 | 0 | -0.00041346 | no | 0 | -0.00076263 | yes | OK | 0.24432 | yes | no outlying indexes |  |  | NTC OK
151223_M00766_0172_000000000-AJRMN | M00766 | 2015-12-23 | 75 | 75 | in date | in date | in date | 92.6132 | no | 23635202 | low | 977.0359 | very low | 0.01686 | no | 92.2318 | no | 0.94189 | 0.85 | no | 0.9657 | no | 0.91522 | no | 0.50766 | no | 8.2817e-14 | nan | 0 | -0.00070693 | yes | 1 | -0.0018884 | yes | OK | 0.42725 | yes | no outlying indexes |  |  | NTC OK
151223_M02641_0060_000000000-AGHV2 | M02641 | 2015-12-23 | 75 | 75 | in date | in date | in date | 87.711 | no | 26108005 | low | 1077.9707 | very low | 0.015686 | no | 87.8311 | no | 0.92979 | 0.85 | no | 0.94545 | no | 0.911 | no | 0.95413 | no | 1.0014e-06 | 6.7599e-86 | 0 | -0.00093076 | yes | 0 | -0.0011369 | yes | OK | 0.6133 | yes | ['15M14273'] | [] | ['15M14273'] | NTC OK
151231_M02641_0061_000000000-AJD8L | M02641 | 2015-12-31 | 75 | 75 | in date | in date | in date | 88.6234 | no | 30986761 | low | 1292.1541 | OK | 0.043562 | no | 89.278 | no | 0.93779 | 0.85 | no | 0.95903 | no | 0.91501 | no | 0.95413 | no | 2.22e-08 | nan | 0 | -0.00055157 | yes | 0 | -0.00062885 | yes | OK | 0.40312 | yes | ['15M14557'] | [] | ['15M14557'] | NTC OK
160107_M00766_0175_000000000-AH6C3 | M00766 | 2016-01-07 | 75 | 75 | in date | in date | in date | 87.9117 | no | 26044869 | low | 1123.4126 | low | 0.012327 | no | 86.8229 | no | 0.95101 | 0.85 | no | 0.96122 | no | 0.9395 | no | 0.95413 | no | 0.056283 | nan | 0 | -0.00056289 | yes | 0 | -0.00070024 | yes | OK | 0.41119 | yes | ['15m15138'] | [] | ['15m15138'] | NTC OK
160108_M00766_0176_000000000-AJDAJ | M00766 | 2016-01-08 | 75 | 75 | in date | in date | in date | 86.9378 | no | 24436074 | low | 1042.2918 | very low | 0.023113 | no | 87.1996 | no | 0.79786 | 0.85 | yes | 0.93046 | no | 0.6482 | yes | 0.15463 | no | 1.2164e-131 | nan | 0 | -0.0018269 | yes | 36 | -0.0074037 | yes | OK | 0.35456 | yes | ['15M15229'] | ['15M15229'] | [] | NTC OK
160114_M00766_0177_000000000-ALYTH | M00766 | 2016-01-14 | 75 | 75 | in date | in date | in date | 87.5409 | no | 27732794 | low | 1180.5507 | low | 0.016589 | no | 87.6219 | no | 0.94367 | 0.85 | no | 0.95391 | no | 0.93177 | no | 0.95413 | no | 0.043231 | nan | 0 | -0.00066231 | yes | 0 | -0.00077358 | yes | OK | 0.52375 | yes | no outlying indexes |  |  | NTC OK
160114_M02641_0065_000000000-ALRU0 | M02641 | 2016-01-14 | 75 | 75 | in date | in date | in date | 91.58 | no | 29247440 | low | 1214.9636 | OK | 0.024503 | no | 91.5468 | no | 0.95131 | 0.85 | no | 0.96723 | no | 0.93592 | no | 0.84095 | no | 7.4878e-05 | nan | 0 | -0.00049951 | no | 0 | -0.00034584 | no | OK | 0.42872 | yes | ['15M03903'] | [] | ['15M03903'] | NTC OK
160128_M02641_0070_000000000-ALAGE | M02641 | 2016-01-28 | 75 | 75 | in date | in date | in date | 91.826 | no | 24305600 | low | 1013.4527 | very low | 0.0159 | no | 91.7655 | no | 0.96174 | 0.85 | no | 0.96994 | no | 0.95369 | no | 0.95413 | no | 0.073098 | nan | 0 | -0.0003153 | no | 0 | -0.00042581 | no | OK | 0.32784 | yes | no outlying indexes |  |  | NTC OK
160129_M00766_0003_000000000-ALAH0 | M00766 | 2016-01-29 | 75 | 75 | in date | in date | in date | 86.7633 | no | 32429499 | low | 1380.9019 | OK | 0.017673 | no | 87.2761 | no | 0.93918 | 0.85 | no | 0.95507 | no | 0.92174 | no | 0.84095 | no | 1.4395e-05 | nan | 0 | -0.00058412 | yes | 0 | -0.00060062 | yes | OK | 0.36577 | yes | ['16M00136'] | [] | ['16M00136'] | NTC OK
160129_M02641_0071_000000000-AMC8N | M02641 | 2016-01-29 | 200 | 200 | in date | in date | in date | 91.7102 | no | 25390554 | low | 1054.7902 | very low | 0.01824 | no | 91.6301 | no | 0.86183 | 0.7 | no | 0.88102 | no | 0.83939 | no | 0.67794 | no | 1.3683e-14 | 6.476e-15 | 1 | -0.0017129 | yes | 0 | -0.0020671 | yes | OK | 0.14555 | yes | ['15M09614'] | ['15M09614'] | [] | NTC OK
160206_M00766_0005_000000000-ALYU1 | M00766 | 2016-02-06 | 75 | 75 | in date | in date | in date | 84.6656 | no | 32152720 | low | 1381.1846 | OK | 0.017535 | no | 83.6833 | yes | 0.93615 | 0.85 | no | 0.9506 | no | 0.92023 | no | 0.84095 | no | 0.0023058 | 1.3476e-13 | 0 | -0.00053805 | yes | 0 | -0.00042151 | no | OK | 0.27786 | yes | no outlying indexes |  |  | NTC OK
160206_M02641_0074_000000000-AL604 | M02641 | 2016-02-06 | 151 | 151 | in date | in date | in date | 88.8163 | no | 18930346 | low | 1014.3724 | OK | 0.015119 | no | 88.4642 | no | 0.92173 | 0.8 | no | 0.9483 | no | 0.89269 | no | 0.67794 | no | 2.5587e-23 | 1.4688e-132 | 0 | -0.00040622 | no | 0 | -0.00095862 | yes | High | 0.80948 | yes | ['16M00477'] | [] | ['16M00477'] | No NTC on run
160209_M00766_0007_000000000-AMERJ | M00766 | 2016-02-09 | 75 | 75 | in date | in date | in date | 79.4588 | yes | 34759702 | low | 1528.9595 | very high | 0.016677 | no | 80.1993 | yes | 0.90703 | 0.85 | no | 0.93282 | no | 0.87982 | no | 0.84095 | no | 1.1867e-08 | 0 | 0 | -0.00081183 | yes | 0 | -0.00073619 | yes | OK | 0.29288 | yes | no outlying indexes |  |  | NTC OK
160216_M00766_0009_000000000-AL73W | M00766 | 2016-02-16 | 151 | 151 | in date | in date | in date | 92.117 | no | 18300588 | low | 961.5089 | low | 0.025014 | no | 92.5591 | no | 0.94116 | 0.8 | no | 0.96156 | no | 0.91989 | no | 0.84095 | no | 1.5361e-15 | 3.496e-24 | 0 | -0.00029836 | no | 0 | -0.00066338 | yes | OK | 0.51285 | yes | ['15M13195'] | [] | ['15M13195'] | NTC OK
160219_M00766_0011_000000000-AL6GM | M00766 | 2016-02-19 | 131 | 131 | in date | in date | in date | 5.6098 | yes | 661574 | low | 22.0869 | very low | 0.9623 | yes | 9.7324 | yes | 0.77871 | 0.8 | yes | 0.79992 | yes | 0.78598 | yes | 0.99584 | no | 1 | 0.73572 | 5 | -0.00056401 | yes | 5 | -0.00059171 | yes | No index information for 160219_M00766_0011_000000000-AL6GM
160220_M02641_0081_000000000-AMETF | M02641 | 2016-02-20 | 200 | 200 | in date | in date | in date | 92.9979 | no | 23074941 | low | 958.5024 | very low | 0.027553 | no | 93.6783 | no | 0.87882 | 0.7 | no | 0.90362 | no | 0.85132 | no | 0.35842 | no | 6.4948e-26 | nan | 0 | -0.0012552 | yes | 0 | -0.0017992 | yes | OK | 0.10059 | yes | ['15M14092'] | [] | ['15M14092'] | NTC OK
160223_M00766_0012_000000000-AL5YC | M00766 | 2016-02-23 | 151 | 151 | in date | in date | in date | 93.2958 | no | 18111084 | low | 951.0894 | low | 0.019086 | no | 93.3336 | no | 0.94516 | 0.8 | no | 0.96589 | no | 0.92393 | no | 0.84095 | no | 9.6635e-16 | 3.8603e-52 | 0 | -0.0002968 | no | 0 | -0.0005917 | yes | OK | 1.8862 | yes | ['15M13471'] | [] | ['15M13471'] | NTC OK

## AutomatedResults
RunID | InstrumentID | RunDate | NumberOfCyclesR1 | NumberofCyclesR2 | FlowCellInDate? | PR2BottleInDate? | ReagentKitInDate? | %ReadsPassingFilter | %ReadsPassingFilterOutsideThreshold? | NumberofReadsPassingFilter | ReadCountWithinThreshold? | MeanClusterDensityPassingFilter | MeanClusterDensityWithinThreshold | COVClusterDensity | COVClusterDensityExceedsThreshold? | %MedianClusterDensityPassingFilter | GapBetweenClusterDensitiesPassingFilterAndNotExceedsThreshold? | ProportionOverQ30 | IlluminaQ30Threshold | ProportionOverQ30ExceedsThreshold? | ProportionOverQ30Read1 | ProportionOverQ30ExceedsThresholdRead1 | ProportionOverQ30Read2 | ProportionOverQ30ExceedsThresholdRead2 | K-SPValue | K-SPValueSignificant? | K-SUnpackPvalue | M-WUnpackPvalue | NumberOfCyclesExceedingThresholdUnderQ30Read1 | SlopeOfLineOfBestFitRead1 | SlopeOfLineOfBestFitRead1ExceedsThreshold? | NumberOfCyclesExceedingThresholdUnderQ30Read2 | SlopeOfLineOfBestFitRead2 | SlopeOfLineOfBestFitRead2ExceedsThreshold? | ProportionOfUndeterminedReads? | COVIndexes | COVIndexesExceedsThreshold? | OutlyingIndexes | OutlyingIndexes(Low) | OutlyingIndexes(High) | StateOfNTC
121009_M00766_0002_000000000-A1U6P | M00766 | 2012-10-09 | 151 | 151 | in date | in date | in date | 45.385 | yes | 3138050 | low | 341.3139 | low | 0.34653 | yes | 0 | yes | 0.87575 | 0.8 | no | 0.87642 | no | 0.87507 | no | 0.99998 | no | 0.99997 | 0.80608 | 1 | 0 | no | 0 | 0 | no | No index information for 121009_M00766_0002_000000000-A1U6P
130206_M00766_0002_000000000-A23JM | M00766 | 2013-02-06 | 250 | 250 | in date | in date | in date | 2.0076 | yes | 160923 | low | 10.2917 | very low | 0.26427 | yes | 1.9741 | yes | 0.27654 | 0.75 | yes | 0.091188 | yes | 0.46189 | yes | 0.35842 | no | 0.0011204 | 0.012662 | 193 | -0.0010965 | yes | 65 | -0.0032351 | yes | No index information for 130206_M00766_0002_000000000-A23JM
130618_M00766_0019_000000000-A4FEU | M00766 | 2013-06-18 | 251 | 251 | in date | in date | in date | 93.703 | no | 10147864 | low | 556.7555 | very low | 0.009214 | no | 93.9257 | no | 0.79545 | 0.75 | no | 0.90383 | no | 0.68024 | yes | 0.35842 | no | 1.6395e-83 | 9.0394e-306 | 0 | -0.0010825 | yes | 47 | -0.0037713 | yes | High | 0.1969 | yes | no outlying indexes |  |  | No NTC on run
130624_M00766_0021_000000000-A53PT | M00766 | 2013-06-24 | 251 | 251 | in date | in date | in date | 88.7381 | no | 17725761 | low | 940.8335 | OK | 0.011629 | no | 89.0017 | no | 0.84736 | 0.75 | no | 0.90275 | no | 0.78576 | no | 0.67794 | no | 1.0912e-52 | 1.2366e-63 | 0 | -0.0009554 | yes | 4 | -0.0019889 | yes | OK | 0.04822 | no | no outlying indexes |  |  | No NTC on run
130712_M00766_0024_000000000-A5AVV | M00766 | 2013-07-12 | 151 | 151 | in date | in date | in date | 82.4401 | no | 13180142 | low | 710.2954 | low | 0.09422 | yes | 82.1308 | no | 0.92195 | 0.8 | no | 0.92377 | no | 0.92292 | no | 0.99998 | no | 0.0096055 | 4.6381e-11 | 0 | -0.00059232 | yes | 0 | -0.00062141 | yes | High | 0.20831 | yes | no outlying indexes |  |  | No NTC on run
130817_M00766_0037_000000000-A4EDD | M00766 | 2013-08-17 | 176 | 176 | in date | in date | in date | 93.613 | no | 14493580 | low | 748.7721 | low | 0.020101 | no | 93.6127 | no | 0.94897 | 0.75 | no | 0.96312 | no | 0.93388 | no | 0.84095 | no | 4.6236e-05 | nan | 0 | -0.00039473 | no | 1 | -0.0006017 | yes | OK | 0.4371 | yes | ['50005-d' '60003-d' '112001-f' '50005-f'] | ['50005-d' '50005-f'] | ['60003-d' '112001-f'] | No NTC on run
130913_M00766_0049_000000000-A5B0E | M00766 | 2013-09-13 | 251 | 251 | in date | in date | in date | 89.8636 | no | 17433281 | low | 913.7303 | OK | 0.024696 | no | 89.1779 | no | 0.83393 | 0.75 | no | 0.88573 | no | 0.77648 | no | 0.84095 | no | 3.8371e-42 | 5.6819e-66 | 0 | -0.0011656 | yes | 4 | -0.0019759 | yes | High | 0.020958 | no | no outlying indexes |  |  | No NTC on run
130916_M00766_0050_000000000-A5M1U | M00766 | 2013-09-16 | 151 | 151 | in date | in date | in date | 86.0618 | no | 22033351 | low | 1185.4173 | very high | 0.028468 | no | 86.0122 | no | 0.9138 | 0.8 | no | 0.93477 | no | 0.89056 | no | 0.84095 | no | 4.3159e-13 | 1.1035e-23 | 0 | -0.00070235 | yes | 0 | -0.0009045 | yes | OK | 0.18994 | yes | ['12M04425'] | ['12M04425'] | [] | No NTC on run
130924_M00766_0052_000000000-A5BE6 | M00766 | 2013-09-24 | 251 | 251 | in date | in date | in date | 67.1952 | yes | 14006895 | low | 805.1388 | high | 0.092176 | yes | 66.8443 | yes | 0.71517 | 0.75 | yes | 0.77069 | no | 0.66348 | yes | 0.99584 | no | 6.3637e-20 | nan | 16 | -0.0019789 | yes | 32 | -0.002186 | yes | High | 0.49178 | yes | no outlying indexes |  |  | No NTC on run
131007_M00766_0055_000000000-A59JT | M00766 | 2013-10-07 | 251 | 251 | in date | in date | in date | 88.9869 | no | 7066429 | low | 464.0561 | very low | 0.11395 | yes | 90.7266 | no | 0.86005 | 0.75 | no | 0.86612 | no | 0.85247 | no | 1 | no | 0.071573 | nan | 1 | -0.0014022 | yes | 1 | -0.0015446 | yes | High | 0.021021 | no | no outlying indexes |  |  | No NTC on run
131025_M00766_0059_000000000-A5P9A | M00766 | 2013-10-25 | 151 | 151 | in date | in date | in date | 90.6341 | no | 17754371 | low | 930.3591 | OK | 0.030836 | no | 91.4036 | no | 0.92242 | 0.8 | no | 0.93549 | no | 0.90546 | no | 0.99584 | no | 1.9427e-05 | nan | 0 | -0.00094972 | yes | 0 | -0.0010274 | yes | OK | 0.51483 | yes | no outlying indexes |  |  | No NTC on run
140207_M00766_0022_000000000-A7PH8 | M00766 | 2014-02-07 | 251 | 251 | in date | in date | in date | 88.3301 | no | 10785103 | low | 558.1003 | low | 0.053446 | no | 89.8657 | no | 0.85142 | 0.75 | no | 0.8548 | no | 0.84338 | no | 1 | no | 0.42818 | nan | 1 | -0.0014636 | yes | 1 | -0.0016776 | yes | High | 0.3944 | yes | no outlying indexes |  |  | No NTC on run
140424_M00766_0033_000000000-A7BNA | M00766 | 2014-04-24 | 151 | 151 | in date | in date | in date | 97.0072 | no | 5094671 | low | 251.0711 | very low | 0.231 | yes | 98.5469 | no | 0.93672 | 0.8 | no | 0.94645 | no | 0.92545 | no | 0.95413 | no | 0.083102 | 0.0055314 | 0 | -0.00096088 | yes | 0 | -0.0010359 | yes | OK | 0.49796 | yes | ['10M1732'] | [] | ['10M1732'] | NTC OK
140514_M00766_0036_000000000-A7BRK | M00766 | 2014-05-14 | 151 | 151 | in date | in date | in date | 81.4684 | no | 20334208 | low | 1124.4827 | very high | 0.017101 | no | 81.2844 | no | 0.79665 | 0.8 | yes | 0.83677 | no | 0.75502 | yes | 0.99584 | no | 4.1492e-14 | nan | 0 | -0.0024166 | yes | 7 | -0.0018547 | yes | High | 0.51173 | yes | no outlying indexes |  |  | No NTC on run
140612_M00766_0039_000000000-A7BNL | M00766 | 2014-06-12 | 251 | 251 | in date | in date | in date | 91.373 | no | 20643048 | low | 1088.1751 | OK | 0.014529 | no | 90.3695 | no | 0.78766 | 0.75 | no | 0.82619 | no | 0.74426 | yes | 0.84095 | no | 1.1316e-41 | nan | 19 | -0.0022137 | yes | 39 | -0.0026621 | yes | OK | 0.11915 | yes | no outlying indexes |  |  | NTC OK
140616_M00766_0040_000000000-A78V9 | M00766 | 2014-06-16 | 251 | 251 | in date | in date | in date | 97.3311 | no | 10710545 | low | 534.5911 | very low | 0.0065097 | no | 97.505 | no | 0.81498 | 0.75 | no | 0.8406 | no | 0.7849 | no | 0.95413 | no | 7.2918e-10 | nan | 1 | -0.0024062 | yes | 30 | -0.0029712 | yes | OK | 0.087674 | yes | no outlying indexes |  |  | NTC OK
140813_M00766_0047_000000000-A8PJL | M00766 | 2014-08-13 | 151 | 151 | in date | in date | in date | 91.9352 | no | 17936069 | low | 943.905 | OK | 0.026213 | no | 91.9798 | no | 0.92465 | 0.8 | no | 0.94352 | no | 0.90463 | no | 0.95413 | no | 3.3998e-09 | 3.0143e-185 | 0 | -0.00068224 | yes | 0 | -0.00075951 | yes | OK | 0.35465 | yes | ['11M08823'] | ['11M08823'] | [] | NTC OK
140901_M00766_0048_000000000-AA63M | M00766 | 2014-09-01 | 151 | 151 | in date | in date | in date | 94.0673 | no | 17640397 | low | 909.8289 | low | 0.02007 | no | 93.8944 | no | 0.93862 | 0.8 | no | 0.96072 | no | 0.91538 | no | 0.84095 | no | 4.2142e-16 | 6.6658e-224 | 0 | -0.00044115 | no | 0 | -0.00052041 | yes | OK | 0.22643 | yes | ['14M05256'] | [] | ['14M05256'] | NTC OK
141014_M00766_0053_000000000-A8R6M | M00766 | 2014-10-14 | 151 | 151 | in date | in date | in date | 89.6387 | no | 19084469 | low | 1011.4321 | OK | 0.014006 | no | 90.0342 | no | 0.89321 | 0.8 | no | 0.91865 | no | 0.86561 | no | 0.95413 | no | 8.0153e-17 | 2.8333e-171 | 0 | -0.0010561 | yes | 0 | -0.0011248 | yes | OK | 0.10441 | yes | no outlying indexes |  |  | NTC OK
141017_M02641_0022_000000000-AA66H | M02641 | 2014-10-17 | 151 | 151 | in date | in date | in date | 79.2948 | no | 22430099 | low | 1210.1492 | very high | 0.017994 | no | 78.9705 | yes | 0.87741 | 0.8 | no | 0.91601 | no | 0.83648 | no | 0.99584 | no | 1.1016e-30 | nan | 0 | -0.00063613 | yes | 0 | -0.00060429 | yes | OK | 0.10826 | yes | ['14M02681'] | ['14M02681'] | [] | NTC OK
141024_M00766_0056_000000000-A8PC5 | M00766 | 2014-10-24 | 151 | 151 | in date | in date | in date | 86.6935 | no | 22517755 | low | 1216.7323 | very high | 0.017442 | no | 86.7307 | no | 0.88873 | 0.8 | no | 0.92881 | no | 0.84641 | no | 0.84095 | no | 1.3784e-37 | 2.2642e-64 | 0 | -0.00069708 | yes | 0 | -0.00081119 | yes | High | 0.51078 | yes | no outlying indexes |  |  | NTC OK
141031_M00766_0058_000000000-AA3GN | M00766 | 2014-10-31 | 75 | 75 | in date | in date | in date | 84.8049 | no | 22830617 | low | 1020.36 | OK | 0.066963 | no | 85.4852 | no | 0.94188 | 0.85 | no | 0.94357 | no | 0.93798 | no | 0.99584 | no | 0.16727 | nan | 0 | -0.00072598 | yes | 0 | -0.0003319 | no | OK | 0.42485 | yes | no outlying indexes |  |  | NTC OK
141107_M00766_0060_000000000-AA8PM | M00766 | 2014-11-07 | 151 | 151 | in date | in date | in date | 90.0798 | no | 11262793 | low | 579.4485 | low | 0.034966 | no | 90.2121 | no | 0.96028 | 0.8 | no | 0.96678 | no | 0.9539 | no | 0.99584 | no | 0.99203 | 0.8978 | 0 | -0.00013806 | no | 0 | -0.00011747 | no | OK | 0.22736 | yes | ['14M08238' '14M09145' '14M09855'] | ['14M08238' '14M09145' '14M09855'] | [] | NTC OK
141118_M00766_0061_000000000-A8P8J | M00766 | 2014-11-18 | 151 | 151 | in date | in date | in date | 92.7322 | no | 19181201 | low | 1009.9074 | OK | 0.027961 | no | 92.312 | no | 0.93572 | 0.8 | no | 0.94773 | no | 0.92218 | no | 0.95413 | no | 1.3904e-05 | nan | 0 | -0.0006231 | yes | 0 | -0.00072415 | yes | OK | 0.13191 | yes | ['14M07799'] | ['14M07799'] | [] | NTC OK
141208_M00766_0063_000000000-A8P7C | M00766 | 2014-12-08 | 151 | 151 | in date | in date | in date | 90.051 | no | 21795042 | low | 1153.0041 | high | 0.020941 | no | 89.941 | no | 0.89826 | 0.8 | no | 0.94464 | no | 0.85869 | no | 0.67794 | no | 1.8648e-37 | nan | 0 | -0.00060048 | yes | 3 | -0.00051121 | yes | OK | 0.32891 | yes | no outlying indexes |  |  | NTC OK
141208_M02641_0026_000000000-A8R55 | M02641 | 2014-12-08 | 151 | 151 | in date | in date | in date | 94.9912 | no | 13800416 | low | 717.5691 | low | 0.027596 | no | 95.1078 | no | 0.84608 | 0.8 | no | 0.94004 | no | 0.74763 | yes | 0.35842 | no | 7.7537e-147 | 6.3912e-157 | 0 | -0.00078235 | yes | 5 | -0.0011709 | yes | OK | 0.52447 | yes | no outlying indexes |  |  | NTC OK
141212_M00766_0064_000000000-ACCEB | M00766 | 2014-12-12 | 151 | 151 | in date | in date | in date | 81.1933 | no | 20826514 | low | 1127.546 | very high | 0.021875 | no | 81.0162 | no | 0.74058 | 0.8 | yes | 0.80816 | no | 0.67993 | yes | 0.99584 | no | 3.8905e-34 | nan | 0 | -0.0028592 | yes | 24 | -0.0015004 | yes | High | 0.28308 | yes | ['6'] | ['6'] | [] | NTC OK
141216_M00766_0065_000000000-ACCDT | M00766 | 2014-12-16 | 151 | 151 | in date | in date | in date | 79.0858 | no | 20548879 | low | 1120.7265 | very high | 0.02252 | no | 79.0848 | yes | 0.77398 | 0.8 | yes | 0.84162 | no | 0.70405 | yes | 0.84095 | no | 1.4032e-49 | 7.9356e-94 | 0 | -0.0021868 | yes | 16 | -0.00079947 | yes | OK | 0.55632 | yes | ['13M11854'] | [] | ['13M11854'] | NTC OK
141219_M00766_0066_000000000-ACCB1 | M00766 | 2014-12-19 | 151 | 151 | in date | in date | in date | 84.1075 | no | 16554828 | low | 882.4454 | OK | 0.030188 | no | 83.544 | no | 0.71631 | 0.8 | yes | 0.77794 | yes | 0.66061 | yes | 0.99998 | no | 9.0168e-19 | 1.7886e-30 | 0 | -0.0034797 | yes | 18 | -0.0018779 | yes | High | 0.8782 | yes | ['13M12119'] | [] | ['13M12119'] | NTC OK
150120_M00766_0071_000000000-AA63K | M00766 | 2015-01-20 | 151 | 151 | in date | in date | in date | 85.9694 | no | 14794255 | low | 793.1929 | low | 0.05599 | no | 87.4592 | no | 0.90779 | 0.8 | no | 0.93487 | no | 0.89285 | no | 0.95413 | no | 0.01591 | nan | 0 | -0.00060834 | yes | 3 | -0.00064423 | yes | OK | 0.44895 | yes | ['9M4318'] | [] | ['9M4318'] | No NTC on run
150127_M02641_0027_000000000-AA65J | M02641 | 2015-01-27 | 151 | 151 | in date | in date | out of date | 94.4613 | no | 15926233 | low | 820.6824 | low | 0.035198 | no | 94.2959 | no | 0.93014 | 0.8 | no | 0.95953 | no | 0.9008 | no | 0.67794 | no | 1.0379e-16 | 3.674e-28 | 0 | -0.0004646 | no | 0 | -0.00058277 | yes | OK | 0.41652 | yes | no outlying indexes |  |  | NTC OK
150203_M00766_0074_000000000-AAUMH | M00766 | 2015-02-03 | 151 | 151 | in date | in date | in date | 81.6068 | no | 16649356 | low | 884.3212 | OK | 0.052428 | no | 81.2451 | no | 0.90342 | 0.8 | no | 0.92886 | no | 0.88527 | no | 0.95413 | no | 0.0090932 | 0.028862 | 0 | -0.00047881 | no | 3 | -0.00020141 | no | OK | 0.3567 | yes | ['13M70724' '13M70411' '13M70386' '13M70429' '13M70467'] | ['13M70724' '13M70386'] | ['13M70411' '13M70429' '13M70467'] | NTC OK
150203_M02641_0028_000000000-ACBYG | M02641 | 2015-02-03 | 151 | 151 | in date | in date | in date | 89.643 | no | 18845203 | low | 988.6001 | OK | 0.014743 | no | 90.0779 | no | 0.82841 | 0.8 | no | 0.86795 | no | 0.78849 | yes | 0.99584 | no | 1.2512e-16 | 0 | 0 | -0.0022305 | yes | 3 | -0.0018175 | yes | OK | 0.3176 | yes | ['13M12250' '15M00579'] | ['15M00579'] | ['13M12250'] | NTC OK
150205_M00766_0075_000000000-ACC43 | M00766 | 2015-02-05 | 151 | 151 | in date | in date | in date | 93.254 | no | 17676890 | low | 914.0169 | low | 0.038617 | no | 93.3074 | no | 0.8595 | 0.8 | no | 0.89452 | no | 0.83021 | no | 0.95413 | no | 5.2726e-15 | 3.8357e-55 | 0 | -0.0019394 | yes | 4 | -0.0017318 | yes | High | 0.58276 | yes | no outlying indexes |  |  | NTC OK
150224_M00766_0078_000000000-ACMP5 | M00766 | 2015-02-24 | 75 | 75 | in date | in date | in date | 82.4896 | no | 36080266 | low | 1546.85 | very high | 0.024283 | no | 82.282 | no | 0.93168 | 0.85 | no | 0.94868 | no | 0.91268 | no | 0.84095 | no | 1.1086e-05 | 0 | 0 | -0.00047772 | no | 0 | -0.00047646 | no | OK | 0.22071 | yes | ['14M09093'] | [] | ['14M09093'] | NTC OK
150309_M02641_0035_000000000-ACC3J | M02641 | 2015-03-09 | 151 | 151 | in date | in date | in date | 88.8745 | no | 22186253 | low | 1157.1795 | very high | 0.017088 | no | 89.0446 | no | 0.79385 | 0.8 | yes | 0.83769 | no | 0.74848 | yes | 0.99584 | no | 2.3483e-15 | nan | 0 | -0.002745 | yes | 7 | -0.0021184 | yes | High | 0.30141 | yes | ['15M00579'] | ['15M00579'] | [] | NTC OK
150311_M00766_0081_000000000-ACN0W | M00766 | 2015-03-11 | 75 | 75 | in date | in date | in date | 92.0795 | no | 28700881 | low | 1177.6318 | OK | 0.016158 | no | 92.5321 | no | 0.96191 | 0.85 | no | 0.9725 | no | 0.95082 | no | 0.84095 | no | 0.010916 | nan | 0 | -0.00034381 | no | 0 | -0.00028909 | no | OK | 0.30879 | yes | no outlying indexes |  |  | NTC OK
150311_M02641_0036_000000000-AD8VB | M02641 | 2015-03-11 | 75 | 75 | in date | in date | in date | 94.0269 | no | 26337766 | low | 1071.9868 | low | 0.027519 | no | 93.3836 | no | 0.95837 | 0.85 | no | 0.9654 | no | 0.95151 | no | 0.95413 | no | 0.49156 | nan | 0 | -0.0006745 | yes | 0 | -0.00014209 | no | OK | 0.38297 | yes | ['14M07894'] | [] | ['14M07894'] | NTC OK
150402_M00766_0087_000000000-ACCAY | M00766 | 2015-04-02 | 151 | 151 | in date | in date | in date | 88.6379 | no | 17224632 | low | 896.3347 | OK | 0.068342 | no | 91.2709 | no | 0.85166 | 0.8 | no | 0.88173 | no | 0.81998 | no | 0.95413 | no | 2.2205e-10 | nan | 0 | -0.0022033 | yes | 5 | -0.0021469 | yes | OK | 0.4939 | yes | no outlying indexes |  |  | NTC OK
150414_M00766_0090_000000000-AAU39 | M00766 | 2015-04-14 | 151 | 151 | in date | in date | in date | 86.0407 | no | 14841878 | low | 771.5068 | low | 0.03988 | no | 87.1861 | no | 0.93627 | 0.8 | no | 0.95058 | no | 0.9214 | no | 0.95413 | no | 0.27827 | 0.21791 | 0 | -0.00025771 | no | 0 | -0.00046336 | no | OK | 0.16948 | yes | ['13M70646'] | ['13M70646'] | [] | NTC OK
150427_M02641_0042_000000000-AD8LB | M02641 | 2015-04-27 | 75 | 75 | in date | in date | in date | 97.4766 | no | 4116334 | low | 164.4147 | very low | 0.057123 | no | 97.3182 | no | 0.95813 | 0.85 | no | 0.97251 | no | 0.95269 | no | 0.95413 | no | 1 | 0.43484 | 0 | -0.00032162 | no | 0 | -0.00011879 | no | High | 0.47354 | yes | no outlying indexes |  |  | NTC OK
150430_M00766_0092_000000000-AD77P | M00766 | 2015-04-30 | 75 | 75 | in date | in date | in date | 86.4782 | no | 33656164 | low | 1420.3657 | high | 0.029215 | no | 86.0606 | no | 0.94127 | 0.85 | no | 0.95639 | no | 0.92495 | no | 0.95413 | no | 2.1517e-05 | 1.5334e-07 | 0 | -0.00053145 | yes | 0 | -0.00060527 | yes | OK | 0.32903 | yes | no outlying indexes |  |  | NTC OK
150506_M00766_0094_000000000-AEVP8 | M00766 | 2015-05-06 | 151 | 151 | in date | in date | in date | 89.4984 | no | 20688072 | low | 1090.2056 | high | 0.029888 | no | 89.7217 | no | 0.92458 | 0.8 | no | 0.94568 | no | 0.90199 | no | 0.84095 | no | 4.8432e-13 | nan | 0 | -0.00051215 | yes | 0 | -0.00090332 | yes | OK | 0.23251 | yes | no outlying indexes |  |  | NTC OK
150508_M00766_0096_000000000-AF9MW | M00766 | 2015-05-08 | 75 | 75 | in date | in date | in date | 91.4565 | no | 29237117 | low | 1200.3221 | OK | 0.049362 | no | 92.6001 | no | 0.95684 | 0.85 | no | 0.97022 | no | 0.94333 | no | 0.84095 | no | 0.0058945 | nan | 0 | -0.00039203 | no | 0 | -0.00040344 | no | OK | 0.37127 | yes | no outlying indexes |  |  | NTC OK
150508_M02641_0047_000000000-AF7NG | M02641 | 2015-05-08 | 75 | 75 | in date | in date | in date | 89.2921 | no | 33452889 | low | 1376.6665 | high | 0.032357 | no | 90.4595 | no | 0.94491 | 0.85 | no | 0.96109 | no | 0.92874 | no | 0.84095 | no | 5.3189e-05 | nan | 0 | -0.00048487 | no | 0 | -0.00044495 | no | OK | 0.37915 | yes | no outlying indexes |  |  | NTC OK
150511_M02641_0048_000000000-AEUGT | M02641 | 2015-05-11 | 250 | 250 | in date | in date | in date | 25.8223 | yes | 3140165 | low | 125.9118 | very low | 0.11338 | yes | 14.4297 | yes | 0.53707 | 0.7 | yes | 0.38185 | yes | 0.6802 | yes | 0.35842 | no | 2.3353e-41 | 4.6079e-40 | 205 | -0.0023104 | yes | 25 | -0.0033967 | yes | OK | 0.17177 | yes | ['14M12235'] | [] | ['14M12235'] | NTC OK
150516_M00766_0099_000000000-AF81F | M00766 | 2015-05-16 | 75 | 75 | in date | in date | in date | 92.0597 | no | 29948630 | low | 1238.4292 | OK | 0.017269 | no | 91.8305 | no | 0.96082 | 0.85 | no | 0.97249 | no | 0.94905 | no | 0.84095 | no | 0.0018358 | 5.1652e-05 | 0 | -0.00036272 | no | 0 | -0.00037743 | no | OK | 0.30556 | yes | no outlying indexes |  |  | NTC OK
150521_M02641_0051_000000000-AF9JH | M02641 | 2015-05-21 | 75 | 75 | in date | in date | in date | 85.4273 | no | 20774323 | low | 847.1236 | low | 0.093549 | yes | 85.9442 | no | 0.92414 | 0.85 | no | 0.93171 | no | 0.91686 | no | 0.99584 | no | 0.99485 | 0.61684 | 0 | -0.00098979 | yes | 0 | -0.00048486 | no | OK | 0.40967 | yes | no outlying indexes |  |  | NTC OK
150522_M00766_0103_000000000-AF9M3 | M00766 | 2015-05-22 | 75 | 75 | in date | in date | in date | 88.2546 | no | 33587117 | low | 1405.6196 | high | 0.026305 | no | 88.4186 | no | 0.94466 | 0.85 | no | 0.95853 | no | 0.92962 | no | 0.84095 | no | 0.0015346 | nan | 0 | -0.00056188 | yes | 0 | -0.00050146 | yes | OK | 0.39895 | yes | no outlying indexes |  |  | NTC OK
150522_M02641_0052_000000000-AF82F | M02641 | 2015-05-22 | 75 | 75 | in date | in date | in date | 77.7302 | no | 26557402 | low | 1103.8322 | high | 0.045966 | no | 76.9724 | yes | 0.90712 | 0.85 | no | 0.92279 | no | 0.89107 | no | 0.99584 | no | 0.060099 | 5.5163e-09 | 0 | -0.00079849 | yes | 0 | -0.00064071 | yes | OK | 0.36661 | yes | ['15M05218'] | [] | ['15M05218'] | NTC OK
150601_M00766_0107_000000000-AF9N8 | M00766 | 2015-06-01 | 75 | 75 | in date | in date | in date | 96.709 | no | 17833767 | low | 714.0694 | very low | 0.036487 | no | 97.1935 | no | 0.97916 | 0.85 | no | 0.98619 | no | 0.97361 | no | 0.84095 | no | 0.47307 | nan | 0 | -0.00024948 | no | 0 | -0.00016826 | no | OK | 0.37843 | yes | ['15M05723'] | ['15M05723'] | [] | NTC OK
150602_M00766_0108_000000000-AFMB8 | M00766 | 2015-06-02 | 200 | 200 | in date | in date | in date | 87.6065 | no | 33648503 | low | 1419.0921 | high | 0.017208 | no | 87.2266 | no | 0.8224 | 0.7 | no | 0.84731 | no | 0.7928 | no | 0.84095 | no | 3.1434e-29 | 0 | 19 | -0.0019091 | yes | 33 | -0.0023749 | yes | OK | 0.36723 | yes | ['14M12202'] | [] | ['14M12202'] | NTC OK
150603_M02641_0056_000000000-AEY8R | M02641 | 2015-06-03 | 151 | 151 | in date | in date | in date | 75.779 | no | 10182405 | low | 516.5408 | low | 0.052494 | no | 76.2721 | yes | 0.92128 | 0.8 | no | 0.93202 | no | 0.91084 | no | 0.99584 | no | 0.72835 | nan | 0 | -0.00018443 | no | 0 | -0.00018378 | no | OK | 0.059979 | yes | ['15M05343' '15M05476-2'] | ['15M05476-2'] | ['15M05343'] | NTC OK
150604_M00766_0109_000000000-AFJ62 | M00766 | 2015-06-04 | 75 | 75 | in date | in date | in date | 90.2876 | no | 30654786 | low | 1283.025 | high | 0.019566 | no | 90.5316 | no | 0.94854 | 0.85 | no | 0.96577 | no | 0.93018 | no | 0.84095 | no | 6.5656e-06 | nan | 0 | -0.0004699 | no | 0 | -0.00054877 | yes | OK | 0.42556 | yes | no outlying indexes |  |  | NTC OK
150604_M02641_0057_000000000-AFL7N | M02641 | 2015-06-04 | 75 | 75 | in date | in date | in date | 83.6851 | no | 20989024 | low | 846.0443 | low | 0.089565 | yes | 87.1228 | no | 0.92879 | 0.85 | no | 0.94412 | no | 0.91514 | no | 0.84095 | no | 0.077245 | nan | 0 | -0.00064783 | yes | 0 | -0.00048886 | no | OK | 0.76195 | yes | no outlying indexes |  |  | NTC OK
150605_M00766_0110_000000000-AF80F | M00766 | 2015-06-05 | 75 | 75 | in date | in date | in date | 83.8601 | no | 32100609 | low | 1379.9375 | high | 0.019606 | no | 84.3276 | no | 0.92414 | 0.85 | no | 0.94606 | no | 0.89846 | no | 0.84095 | no | 2.0225e-07 | nan | 0 | -0.00071066 | yes | 0 | -0.00077961 | yes | OK | 1.153 | yes | ['15M01355'] | [] | ['15M01355'] | NTC OK
150610_M00766_0111_000000000-AFMWF | M00766 | 2015-06-10 | 75 | 75 | in date | in date | in date | 89.8899 | no | 32008549 | low | 1328.1731 | high | 0.015813 | no | 89.8802 | no | 0.94912 | 0.85 | no | 0.96373 | no | 0.93447 | no | 0.84095 | no | 0.0020072 | nan | 0 | -0.00052315 | yes | 0 | -0.00050342 | yes | OK | 0.42926 | yes | ['14M10378'] | [] | ['14M10378'] | NTC OK
150611_M02641_0060_000000000-AFN2Y | M02641 | 2015-06-11 | 75 | 75 | in date | in date | in date | 76.6898 | no | 23799396 | low | 947.2643 | OK | 0.054797 | no | 75.8247 | yes | 0.88682 | 0.85 | no | 0.90551 | no | 0.86738 | no | 0.99584 | no | 0.046537 | 2.6959e-66 | 0 | -0.001158 | yes | 0 | -0.00077535 | yes | OK | 0.39558 | yes | no outlying indexes |  |  | NTC OK
150615_M02641_0062_000000000-AFH7C | M02641 | 2015-06-15 | 75 | 75 | in date | in date | in date | 76.7102 | no | 21607117 | low | 847.9618 | low | 0.051033 | no | 78.099 | yes | 0.89692 | 0.85 | no | 0.91029 | no | 0.88432 | no | 0.99584 | no | 0.37812 | 0.020859 | 0 | -0.00092944 | yes | 0 | -0.00061072 | yes | OK | 0.23706 | yes | ['15M04473'] | [] | ['15M04473'] | NTC OK
150619_M02641_0063_000000000-AFN2W | M02641 | 2015-06-19 | 75 | 75 | in date | in date | in date | 86.313 | no | 22890237 | low | 921.1523 | low | 0.031622 | no | 86.5093 | no | 0.93054 | 0.85 | no | 0.94446 | no | 0.91735 | no | 0.95413 | no | 0.013174 | nan | 0 | -0.00066355 | yes | 0 | -0.0006793 | yes | OK | 0.37302 | yes | no outlying indexes |  |  | NTC OK
150625_M02641_0065_000000000-AEY2A | M02641 | 2015-06-25 | 151 | 151 | in date | in date | in date | 67.3086 | yes | 7123794 | low | 354.4358 | very low | 0.027403 | no | 65.5091 | yes | 0.90035 | 0.8 | no | 0.9059 | no | 0.89467 | no | 0.99998 | no | 0.37708 | 0.0069801 | 0 | -0.00018719 | no | 0 | -0.00016475 | no | OK | 0.2852 | yes | ['15M06230-2' '15M06230' '15M06510-2'] | ['15M06230-2' '15M06230' '15M06510-2'] | [] | NTC OK
150703_M00766_0117_000000000-AFN2F | M00766 | 2015-07-03 | 75 | 75 | in date | in date | in date | 89.9328 | no | 28300840 | low | 1173.6259 | OK | 0.036091 | no | 90.4158 | no | 0.94919 | 0.85 | no | 0.95814 | no | 0.93984 | no | 0.84095 | no | 0.81385 | nan | 0 | -0.00062428 | yes | 0 | -0.00041123 | no | OK | 0.27389 | yes | ['15M06941'] | [] | ['15M06941'] | NTC OK
150703_M02641_0001_000000000-AFMW2 | M02641 | 2015-07-03 | 75 | 75 | in date | in date | in date | 92.3482 | no | 29421610 | low | 1208.9627 | OK | 0.019251 | no | 92.6367 | no | 0.95661 | 0.85 | no | 0.96958 | no | 0.94384 | no | 0.84095 | no | 0.001929 | 2.1798e-15 | 0 | -0.0004435 | no | 0 | -0.00035874 | no | OK | 0.64377 | yes | no outlying indexes |  |  | NTC OK
150709_M02641_0004_000000000-AFMPC | M02641 | 2015-07-09 | 75 | 75 | in date | in date | in date | 94.2018 | no | 27443669 | low | 1115.769 | low | 0.036555 | no | 95.4991 | no | 0.96502 | 0.85 | no | 0.9763 | no | 0.95448 | no | 0.84095 | no | 0.016984 | nan | 0 | -0.00039431 | no | 0 | -0.00024497 | no | OK | 0.38492 | yes | no outlying indexes |  |  | NTC OK
150727_M00766_0119_000000000-AF3TM | M00766 | 2015-07-27 | 151 | 151 | in date | in date | in date | 88.8616 | no | 15391937 | low | 825.3634 | low | 0.055595 | no | 88.285 | no | 0.9196 | 0.8 | no | 0.9501 | no | 0.88668 | no | 0.84095 | no | 1.2664e-16 | nan | 0 | -0.00038471 | no | 0 | -0.00086889 | yes | OK | 0.60581 | yes | no outlying indexes |  |  | NTC OK
150727_M02641_0007_000000000-AGEW7 | M02641 | 2015-07-27 | 230 | 230 | in date | in date | in date | 91.5639 | no | 31830282 | low | 1309.0338 | high | 0.015312 | no | 91.8906 | no | 0.76141 | 0.7 | no | 0.80125 | no | 0.71555 | no | 0.35842 | no | 5.1559e-54 | 2.0554e-121 | 29 | -0.0026602 | yes | 56 | -0.0035322 | yes | OK | 0.33395 | yes | ['15M06635'] | [] | ['15M06635'] | NTC OK
150807_M00766_0122_000000000-AFMW7 | M00766 | 2015-08-07 | 75 | 75 | in date | in date | in date | 95.8884 | no | 19129362 | low | 771.607 | low | 0.031184 | no | 96.3208 | no | 0.9746 | 0.85 | no | 0.98266 | no | 0.96788 | no | 0.84095 | no | 0.16299 | nan | 0 | -0.00031347 | no | 0 | -0.0002468 | no | OK | 0.36606 | yes | no outlying indexes |  |  | NTC OK
150813_M00766_0124_000000000-AFYEH | M00766 | 2015-08-13 | 75 | 75 | in date | in date | in date | 94.3227 | no | 25145640 | low | 1025.335 | low | 0.021357 | no | 94.2549 | no | 0.96769 | 0.85 | no | 0.97825 | no | 0.95812 | no | 0.84095 | no | 0.0084968 | nan | 0 | -0.00033121 | no | 0 | -0.00034652 | no | OK | 0.23439 | yes | no outlying indexes |  |  | NTC OK
150818_M00766_0125_000000000-AE8BP | M00766 | 2015-08-18 | 151 | 151 | in date | in date | in date | 91.3886 | no | 15351123 | low | 804.6945 | low | 0.058805 | no | 90.8408 | no | 0.92459 | 0.8 | no | 0.94957 | no | 0.89589 | no | 0.84095 | no | 6.1155e-16 | 3.3622e-25 | 0 | -0.00052265 | yes | 0 | -0.0011087 | yes | OK | 0.31851 | yes | ['15M08233'] | [] | ['15M08233'] | NTC OK
150826_M00766_0127_000000000-AGKLT | M00766 | 2015-08-26 | 200 | 200 | in date | in date | in date | 93.1048 | no | 24799829 | low | 1017.4002 | low | 0.016826 | no | 93.586 | no | 0.88995 | 0.7 | no | 0.91037 | no | 0.86811 | no | 0.67794 | no | 4.0958e-20 | nan | 0 | -0.0013198 | yes | 0 | -0.0017885 | yes | OK | 0.094928 | yes | no outlying indexes |  |  | NTC OK
150904_M00766_0129_000000000-AFNA2 | M00766 | 2015-09-04 | 151 | 151 | in date | in date | in date | 84.0182 | no | 18908531 | low | 1029.9385 | high | 0.024332 | no | 83.7698 | no | 0.81255 | 0.8 | no | 0.86788 | no | 0.75533 | yes | 0.95413 | no | 5.0649e-39 | nan | 0 | -0.0022068 | yes | 9 | -0.0021487 | yes | OK | 0.37398 | yes | ['15M04668' '15M04664'] | ['15M04664'] | ['15M04668'] | NTC OK
150910_M00766_0130_000000000-AGJGG | M00766 | 2015-09-10 | 75 | 75 | in date | in date | in date | 88.355 | no | 24463097 | low | 1059.3697 | low | 0.0095777 | no | 88.2815 | no | 0.94136 | 0.85 | no | 0.95905 | no | 0.92225 | no | 0.95413 | no | 1.5264e-06 | 1.3678e-23 | 0 | -0.00056717 | yes | 0 | -0.00075898 | yes | OK | 0.28101 | yes | no outlying indexes |  |  | NTC OK
150916_M02641_0025_000000000-AFN33 | M02641 | 2015-09-16 | 151 | 151 | in date | in date | in date | 91.2782 | no | 21293722 | low | 1112.3198 | high | 0.023575 | no | 90.9013 | no | 0.8777 | 0.8 | no | 0.90422 | no | 0.84924 | no | 0.95413 | no | 3.9616e-16 | nan | 0 | -0.0018497 | yes | 5 | -0.0021883 | yes | OK | 0.10958 | yes | ['15M04682'] | ['15M04682'] | [] | NTC OK
150917_M02641_0026_000000000-AFN0A | M02641 | 2015-09-17 | 151 | 151 | in date | in date | in date | 91.5269 | no | 20122988 | low | 1048.5468 | OK | 0.02524 | no | 91.3116 | no | 0.86192 | 0.8 | no | 0.89213 | no | 0.82928 | no | 0.84095 | no | 2.1049e-19 | nan | 1 | -0.0022155 | yes | 10 | -0.0026282 | yes | OK | 0.18217 | yes | ['15M08778'] | ['15M08778'] | [] | NTC OK
150918_M00766_0133_000000000-AH985 | M00766 | 2015-09-18 | 200 | 200 | in date | in date | in date | 91.3077 | no | 30062610 | low | 1253.19 | OK | 0.031977 | no | 91.5655 | no | 0.84979 | 0.7 | no | 0.88478 | no | 0.81133 | no | 0.67794 | no | 4.0326e-70 | nan | 11 | -0.0017023 | yes | 24 | -0.0021771 | yes | OK | 0.15671 | yes | ['15M06627'] | [] | ['15M06627'] | NTC OK
150923_M00766_0134_000000000-AFN32 | M00766 | 2015-09-23 | 151 | 151 | in date | in date | in date | 89.0495 | no | 20113067 | low | 1072.9998 | high | 0.033799 | no | 88.906 | no | 0.86242 | 0.8 | no | 0.89294 | no | 0.82952 | no | 0.95413 | no | 1.619e-18 | nan | 0 | -0.002005 | yes | 9 | -0.0025255 | yes | OK | 0.35061 | yes | ['15M08810' '15M08808'] | ['15M08810'] | ['15M08808'] | NTC OK
150924_M02641_0027_000000000-AGJGM | M02641 | 2015-09-24 | 75 | 75 | in date | in date | in date | 95.0401 | no | 24554073 | low | 997.0288 | low | 0.0085367 | no | 95.0918 | no | 0.97292 | 0.85 | no | 0.98196 | no | 0.96511 | no | 0.84095 | no | 0.019748 | nan | 0 | -0.00019383 | no | 0 | -0.00024233 | no | OK | 0.30211 | yes | ['15M10058'] | [] | ['15M10058'] | NTC OK
150925_M00766_0135_000000000-AFL75 | M00766 | 2015-09-25 | 151 | 151 | in date | in date | in date | 97.708 | no | 9080878 | low | 460.5892 | very low | 0.039145 | no | 97.9623 | no | 0.94702 | 0.8 | no | 0.95724 | no | 0.93649 | no | 0.95413 | no | 0.0051842 | nan | 0 | -0.00069048 | yes | 0 | -0.0008509 | yes | OK | 0.46239 | yes | no outlying indexes |  |  | NTC OK
150929_M02641_0030_000000000-AFWJ3 | M02641 | 2015-09-29 | 151 | 151 | in date | in date | in date | 95.2167 | no | 14824546 | low | 765.6339 | low | 0.037672 | no | 95.287 | no | 0.94667 | 0.8 | no | 0.96411 | no | 0.92657 | no | 0.84095 | no | 2.2694e-08 | nan | 0 | -0.00044841 | no | 0 | -0.00094835 | yes | OK | 0.48248 | yes | ['15M10676' '15M09021'] | ['15M10676' '15M09021'] | [] | NTC OK
150930_M00766_0138_000000000-AH4BR | M00766 | 2015-09-30 | 75 | 75 | in date | in date | in date | 90.8515 | no | 29081266 | low | 1203.4983 | OK | 0.018375 | no | 90.911 | no | 0.93455 | 0.85 | no | 0.91651 | no | 0.95034 | no | 0.99584 | no | 5.9019e-07 | nan | 0 | -0.0013505 | yes | 0 | -0.00035323 | no | OK | 0.37912 | yes | no outlying indexes |  |  | NTC OK
151005_M02641_0034_000000000-AH9ML | M02641 | 2015-10-05 | 200 | 200 | in date | in date | in date | 91.5287 | no | 33902064 | low | 1389.8402 | high | 0.014042 | no | 91.6743 | no | 0.83382 | 0.7 | no | 0.86254 | no | 0.80114 | no | 0.50766 | no | 1.3625e-44 | nan | 23 | -0.0021037 | yes | 30 | -0.0027427 | yes | OK | 0.17597 | yes | ['15M06635'] | [] | ['15M06635'] | NTC OK
151008_M02641_0035_000000000-AFL73 | M02641 | 2015-10-08 | 151 | 151 | in date | in date | in date | 92.9471 | no | 20624078 | low | 1070.1471 | OK | 0.02054 | no | 93.2815 | no | 0.88534 | 0.8 | no | 0.91095 | no | 0.85745 | no | 0.95413 | no | 1.7378e-17 | nan | 0 | -0.0018668 | yes | 5 | -0.0023798 | yes | OK | 1.2158 | yes | ['15M09860' '15M09850' '15M09844'] | [] | ['15M09860' '15M09850' '15M09844'] | NTC OK
151013_M02641_0038_000000000-AFLW0 | M02641 | 2015-10-13 | 151 | 151 | in date | in date | in date | 92.2835 | no | 19868115 | low | 1036.6449 | OK | 0.015868 | no | 92.3144 | no | 0.85852 | 0.8 | no | 0.89862 | no | 0.81584 | no | 0.84095 | no | 9.4078e-32 | 1.0701e-96 | 0 | -0.0019454 | yes | 7 | -0.0029862 | yes | OK | 0.1712 | yes | no outlying indexes |  |  | NTC OK
151014_M02641_0039_000000000-AFLFV | M02641 | 2015-10-14 | 151 | 151 | in date | in date | in date | 93.1322 | no | 20043314 | low | 1036.9175 | OK | 0.024404 | no | 93.1321 | no | 0.86964 | 0.8 | no | 0.90574 | no | 0.83081 | no | 0.84095 | no | 1.4818e-31 | nan | 0 | -0.0019202 | yes | 8 | -0.0027745 | yes | OK | 0.27897 | yes | ['15M10686'] | ['15M10686'] | [] | NTC OK
151015_M00766_0142_000000000-AJD8B | M00766 | 2015-10-15 | 151 | 151 | in date | in date | in date | 88.82 | no | 16091719 | low | 856.7867 | low | 0.030984 | no | 88.9965 | no | 0.9206 | 0.8 | no | 0.95153 | no | 0.88855 | no | 0.67794 | no | 9.7576e-25 | nan | 0 | -0.0004536 | no | 1 | -0.0013313 | yes | OK | 0.2848 | yes | ['15M10752' '15M10508'] | ['15M10508'] | ['15M10752'] | NTC OK
151021_M02641_0041_000000000-AGHAA | M02641 | 2015-10-21 | 75 | 75 | in date | in date | in date | 93.3815 | no | 29221047 | low | 1189.4531 | OK | 0.018033 | no | 93.5119 | no | 0.96054 | 0.85 | no | 0.97325 | no | 0.9489 | no | 0.84095 | no | 0.001648 | 0.00019262 | 0 | -0.00043517 | no | 0 | -0.00032505 | no | OK | 0.42919 | yes | ['15M11968'] | [] | ['15M11968'] | NTC OK
151022_M00766_0146_000000000-AFN3B | M00766 | 2015-10-22 | 151 | 151 | in date | in date | in date | 81.4565 | no | 19862739 | low | 1107.5448 | very high | 0.031521 | no | 80.9673 | no | 0.90849 | 0.8 | no | 0.93308 | no | 0.88275 | no | 0.95413 | no | 6.0372e-16 | 1.7055e-184 | 0 | -0.00050949 | yes | 0 | -0.00097367 | yes | OK | 0.21273 | yes | no outlying indexes |  |  | NTC OK
151023_M00766_0147_000000000-AFN37 | M00766 | 2015-10-23 | 151 | 151 | in date | in date | in date | 84.6509 | no | 20693511 | low | 1134.0694 | very high | 0.039199 | no | 85.3046 | no | 0.92707 | 0.8 | no | 0.95123 | no | 0.90247 | no | 0.95413 | no | 6.8289e-18 | nan | 0 | -0.00030046 | no | 0 | -0.00056736 | yes | OK | 0.30167 | yes | ['15M10999' '15M10996'] | ['15M10996'] | ['15M10999'] | NTC OK
151028_M02641_0043_000000000-AFL74 | M02641 | 2015-10-28 | 151 | 151 | in date | in date | in date | 96.8353 | no | 15543815 | low | 805.7896 | low | 0.035546 | no | 96.7917 | no | 0.96903 | 0.8 | no | 0.97972 | no | 0.95847 | no | 0.67794 | no | 8.1444e-08 | 1.4081e-14 | 0 | -0.00020911 | no | 0 | -0.0004479 | no | OK | 0.15867 | yes | no outlying indexes |  |  | NTC OK
151030_M00766_0150_000000000-AFLDT | M00766 | 2015-10-30 | 151 | 151 | in date | in date | in date | 89.1258 | no | 19498823 | low | 1043.7728 | OK | 0.022316 | no | 89.1735 | no | 0.94282 | 0.8 | no | 0.95917 | no | 0.92637 | no | 0.95413 | no | 6.2812e-11 | 1.3707e-27 | 0 | -0.00027197 | no | 0 | -0.00043246 | no | OK | 0.50206 | yes | ['15M11131'] | [] | ['15M11131'] | NTC OK
151030_M02641_0044_000000000-AJHLJ | M02641 | 2015-10-30 | 75 | 75 | in date | in date | in date | 89.5684 | no | 33168068 | low | 1381.3797 | high | 0.015021 | no | 89.5367 | no | 0.94181 | 0.85 | no | 0.96027 | no | 0.92327 | no | 0.84095 | no | 1.9212e-09 | 1.7277e-37 | 0 | -0.00053639 | yes | 0 | -0.00046394 | no | OK | 0.50917 | yes | ['15M12049'] | [] | ['15M12049'] | NTC OK
151103_M02641_0046_000000000-AJ5Y7 | M02641 | 2015-11-03 | 151 | 151 | in date | in date | in date | 96.1111 | no | 17192399 | low | 883.2398 | low | 0.021428 | no | 96.0917 | no | 0.96652 | 0.8 | no | 0.97476 | no | 0.95884 | no | 0.95413 | no | 0.00063959 | nan | 0 | -0.00027215 | no | 0 | -0.00034016 | no | OK | 0.41562 | yes | ['15M11227'] | [] | ['15M11227'] | NTC OK
151105_M02641_0047_000000000-AJF4E | M02641 | 2015-11-05 | 151 | 151 | in date | in date | in date | 96.0812 | no | 17316625 | low | 888.5653 | low | 0.024517 | no | 96.1753 | no | 0.96613 | 0.8 | no | 0.97668 | no | 0.95675 | no | 0.84095 | no | 7.3533e-09 | 2.5766e-25 | 0 | -0.00020337 | no | 0 | -0.00027243 | no | OK | 1.7725 | yes | ['15M11323'] | [] | ['15M11323'] | NTC OK
151116_M02641_0050_000000000-AJ6L3 | M02641 | 2015-11-16 | 151 | 151 | in date | in date | in date | 95.1303 | no | 13144452 | low | 672.6944 | low | 0.074365 | no | 94.8898 | no | 0.95687 | 0.8 | no | 0.96857 | no | 0.94432 | no | 0.95413 | no | 0.0016197 | 1.4817e-06 | 0 | -0.00029535 | no | 0 | -0.00053054 | yes | OK | 0.17711 | yes | no outlying indexes |  |  | NTC OK
151117_M00766_0155_000000000-AJ6LL | M00766 | 2015-11-17 | 151 | 151 | in date | in date | in date | 87.0005 | no | 16509092 | low | 884.2256 | OK | 0.024563 | no | 86.7919 | no | 0.86377 | 0.8 | no | 0.9042 | no | 0.821 | no | 0.84095 | no | 8.628e-16 | nan | 0 | -0.0012235 | yes | 1 | -0.00063026 | yes | OK | 0.51279 | yes | ['15M11867'] | [] | ['15M11867'] | NTC OK
151117_M02641_0051_000000000-AJF4A | M02641 | 2015-11-17 | 151 | 151 | in date | in date | in date | 94.2241 | no | 17785498 | low | 905.3067 | low | 0.03067 | no | 94.2545 | no | 0.88313 | 0.8 | no | 0.91617 | no | 0.84833 | no | 0.84095 | no | 7.9592e-12 | nan | 0 | -0.0012532 | yes | 0 | -0.00072138 | yes | OK | 0.56161 | yes | ['15M12552'] | [] | ['15M12552'] | NTC OK
151125_M02641_0053_000000000-AJJ3G | M02641 | 2015-11-25 | 151 | 151 | in date | in date | in date | 96.6504 | no | 3965190 | low | 201.2481 | very low | 0.02812 | no | 96.9381 | no | 0.97423 | 0.8 | no | 0.98297 | no | 0.96916 | no | 0.67794 | no | 0.069974 | 0.0030109 | 0 | -0.00019201 | no | 0 | -0.00032396 | no | OK | 1.7978 | yes | ['15M12551'] | [] | ['15M12551'] | NTC OK
151126_M00766_0159_000000000-AJD7L | M00766 | 2015-11-26 | 75 | 75 | in date | in date | in date | 88.0477 | no | 24133209 | low | 1032.2091 | low | 0.017289 | no | 87.5838 | no | 0.95346 | 0.85 | no | 0.96296 | no | 0.94323 | no | 0.84095 | no | 0.093829 | 1.6996e-15 | 0 | -0.0004886 | no | 0 | -0.00047968 | no | OK | 0.27264 | yes | no outlying indexes |  |  | NTC OK
151126_M02641_0054_000000000-AJD7T | M02641 | 2015-11-26 | 75 | 75 | in date | in date | in date | 94.0217 | no | 24949166 | low | 1025.4829 | low | 0.029656 | no | 93.8083 | no | 0.96822 | 0.85 | no | 0.97532 | no | 0.96185 | no | 0.95413 | no | 0.1922 | nan | 0 | -0.00030117 | no | 0 | -0.00033189 | no | OK | 0.27535 | yes | ['15M13666'] | [] | ['15M13666'] | NTC OK
151127_M00766_0161_000000000-AJERP | M00766 | 2015-11-27 | 75 | 75 | in date | in date | in date | 90.539 | no | 24073552 | low | 1002.8265 | low | 0.022959 | no | 90.0889 | no | 0.95458 | 0.85 | no | 0.96258 | no | 0.94625 | no | 0.95413 | no | 0.1243 | nan | 0 | -0.00057536 | yes | 0 | -0.00051622 | yes | OK | 0.56739 | yes | no outlying indexes |  |  | NTC OK
151201_M02641_0055_000000000-AJ6B2 | M02641 | 2015-12-01 | 151 | 151 | in date | in date | in date | 93.147 | no | 20354768 | low | 1062.1404 | OK | 0.019929 | no | 93.6277 | no | 0.95729 | 0.8 | no | 0.96997 | no | 0.94485 | no | 0.95413 | no | 2.0042e-09 | nan | 0 | -0.00022088 | no | 0 | -0.0003646 | no | OK | 1.8796 | yes | ['15M12551'] | [] | ['15M12551'] | NTC OK
151203_M00766_0165_000000000-AJFAA | M00766 | 2015-12-03 | 151 | 151 | in date | in date | in date | 86.1622 | no | 16734072 | low | 890.0657 | OK | 0.048269 | no | 84.2361 | no | 0.92678 | 0.8 | no | 0.94416 | no | 0.90681 | no | 0.95413 | no | 1.2134e-06 | nan | 0 | -0.0003637 | no | 0 | -0.00065096 | yes | OK | 0.47912 | yes | ['15M13694'] | ['15M13694'] | [] | NTC OK
151208_M00766_0167_000000000-AJ5JR | M00766 | 2015-12-08 | 151 | 151 | in date | in date | in date | 97.4409 | no | 5807543 | low | 293.2064 | very low | 0.036817 | no | 97.4486 | no | 0.96505 | 0.8 | no | 0.973 | no | 0.95786 | no | 0.95413 | no | 0.12937 | nan | 0 | -0.00036542 | no | 0 | -0.00043496 | no | OK | 0.48404 | yes | no outlying indexes |  |  | NTC OK
151210_M02641_0058_000000000-AJ5YJ | M02641 | 2015-12-10 | 151 | 151 | in date | in date | in date | 91.7526 | no | 18225198 | low | 951.6353 | OK | 0.049608 | no | 91.4934 | no | 0.93902 | 0.8 | no | 0.95565 | no | 0.92066 | no | 0.95413 | no | 6.5736e-05 | nan | 0 | -0.00040902 | no | 0 | -0.00072147 | yes | OK | 0.29821 | yes | no outlying indexes |  |  | NTC OK
151216_M00766_0170_000000000-ALEAF | M00766 | 2015-12-16 | 151 | 151 | in date | in date | in date | 93.7279 | no | 17003866 | low | 880.7103 | low | 0.036745 | no | 93.9642 | no | 0.94991 | 0.8 | no | 0.96257 | no | 0.937 | no | 0.95413 | no | 4.5633e-06 | nan | 0 | -0.00035203 | no | 0 | -0.00057532 | yes | OK | 0.58674 | yes | ['15M13059'] | [] | ['15M13059'] | NTC OK
151217_M02641_0059_000000000-AJFAD | M02641 | 2015-12-17 | 151 | 151 | in date | in date | in date | 58.3542 | yes | 4092271 | low | 204.0891 | very low | 0.28007 | yes | 55.6594 | yes | 0.92478 | 0.8 | no | 0.94271 | no | 0.91061 | no | 0.84095 | no | 0.016952 | 0.00029975 | 0 | -0.00041346 | no | 0 | -0.00076263 | yes | OK | 0.24432 | yes | no outlying indexes |  |  | NTC OK
151223_M00766_0172_000000000-AJRMN | M00766 | 2015-12-23 | 75 | 75 | in date | in date | in date | 92.6132 | no | 23635202 | low | 977.0359 | low | 0.01686 | no | 92.2318 | no | 0.94189 | 0.85 | no | 0.9657 | no | 0.91522 | no | 0.50766 | no | 8.2817e-14 | nan | 0 | -0.00070693 | yes | 1 | -0.0018884 | yes | OK | 0.42725 | yes | no outlying indexes |  |  | NTC OK
151223_M02641_0060_000000000-AGHV2 | M02641 | 2015-12-23 | 75 | 75 | in date | in date | in date | 87.711 | no | 26108005 | low | 1077.9707 | OK | 0.015686 | no | 87.8311 | no | 0.92979 | 0.85 | no | 0.94545 | no | 0.911 | no | 0.95413 | no | 1.0014e-06 | 6.7599e-86 | 0 | -0.00093076 | yes | 0 | -0.0011369 | yes | OK | 0.6133 | yes | ['15M14273'] | [] | ['15M14273'] | NTC OK
151231_M02641_0061_000000000-AJD8L | M02641 | 2015-12-31 | 75 | 75 | in date | in date | in date | 88.6234 | no | 30986761 | low | 1292.1541 | high | 0.043562 | no | 89.278 | no | 0.93779 | 0.85 | no | 0.95903 | no | 0.91501 | no | 0.95413 | no | 2.22e-08 | nan | 0 | -0.00055157 | yes | 0 | -0.00062885 | yes | OK | 0.40312 | yes | ['15M14557'] | [] | ['15M14557'] | NTC OK
160107_M00766_0175_000000000-AH6C3 | M00766 | 2016-01-07 | 75 | 75 | in date | in date | in date | 87.9117 | no | 26044869 | low | 1123.4126 | OK | 0.012327 | no | 86.8229 | no | 0.95101 | 0.85 | no | 0.96122 | no | 0.9395 | no | 0.95413 | no | 0.056283 | nan | 0 | -0.00056289 | yes | 0 | -0.00070024 | yes | OK | 0.41119 | yes | ['15m15138'] | [] | ['15m15138'] | NTC OK
160108_M00766_0176_000000000-AJDAJ | M00766 | 2016-01-08 | 75 | 75 | in date | in date | in date | 86.9378 | no | 24436074 | low | 1042.2918 | low | 0.023113 | no | 87.1996 | no | 0.79786 | 0.85 | yes | 0.93046 | no | 0.6482 | yes | 0.15463 | no | 1.2164e-131 | nan | 0 | -0.0018269 | yes | 36 | -0.0074037 | yes | OK | 0.35456 | yes | ['15M15229'] | ['15M15229'] | [] | NTC OK
160114_M00766_0177_000000000-ALYTH | M00766 | 2016-01-14 | 75 | 75 | in date | in date | in date | 87.5409 | no | 27732794 | low | 1180.5507 | OK | 0.016589 | no | 87.6219 | no | 0.94367 | 0.85 | no | 0.95391 | no | 0.93177 | no | 0.95413 | no | 0.043231 | nan | 0 | -0.00066231 | yes | 0 | -0.00077358 | yes | OK | 0.52375 | yes | no outlying indexes |  |  | NTC OK
160114_M02641_0065_000000000-ALRU0 | M02641 | 2016-01-14 | 75 | 75 | in date | in date | in date | 91.58 | no | 29247440 | low | 1214.9636 | OK | 0.024503 | no | 91.5468 | no | 0.95131 | 0.85 | no | 0.96723 | no | 0.93592 | no | 0.84095 | no | 7.4878e-05 | nan | 0 | -0.00049951 | no | 0 | -0.00034584 | no | OK | 0.42872 | yes | ['15M03903'] | [] | ['15M03903'] | NTC OK
160128_M02641_0070_000000000-ALAGE | M02641 | 2016-01-28 | 75 | 75 | in date | in date | in date | 91.826 | no | 24305600 | low | 1013.4527 | low | 0.0159 | no | 91.7655 | no | 0.96174 | 0.85 | no | 0.96994 | no | 0.95369 | no | 0.95413 | no | 0.073098 | nan | 0 | -0.0003153 | no | 0 | -0.00042581 | no | OK | 0.32784 | yes | no outlying indexes |  |  | NTC OK
160129_M00766_0003_000000000-ALAH0 | M00766 | 2016-01-29 | 75 | 75 | in date | in date | in date | 86.7633 | no | 32429499 | low | 1380.9019 | high | 0.017673 | no | 87.2761 | no | 0.93918 | 0.85 | no | 0.95507 | no | 0.92174 | no | 0.84095 | no | 1.4395e-05 | nan | 0 | -0.00058412 | yes | 0 | -0.00060062 | yes | OK | 0.36577 | yes | ['16M00136'] | [] | ['16M00136'] | NTC OK
160129_M02641_0071_000000000-AMC8N | M02641 | 2016-01-29 | 200 | 200 | in date | in date | in date | 91.7102 | no | 25390554 | low | 1054.7902 | low | 0.01824 | no | 91.6301 | no | 0.86183 | 0.7 | no | 0.88102 | no | 0.83939 | no | 0.67794 | no | 1.3683e-14 | 6.476e-15 | 1 | -0.0017129 | yes | 0 | -0.0020671 | yes | OK | 0.14555 | yes | ['15M09614'] | ['15M09614'] | [] | NTC OK
160206_M00766_0005_000000000-ALYU1 | M00766 | 2016-02-06 | 75 | 75 | in date | in date | in date | 84.6656 | no | 32152720 | low | 1381.1846 | high | 0.017535 | no | 83.6833 | no | 0.93615 | 0.85 | no | 0.9506 | no | 0.92023 | no | 0.84095 | no | 0.0023058 | 1.3476e-13 | 0 | -0.00053805 | yes | 0 | -0.00042151 | no | OK | 0.27786 | yes | no outlying indexes |  |  | NTC OK
160206_M02641_0074_000000000-AL604 | M02641 | 2016-02-06 | 151 | 151 | in date | in date | in date | 88.8163 | no | 18930346 | low | 1014.3724 | OK | 0.015119 | no | 88.4642 | no | 0.92173 | 0.8 | no | 0.9483 | no | 0.89269 | no | 0.67794 | no | 2.5587e-23 | 1.4688e-132 | 0 | -0.00040622 | no | 0 | -0.00095862 | yes | High | 0.80948 | yes | ['16M00477'] | [] | ['16M00477'] | No NTC on run
160209_M00766_0007_000000000-AMERJ | M00766 | 2016-02-09 | 75 | 75 | in date | in date | in date | 79.4588 | no | 34759702 | low | 1528.9595 | very high | 0.016677 | no | 80.1993 | no | 0.90703 | 0.85 | no | 0.93282 | no | 0.87982 | no | 0.84095 | no | 1.1867e-08 | 0 | 0 | -0.00081183 | yes | 0 | -0.00073619 | yes | OK | 0.29288 | yes | no outlying indexes |  |  | NTC OK
160216_M00766_0009_000000000-AL73W | M00766 | 2016-02-16 | 151 | 151 | in date | in date | in date | 92.117 | no | 18300588 | low | 961.5089 | OK | 0.025014 | no | 92.5591 | no | 0.94116 | 0.8 | no | 0.96156 | no | 0.91989 | no | 0.84095 | no | 1.5361e-15 | 3.496e-24 | 0 | -0.00029836 | no | 0 | -0.00066338 | yes | OK | 0.51285 | yes | ['15M13195'] | [] | ['15M13195'] | NTC OK
160219_M00766_0011_000000000-AL6GM | M00766 | 2016-02-19 | 131 | 131 | in date | in date | in date | 5.6098 | yes | 661574 | low | 22.0869 | very low | 0.9623 | yes | 9.7324 | yes | 0.77871 | 0.8 | yes | 0.79992 | yes | 0.78598 | yes | 0.99584 | no | 1 | 0.73572 | 5 | -0.00056401 | yes | 5 | -0.00059171 | yes | No index information for 160219_M00766_0011_000000000-AL6GM
160220_M02641_0081_000000000-AMETF | M02641 | 2016-02-20 | 200 | 200 | in date | in date | in date | 92.9979 | no | 23074941 | low | 958.5024 | low | 0.027553 | no | 93.6783 | no | 0.87882 | 0.7 | no | 0.90362 | no | 0.85132 | no | 0.35842 | no | 6.4948e-26 | nan | 0 | -0.0012552 | yes | 0 | -0.0017992 | yes | OK | 0.10059 | yes | ['15M14092'] | [] | ['15M14092'] | NTC OK
160223_M00766_0012_000000000-AL5YC | M00766 | 2016-02-23 | 151 | 151 | in date | in date | in date | 93.2958 | no | 18111084 | low | 951.0894 | OK | 0.019086 | no | 93.3336 | no | 0.94516 | 0.8 | no | 0.96589 | no | 0.92393 | no | 0.84095 | no | 9.6635e-16 | 3.8603e-52 | 0 | -0.0002968 | no | 0 | -0.0005917 | yes | OK | 1.8862 | yes | ['15M13471'] | [] | ['15M13471'] | NTC OK

## AnalyData
RunID | InstrumentID | RunDate | NumberOfCyclesR1 | NumberofCyclesR2 | %ReadsPassingFilter | %ReadsPassingFilterOutsideThreshold? | NumberofReadsPassingFilter | ReadCountWithinThreshold? | MeanClusterDensityPassingFilter | MeanClusterDensityWithinThreshold | COVClusterDensity | COVClusterDensityExceedsThreshold? | %MedianClusterDensityPassingFilter | GapBetweenClusterDensitiesPassingFilterAndNotExceedsThreshold? | ProportionOverQ30 | IlluminaQ30Threshold | ProportionOverQ30ExceedsThreshold? | ProportionOverQ30Read1 | ProportionOverQ30ExceedsThresholdRead1 | ProportionOverQ30Read2 | ProportionOverQ30ExceedsThresholdRead2 | K-SPValue | K-SPValueSignificant? | K-SUnpackPvalue | M-WUnpackPvalue | NumberOfCyclesExceedingThresholdUnderQ30Read1 | SlopeOfLineOfBestFitRead1 | SlopeOfLineOfBestFitRead1ExceedsThreshold? | NumberOfCyclesExceedingThresholdUnderQ30Read2 | SlopeOfLineOfBestFitRead2 | SlopeOfLineOfBestFitRead2ExceedsThreshold? | Fail? | Fail2? | Fail? No CD | Fail2? No CD | R1vsR2PropQ30 | R1vsR2K-S | NumCycleR1Threshold | NumCycleR2Threshold | R1vsR2CyclesQ30 | R1vsR2bestfit
121009_M00766_0002_000000000-A1U6P | M00766 | 2012-10-09 | 151 | 151 | 45.385 | yes | 3138050 | low | 341.3139 | low | 0.34653 | yes | 0 | yes | 0.87575 | 0.8 | no | 0.87642 | no | 0.87507 | no | 0.99998 | no | 0.99997 | 0.80608 | 1 | 0 | no | 0 | 0 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
130206_M00766_0002_000000000-A23JM | M00766 | 2013-02-06 | 250 | 250 | 2.0076 | yes | 160923 | low | 10.2917 | very low | 0.26427 | yes | 1.9741 | yes | 0.27654 | 0.75 | yes | 0.091188 | yes | 0.46189 | yes | 0.35842 | no | 0.0011204 | 0.012662 | 193 | -0.0010965 | yes | 65 | -0.0032351 | yes | fail | fail | fail | fail | same | pass | exceeded | exceeded | same | same
130618_M00766_0019_000000000-A4FEU | M00766 | 2013-06-18 | 251 | 251 | 93.703 | no | 10147864 | low | 556.7555 | very low | 0.009214 | no | 93.9257 | no | 0.79545 | 0.75 | no | 0.90383 | no | 0.68024 | yes | 0.35842 | no | 1.6395e-83 | 9.0394e-306 | 0 | -0.0010825 | yes | 47 | -0.0037713 | yes | fail | fail | pass | fail | diff | pass | not exceeded | exceeded | diff | same
130624_M00766_0021_000000000-A53PT | M00766 | 2013-06-24 | 251 | 251 | 88.7381 | no | 17725761 | low | 940.8335 | OK | 0.011629 | no | 89.0017 | no | 0.84736 | 0.75 | no | 0.90275 | no | 0.78576 | no | 0.67794 | no | 1.0912e-52 | 1.2366e-63 | 0 | -0.0009554 | yes | 4 | -0.0019889 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
130712_M00766_0024_000000000-A5AVV | M00766 | 2013-07-12 | 151 | 151 | 82.4401 | no | 13180142 | low | 710.2954 | low | 0.09422 | yes | 82.1308 | no | 0.92195 | 0.8 | no | 0.92377 | no | 0.92292 | no | 0.99998 | no | 0.0096055 | 4.6381e-11 | 0 | -0.00059232 | yes | 0 | -0.00062141 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
130817_M00766_0037_000000000-A4EDD | M00766 | 2013-08-17 | 176 | 176 | 93.613 | no | 14493580 | low | 748.7721 | low | 0.020101 | no | 93.6127 | no | 0.94897 | 0.75 | no | 0.96312 | no | 0.93388 | no | 0.84095 | no | 4.6236e-05 | nan | 0 | -0.00039473 | no | 1 | -0.0006017 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
130913_M00766_0049_000000000-A5B0E | M00766 | 2013-09-13 | 251 | 251 | 89.8636 | no | 17433281 | low | 913.7303 | OK | 0.024696 | no | 89.1779 | no | 0.83393 | 0.75 | no | 0.88573 | no | 0.77648 | no | 0.84095 | no | 3.8371e-42 | 5.6819e-66 | 0 | -0.0011656 | yes | 4 | -0.0019759 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
130916_M00766_0050_000000000-A5M1U | M00766 | 2013-09-16 | 151 | 151 | 86.0618 | no | 22033351 | low | 1185.4173 | very high | 0.028468 | no | 86.0122 | no | 0.9138 | 0.8 | no | 0.93477 | no | 0.89056 | no | 0.84095 | no | 4.3159e-13 | 1.1035e-23 | 0 | -0.00070235 | yes | 0 | -0.0009045 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
130924_M00766_0052_000000000-A5BE6 | M00766 | 2013-09-24 | 251 | 251 | 67.1952 | yes | 14006895 | low | 805.1388 | high | 0.092176 | yes | 66.8443 | yes | 0.71517 | 0.75 | yes | 0.77069 | no | 0.66348 | yes | 0.99584 | no | 6.3637e-20 | nan | 16 | -0.0019789 | yes | 32 | -0.002186 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same
131007_M00766_0055_000000000-A59JT | M00766 | 2013-10-07 | 251 | 251 | 88.9869 | no | 7066429 | low | 464.0561 | very low | 0.11395 | yes | 90.7266 | no | 0.86005 | 0.75 | no | 0.86612 | no | 0.85247 | no | 1 | no | 0.071573 | nan | 1 | -0.0014022 | yes | 1 | -0.0015446 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
131025_M00766_0059_000000000-A5P9A | M00766 | 2013-10-25 | 151 | 151 | 90.6341 | no | 17754371 | low | 930.3591 | OK | 0.030836 | no | 91.4036 | no | 0.92242 | 0.8 | no | 0.93549 | no | 0.90546 | no | 0.99584 | no | 1.9427e-05 | nan | 0 | -0.00094972 | yes | 0 | -0.0010274 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
140207_M00766_0022_000000000-A7PH8 | M00766 | 2014-02-07 | 251 | 251 | 88.3301 | no | 10785103 | low | 558.1003 | low | 0.053446 | no | 89.8657 | no | 0.85142 | 0.75 | no | 0.8548 | no | 0.84338 | no | 1 | no | 0.42818 | nan | 1 | -0.0014636 | yes | 1 | -0.0016776 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
140424_M00766_0033_000000000-A7BNA | M00766 | 2014-04-24 | 151 | 151 | 97.0072 | no | 5094671 | low | 251.0711 | very low | 0.231 | yes | 98.5469 | no | 0.93672 | 0.8 | no | 0.94645 | no | 0.92545 | no | 0.95413 | no | 0.083102 | 0.0055314 | 0 | -0.00096088 | yes | 0 | -0.0010359 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
140514_M00766_0036_000000000-A7BRK | M00766 | 2014-05-14 | 151 | 151 | 81.4684 | no | 20334208 | low | 1124.4827 | very high | 0.017101 | no | 81.2844 | no | 0.79665 | 0.8 | yes | 0.83677 | no | 0.75502 | yes | 0.99584 | no | 4.1492e-14 | nan | 0 | -0.0024166 | yes | 7 | -0.0018547 | yes | fail | fail | fail | fail | diff | pass | not exceeded | not exceeded | same | same
140612_M00766_0039_000000000-A7BNL | M00766 | 2014-06-12 | 251 | 251 | 91.373 | no | 20643048 | low | 1088.1751 | OK | 0.014529 | no | 90.3695 | no | 0.78766 | 0.75 | no | 0.82619 | no | 0.74426 | yes | 0.84095 | no | 1.1316e-41 | nan | 19 | -0.0022137 | yes | 39 | -0.0026621 | yes | pass | fail | pass | fail | diff | pass | not exceeded | exceeded | diff | same
140616_M00766_0040_000000000-A78V9 | M00766 | 2014-06-16 | 251 | 251 | 97.3311 | no | 10710545 | low | 534.5911 | very low | 0.0065097 | no | 97.505 | no | 0.81498 | 0.75 | no | 0.8406 | no | 0.7849 | no | 0.95413 | no | 7.2918e-10 | nan | 1 | -0.0024062 | yes | 30 | -0.0029712 | yes | fail | fail | pass | pass | same | pass | not exceeded | exceeded | diff | same
140813_M00766_0047_000000000-A8PJL | M00766 | 2014-08-13 | 151 | 151 | 91.9352 | no | 17936069 | low | 943.905 | OK | 0.026213 | no | 91.9798 | no | 0.92465 | 0.8 | no | 0.94352 | no | 0.90463 | no | 0.95413 | no | 3.3998e-09 | 3.0143e-185 | 0 | -0.00068224 | yes | 0 | -0.00075951 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
140901_M00766_0048_000000000-AA63M | M00766 | 2014-09-01 | 151 | 151 | 94.0673 | no | 17640397 | low | 909.8289 | low | 0.02007 | no | 93.8944 | no | 0.93862 | 0.8 | no | 0.96072 | no | 0.91538 | no | 0.84095 | no | 4.2142e-16 | 6.6658e-224 | 0 | -0.00044115 | no | 0 | -0.00052041 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
141014_M00766_0053_000000000-A8R6M | M00766 | 2014-10-14 | 151 | 151 | 89.6387 | no | 19084469 | low | 1011.4321 | OK | 0.014006 | no | 90.0342 | no | 0.89321 | 0.8 | no | 0.91865 | no | 0.86561 | no | 0.95413 | no | 8.0153e-17 | 2.8333e-171 | 0 | -0.0010561 | yes | 0 | -0.0011248 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
141017_M02641_0022_000000000-AA66H | M02641 | 2014-10-17 | 151 | 151 | 79.2948 | no | 22430099 | low | 1210.1492 | very high | 0.017994 | no | 78.9705 | yes | 0.87741 | 0.8 | no | 0.91601 | no | 0.83648 | no | 0.99584 | no | 1.1016e-30 | nan | 0 | -0.00063613 | yes | 0 | -0.00060429 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
141024_M00766_0056_000000000-A8PC5 | M00766 | 2014-10-24 | 151 | 151 | 86.6935 | no | 22517755 | low | 1216.7323 | very high | 0.017442 | no | 86.7307 | no | 0.88873 | 0.8 | no | 0.92881 | no | 0.84641 | no | 0.84095 | no | 1.3784e-37 | 2.2642e-64 | 0 | -0.00069708 | yes | 0 | -0.00081119 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
141031_M00766_0058_000000000-AA3GN | M00766 | 2014-10-31 | 75 | 75 | 84.8049 | no | 22830617 | low | 1020.36 | OK | 0.066963 | no | 85.4852 | no | 0.94188 | 0.85 | no | 0.94357 | no | 0.93798 | no | 0.99584 | no | 0.16727 | nan | 0 | -0.00072598 | yes | 0 | -0.0003319 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
141107_M00766_0060_000000000-AA8PM | M00766 | 2014-11-07 | 151 | 151 | 90.0798 | no | 11262793 | low | 579.4485 | low | 0.034966 | no | 90.2121 | no | 0.96028 | 0.8 | no | 0.96678 | no | 0.9539 | no | 0.99584 | no | 0.99203 | 0.8978 | 0 | -0.00013806 | no | 0 | -0.00011747 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
141118_M00766_0061_000000000-A8P8J | M00766 | 2014-11-18 | 151 | 151 | 92.7322 | no | 19181201 | low | 1009.9074 | OK | 0.027961 | no | 92.312 | no | 0.93572 | 0.8 | no | 0.94773 | no | 0.92218 | no | 0.95413 | no | 1.3904e-05 | nan | 0 | -0.0006231 | yes | 0 | -0.00072415 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
141208_M00766_0063_000000000-A8P7C | M00766 | 2014-12-08 | 151 | 151 | 90.051 | no | 21795042 | low | 1153.0041 | high | 0.020941 | no | 89.941 | no | 0.89826 | 0.8 | no | 0.94464 | no | 0.85869 | no | 0.67794 | no | 1.8648e-37 | nan | 0 | -0.00060048 | yes | 3 | -0.00051121 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
141208_M02641_0026_000000000-A8R55 | M02641 | 2014-12-08 | 151 | 151 | 94.9912 | no | 13800416 | low | 717.5691 | low | 0.027596 | no | 95.1078 | no | 0.84608 | 0.8 | no | 0.94004 | no | 0.74763 | yes | 0.35842 | no | 7.7537e-147 | 6.3912e-157 | 0 | -0.00078235 | yes | 5 | -0.0011709 | yes | pass | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same
141212_M00766_0064_000000000-ACCEB | M00766 | 2014-12-12 | 151 | 151 | 81.1933 | no | 20826514 | low | 1127.546 | very high | 0.021875 | no | 81.0162 | no | 0.74058 | 0.8 | yes | 0.80816 | no | 0.67993 | yes | 0.99584 | no | 3.8905e-34 | nan | 0 | -0.0028592 | yes | 24 | -0.0015004 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same
141216_M00766_0065_000000000-ACCDT | M00766 | 2014-12-16 | 151 | 151 | 79.0858 | no | 20548879 | low | 1120.7265 | very high | 0.02252 | no | 79.0848 | yes | 0.77398 | 0.8 | yes | 0.84162 | no | 0.70405 | yes | 0.84095 | no | 1.4032e-49 | 7.9356e-94 | 0 | -0.0021868 | yes | 16 | -0.00079947 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same
141219_M00766_0066_000000000-ACCB1 | M00766 | 2014-12-19 | 151 | 151 | 84.1075 | no | 16554828 | low | 882.4454 | OK | 0.030188 | no | 83.544 | no | 0.71631 | 0.8 | yes | 0.77794 | yes | 0.66061 | yes | 0.99998 | no | 9.0168e-19 | 1.7886e-30 | 0 | -0.0034797 | yes | 18 | -0.0018779 | yes | fail | fail | fail | fail | same | pass | not exceeded | exceeded | diff | same
150120_M00766_0071_000000000-AA63K | M00766 | 2015-01-20 | 151 | 151 | 85.9694 | no | 14794255 | low | 793.1929 | low | 0.05599 | no | 87.4592 | no | 0.90779 | 0.8 | no | 0.93487 | no | 0.89285 | no | 0.95413 | no | 0.01591 | nan | 0 | -0.00060834 | yes | 3 | -0.00064423 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150127_M02641_0027_000000000-AA65J | M02641 | 2015-01-27 | 151 | 151 | 94.4613 | no | 15926233 | low | 820.6824 | low | 0.035198 | no | 94.2959 | no | 0.93014 | 0.8 | no | 0.95953 | no | 0.9008 | no | 0.67794 | no | 1.0379e-16 | 3.674e-28 | 0 | -0.0004646 | no | 0 | -0.00058277 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150203_M00766_0074_000000000-AAUMH | M00766 | 2015-02-03 | 151 | 151 | 81.6068 | no | 16649356 | low | 884.3212 | OK | 0.052428 | no | 81.2451 | no | 0.90342 | 0.8 | no | 0.92886 | no | 0.88527 | no | 0.95413 | no | 0.0090932 | 0.028862 | 0 | -0.00047881 | no | 3 | -0.00020141 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150203_M02641_0028_000000000-ACBYG | M02641 | 2015-02-03 | 151 | 151 | 89.643 | no | 18845203 | low | 988.6001 | OK | 0.014743 | no | 90.0779 | no | 0.82841 | 0.8 | no | 0.86795 | no | 0.78849 | yes | 0.99584 | no | 1.2512e-16 | 0 | 0 | -0.0022305 | yes | 3 | -0.0018175 | yes | pass | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same
150205_M00766_0075_000000000-ACC43 | M00766 | 2015-02-05 | 151 | 151 | 93.254 | no | 17676890 | low | 914.0169 | low | 0.038617 | no | 93.3074 | no | 0.8595 | 0.8 | no | 0.89452 | no | 0.83021 | no | 0.95413 | no | 5.2726e-15 | 3.8357e-55 | 0 | -0.0019394 | yes | 4 | -0.0017318 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150224_M00766_0078_000000000-ACMP5 | M00766 | 2015-02-24 | 75 | 75 | 82.4896 | no | 36080266 | low | 1546.85 | very high | 0.024283 | no | 82.282 | no | 0.93168 | 0.85 | no | 0.94868 | no | 0.91268 | no | 0.84095 | no | 1.1086e-05 | 0 | 0 | -0.00047772 | no | 0 | -0.00047646 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150309_M02641_0035_000000000-ACC3J | M02641 | 2015-03-09 | 151 | 151 | 88.8745 | no | 22186253 | low | 1157.1795 | very high | 0.017088 | no | 89.0446 | no | 0.79385 | 0.8 | yes | 0.83769 | no | 0.74848 | yes | 0.99584 | no | 2.3483e-15 | nan | 0 | -0.002745 | yes | 7 | -0.0021184 | yes | fail | fail | fail | fail | diff | pass | not exceeded | not exceeded | same | same
150311_M00766_0081_000000000-ACN0W | M00766 | 2015-03-11 | 75 | 75 | 92.0795 | no | 28700881 | low | 1177.6318 | OK | 0.016158 | no | 92.5321 | no | 0.96191 | 0.85 | no | 0.9725 | no | 0.95082 | no | 0.84095 | no | 0.010916 | nan | 0 | -0.00034381 | no | 0 | -0.00028909 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150311_M02641_0036_000000000-AD8VB | M02641 | 2015-03-11 | 75 | 75 | 94.0269 | no | 26337766 | low | 1071.9868 | low | 0.027519 | no | 93.3836 | no | 0.95837 | 0.85 | no | 0.9654 | no | 0.95151 | no | 0.95413 | no | 0.49156 | nan | 0 | -0.0006745 | yes | 0 | -0.00014209 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150402_M00766_0087_000000000-ACCAY | M00766 | 2015-04-02 | 151 | 151 | 88.6379 | no | 17224632 | low | 896.3347 | OK | 0.068342 | no | 91.2709 | no | 0.85166 | 0.8 | no | 0.88173 | no | 0.81998 | no | 0.95413 | no | 2.2205e-10 | nan | 0 | -0.0022033 | yes | 5 | -0.0021469 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150414_M00766_0090_000000000-AAU39 | M00766 | 2015-04-14 | 151 | 151 | 86.0407 | no | 14841878 | low | 771.5068 | low | 0.03988 | no | 87.1861 | no | 0.93627 | 0.8 | no | 0.95058 | no | 0.9214 | no | 0.95413 | no | 0.27827 | 0.21791 | 0 | -0.00025771 | no | 0 | -0.00046336 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150427_M02641_0042_000000000-AD8LB | M02641 | 2015-04-27 | 75 | 75 | 97.4766 | no | 4116334 | low | 164.4147 | very low | 0.057123 | no | 97.3182 | no | 0.95813 | 0.85 | no | 0.97251 | no | 0.95269 | no | 0.95413 | no | 1 | 0.43484 | 0 | -0.00032162 | no | 0 | -0.00011879 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150430_M00766_0092_000000000-AD77P | M00766 | 2015-04-30 | 75 | 75 | 86.4782 | no | 33656164 | low | 1420.3657 | high | 0.029215 | no | 86.0606 | no | 0.94127 | 0.85 | no | 0.95639 | no | 0.92495 | no | 0.95413 | no | 2.1517e-05 | 1.5334e-07 | 0 | -0.00053145 | yes | 0 | -0.00060527 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150506_M00766_0094_000000000-AEVP8 | M00766 | 2015-05-06 | 151 | 151 | 89.4984 | no | 20688072 | low | 1090.2056 | high | 0.029888 | no | 89.7217 | no | 0.92458 | 0.8 | no | 0.94568 | no | 0.90199 | no | 0.84095 | no | 4.8432e-13 | nan | 0 | -0.00051215 | yes | 0 | -0.00090332 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150508_M00766_0096_000000000-AF9MW | M00766 | 2015-05-08 | 75 | 75 | 91.4565 | no | 29237117 | low | 1200.3221 | OK | 0.049362 | no | 92.6001 | no | 0.95684 | 0.85 | no | 0.97022 | no | 0.94333 | no | 0.84095 | no | 0.0058945 | nan | 0 | -0.00039203 | no | 0 | -0.00040344 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150508_M02641_0047_000000000-AF7NG | M02641 | 2015-05-08 | 75 | 75 | 89.2921 | no | 33452889 | low | 1376.6665 | high | 0.032357 | no | 90.4595 | no | 0.94491 | 0.85 | no | 0.96109 | no | 0.92874 | no | 0.84095 | no | 5.3189e-05 | nan | 0 | -0.00048487 | no | 0 | -0.00044495 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150511_M02641_0048_000000000-AEUGT | M02641 | 2015-05-11 | 250 | 250 | 25.8223 | yes | 3140165 | low | 125.9118 | very low | 0.11338 | yes | 14.4297 | yes | 0.53707 | 0.7 | yes | 0.38185 | yes | 0.6802 | yes | 0.35842 | no | 2.3353e-41 | 4.6079e-40 | 205 | -0.0023104 | yes | 25 | -0.0033967 | yes | fail | fail | fail | fail | same | pass | exceeded | not exceeded | diff | same
150516_M00766_0099_000000000-AF81F | M00766 | 2015-05-16 | 75 | 75 | 92.0597 | no | 29948630 | low | 1238.4292 | OK | 0.017269 | no | 91.8305 | no | 0.96082 | 0.85 | no | 0.97249 | no | 0.94905 | no | 0.84095 | no | 0.0018358 | 5.1652e-05 | 0 | -0.00036272 | no | 0 | -0.00037743 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150521_M02641_0051_000000000-AF9JH | M02641 | 2015-05-21 | 75 | 75 | 85.4273 | no | 20774323 | low | 847.1236 | low | 0.093549 | yes | 85.9442 | no | 0.92414 | 0.85 | no | 0.93171 | no | 0.91686 | no | 0.99584 | no | 0.99485 | 0.61684 | 0 | -0.00098979 | yes | 0 | -0.00048486 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff
150522_M00766_0103_000000000-AF9M3 | M00766 | 2015-05-22 | 75 | 75 | 88.2546 | no | 33587117 | low | 1405.6196 | high | 0.026305 | no | 88.4186 | no | 0.94466 | 0.85 | no | 0.95853 | no | 0.92962 | no | 0.84095 | no | 0.0015346 | nan | 0 | -0.00056188 | yes | 0 | -0.00050146 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150522_M02641_0052_000000000-AF82F | M02641 | 2015-05-22 | 75 | 75 | 77.7302 | no | 26557402 | low | 1103.8322 | high | 0.045966 | no | 76.9724 | yes | 0.90712 | 0.85 | no | 0.92279 | no | 0.89107 | no | 0.99584 | no | 0.060099 | 5.5163e-09 | 0 | -0.00079849 | yes | 0 | -0.00064071 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150601_M00766_0107_000000000-AF9N8 | M00766 | 2015-06-01 | 75 | 75 | 96.709 | no | 17833767 | low | 714.0694 | very low | 0.036487 | no | 97.1935 | no | 0.97916 | 0.85 | no | 0.98619 | no | 0.97361 | no | 0.84095 | no | 0.47307 | nan | 0 | -0.00024948 | no | 0 | -0.00016826 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150602_M00766_0108_000000000-AFMB8 | M00766 | 2015-06-02 | 200 | 200 | 87.6065 | no | 33648503 | low | 1419.0921 | high | 0.017208 | no | 87.2266 | no | 0.8224 | 0.7 | no | 0.84731 | no | 0.7928 | no | 0.84095 | no | 3.1434e-29 | 0 | 19 | -0.0019091 | yes | 33 | -0.0023749 | yes | pass | pass | pass | pass | same | pass | not exceeded | exceeded | diff | same
150603_M02641_0056_000000000-AEY8R | M02641 | 2015-06-03 | 151 | 151 | 75.779 | no | 10182405 | low | 516.5408 | low | 0.052494 | no | 76.2721 | yes | 0.92128 | 0.8 | no | 0.93202 | no | 0.91084 | no | 0.99584 | no | 0.72835 | nan | 0 | -0.00018443 | no | 0 | -0.00018378 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150604_M00766_0109_000000000-AFJ62 | M00766 | 2015-06-04 | 75 | 75 | 90.2876 | no | 30654786 | low | 1283.025 | high | 0.019566 | no | 90.5316 | no | 0.94854 | 0.85 | no | 0.96577 | no | 0.93018 | no | 0.84095 | no | 6.5656e-06 | nan | 0 | -0.0004699 | no | 0 | -0.00054877 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150604_M02641_0057_000000000-AFL7N | M02641 | 2015-06-04 | 75 | 75 | 83.6851 | no | 20989024 | low | 846.0443 | low | 0.089565 | yes | 87.1228 | no | 0.92879 | 0.85 | no | 0.94412 | no | 0.91514 | no | 0.84095 | no | 0.077245 | nan | 0 | -0.00064783 | yes | 0 | -0.00048886 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff
150605_M00766_0110_000000000-AF80F | M00766 | 2015-06-05 | 75 | 75 | 83.8601 | no | 32100609 | low | 1379.9375 | high | 0.019606 | no | 84.3276 | no | 0.92414 | 0.85 | no | 0.94606 | no | 0.89846 | no | 0.84095 | no | 2.0225e-07 | nan | 0 | -0.00071066 | yes | 0 | -0.00077961 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150610_M00766_0111_000000000-AFMWF | M00766 | 2015-06-10 | 75 | 75 | 89.8899 | no | 32008549 | low | 1328.1731 | high | 0.015813 | no | 89.8802 | no | 0.94912 | 0.85 | no | 0.96373 | no | 0.93447 | no | 0.84095 | no | 0.0020072 | nan | 0 | -0.00052315 | yes | 0 | -0.00050342 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150611_M02641_0060_000000000-AFN2Y | M02641 | 2015-06-11 | 75 | 75 | 76.6898 | no | 23799396 | low | 947.2643 | OK | 0.054797 | no | 75.8247 | yes | 0.88682 | 0.85 | no | 0.90551 | no | 0.86738 | no | 0.99584 | no | 0.046537 | 2.6959e-66 | 0 | -0.001158 | yes | 0 | -0.00077535 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150615_M02641_0062_000000000-AFH7C | M02641 | 2015-06-15 | 75 | 75 | 76.7102 | no | 21607117 | low | 847.9618 | low | 0.051033 | no | 78.099 | yes | 0.89692 | 0.85 | no | 0.91029 | no | 0.88432 | no | 0.99584 | no | 0.37812 | 0.020859 | 0 | -0.00092944 | yes | 0 | -0.00061072 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150619_M02641_0063_000000000-AFN2W | M02641 | 2015-06-19 | 75 | 75 | 86.313 | no | 22890237 | low | 921.1523 | low | 0.031622 | no | 86.5093 | no | 0.93054 | 0.85 | no | 0.94446 | no | 0.91735 | no | 0.95413 | no | 0.013174 | nan | 0 | -0.00066355 | yes | 0 | -0.0006793 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150625_M02641_0065_000000000-AEY2A | M02641 | 2015-06-25 | 151 | 151 | 67.3086 | yes | 7123794 | low | 354.4358 | very low | 0.027403 | no | 65.5091 | yes | 0.90035 | 0.8 | no | 0.9059 | no | 0.89467 | no | 0.99998 | no | 0.37708 | 0.0069801 | 0 | -0.00018719 | no | 0 | -0.00016475 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150703_M00766_0117_000000000-AFN2F | M00766 | 2015-07-03 | 75 | 75 | 89.9328 | no | 28300840 | low | 1173.6259 | OK | 0.036091 | no | 90.4158 | no | 0.94919 | 0.85 | no | 0.95814 | no | 0.93984 | no | 0.84095 | no | 0.81385 | nan | 0 | -0.00062428 | yes | 0 | -0.00041123 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150703_M02641_0001_000000000-AFMW2 | M02641 | 2015-07-03 | 75 | 75 | 92.3482 | no | 29421610 | low | 1208.9627 | OK | 0.019251 | no | 92.6367 | no | 0.95661 | 0.85 | no | 0.96958 | no | 0.94384 | no | 0.84095 | no | 0.001929 | 2.1798e-15 | 0 | -0.0004435 | no | 0 | -0.00035874 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150709_M02641_0004_000000000-AFMPC | M02641 | 2015-07-09 | 75 | 75 | 94.2018 | no | 27443669 | low | 1115.769 | low | 0.036555 | no | 95.4991 | no | 0.96502 | 0.85 | no | 0.9763 | no | 0.95448 | no | 0.84095 | no | 0.016984 | nan | 0 | -0.00039431 | no | 0 | -0.00024497 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150727_M00766_0119_000000000-AF3TM | M00766 | 2015-07-27 | 151 | 151 | 88.8616 | no | 15391937 | low | 825.3634 | low | 0.055595 | no | 88.285 | no | 0.9196 | 0.8 | no | 0.9501 | no | 0.88668 | no | 0.84095 | no | 1.2664e-16 | nan | 0 | -0.00038471 | no | 0 | -0.00086889 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150727_M02641_0007_000000000-AGEW7 | M02641 | 2015-07-27 | 230 | 230 | 91.5639 | no | 31830282 | low | 1309.0338 | high | 0.015312 | no | 91.8906 | no | 0.76141 | 0.7 | no | 0.80125 | no | 0.71555 | no | 0.35842 | no | 5.1559e-54 | 2.0554e-121 | 29 | -0.0026602 | yes | 56 | -0.0035322 | yes | pass | pass | pass | pass | same | pass | exceeded | exceeded | same | same
150807_M00766_0122_000000000-AFMW7 | M00766 | 2015-08-07 | 75 | 75 | 95.8884 | no | 19129362 | low | 771.607 | low | 0.031184 | no | 96.3208 | no | 0.9746 | 0.85 | no | 0.98266 | no | 0.96788 | no | 0.84095 | no | 0.16299 | nan | 0 | -0.00031347 | no | 0 | -0.0002468 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150813_M00766_0124_000000000-AFYEH | M00766 | 2015-08-13 | 75 | 75 | 94.3227 | no | 25145640 | low | 1025.335 | low | 0.021357 | no | 94.2549 | no | 0.96769 | 0.85 | no | 0.97825 | no | 0.95812 | no | 0.84095 | no | 0.0084968 | nan | 0 | -0.00033121 | no | 0 | -0.00034652 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150818_M00766_0125_000000000-AE8BP | M00766 | 2015-08-18 | 151 | 151 | 91.3886 | no | 15351123 | low | 804.6945 | low | 0.058805 | no | 90.8408 | no | 0.92459 | 0.8 | no | 0.94957 | no | 0.89589 | no | 0.84095 | no | 6.1155e-16 | 3.3622e-25 | 0 | -0.00052265 | yes | 0 | -0.0011087 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150826_M00766_0127_000000000-AGKLT | M00766 | 2015-08-26 | 200 | 200 | 93.1048 | no | 24799829 | low | 1017.4002 | low | 0.016826 | no | 93.586 | no | 0.88995 | 0.7 | no | 0.91037 | no | 0.86811 | no | 0.67794 | no | 4.0958e-20 | nan | 0 | -0.0013198 | yes | 0 | -0.0017885 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150904_M00766_0129_000000000-AFNA2 | M00766 | 2015-09-04 | 151 | 151 | 84.0182 | no | 18908531 | low | 1029.9385 | high | 0.024332 | no | 83.7698 | no | 0.81255 | 0.8 | no | 0.86788 | no | 0.75533 | yes | 0.95413 | no | 5.0649e-39 | nan | 0 | -0.0022068 | yes | 9 | -0.0021487 | yes | pass | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same
150910_M00766_0130_000000000-AGJGG | M00766 | 2015-09-10 | 75 | 75 | 88.355 | no | 24463097 | low | 1059.3697 | low | 0.0095777 | no | 88.2815 | no | 0.94136 | 0.85 | no | 0.95905 | no | 0.92225 | no | 0.95413 | no | 1.5264e-06 | 1.3678e-23 | 0 | -0.00056717 | yes | 0 | -0.00075898 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150916_M02641_0025_000000000-AFN33 | M02641 | 2015-09-16 | 151 | 151 | 91.2782 | no | 21293722 | low | 1112.3198 | high | 0.023575 | no | 90.9013 | no | 0.8777 | 0.8 | no | 0.90422 | no | 0.84924 | no | 0.95413 | no | 3.9616e-16 | nan | 0 | -0.0018497 | yes | 5 | -0.0021883 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150917_M02641_0026_000000000-AFN0A | M02641 | 2015-09-17 | 151 | 151 | 91.5269 | no | 20122988 | low | 1048.5468 | OK | 0.02524 | no | 91.3116 | no | 0.86192 | 0.8 | no | 0.89213 | no | 0.82928 | no | 0.84095 | no | 2.1049e-19 | nan | 1 | -0.0022155 | yes | 10 | -0.0026282 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150918_M00766_0133_000000000-AH985 | M00766 | 2015-09-18 | 200 | 200 | 91.3077 | no | 30062610 | low | 1253.19 | OK | 0.031977 | no | 91.5655 | no | 0.84979 | 0.7 | no | 0.88478 | no | 0.81133 | no | 0.67794 | no | 4.0326e-70 | nan | 11 | -0.0017023 | yes | 24 | -0.0021771 | yes | pass | pass | pass | pass | same | pass | not exceeded | exceeded | diff | same
150923_M00766_0134_000000000-AFN32 | M00766 | 2015-09-23 | 151 | 151 | 89.0495 | no | 20113067 | low | 1072.9998 | high | 0.033799 | no | 88.906 | no | 0.86242 | 0.8 | no | 0.89294 | no | 0.82952 | no | 0.95413 | no | 1.619e-18 | nan | 0 | -0.002005 | yes | 9 | -0.0025255 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150924_M02641_0027_000000000-AGJGM | M02641 | 2015-09-24 | 75 | 75 | 95.0401 | no | 24554073 | low | 997.0288 | low | 0.0085367 | no | 95.0918 | no | 0.97292 | 0.85 | no | 0.98196 | no | 0.96511 | no | 0.84095 | no | 0.019748 | nan | 0 | -0.00019383 | no | 0 | -0.00024233 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150925_M00766_0135_000000000-AFL75 | M00766 | 2015-09-25 | 151 | 151 | 97.708 | no | 9080878 | low | 460.5892 | very low | 0.039145 | no | 97.9623 | no | 0.94702 | 0.8 | no | 0.95724 | no | 0.93649 | no | 0.95413 | no | 0.0051842 | nan | 0 | -0.00069048 | yes | 0 | -0.0008509 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150929_M02641_0030_000000000-AFWJ3 | M02641 | 2015-09-29 | 151 | 151 | 95.2167 | no | 14824546 | low | 765.6339 | low | 0.037672 | no | 95.287 | no | 0.94667 | 0.8 | no | 0.96411 | no | 0.92657 | no | 0.84095 | no | 2.2694e-08 | nan | 0 | -0.00044841 | no | 0 | -0.00094835 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150930_M00766_0138_000000000-AH4BR | M00766 | 2015-09-30 | 75 | 75 | 90.8515 | no | 29081266 | low | 1203.4983 | OK | 0.018375 | no | 90.911 | no | 0.93455 | 0.85 | no | 0.91651 | no | 0.95034 | no | 0.99584 | no | 5.9019e-07 | nan | 0 | -0.0013505 | yes | 0 | -0.00035323 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151005_M02641_0034_000000000-AH9ML | M02641 | 2015-10-05 | 200 | 200 | 91.5287 | no | 33902064 | low | 1389.8402 | high | 0.014042 | no | 91.6743 | no | 0.83382 | 0.7 | no | 0.86254 | no | 0.80114 | no | 0.50766 | no | 1.3625e-44 | nan | 23 | -0.0021037 | yes | 30 | -0.0027427 | yes | pass | pass | pass | pass | same | pass | exceeded | exceeded | same | same
151008_M02641_0035_000000000-AFL73 | M02641 | 2015-10-08 | 151 | 151 | 92.9471 | no | 20624078 | low | 1070.1471 | OK | 0.02054 | no | 93.2815 | no | 0.88534 | 0.8 | no | 0.91095 | no | 0.85745 | no | 0.95413 | no | 1.7378e-17 | nan | 0 | -0.0018668 | yes | 5 | -0.0023798 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151013_M02641_0038_000000000-AFLW0 | M02641 | 2015-10-13 | 151 | 151 | 92.2835 | no | 19868115 | low | 1036.6449 | OK | 0.015868 | no | 92.3144 | no | 0.85852 | 0.8 | no | 0.89862 | no | 0.81584 | no | 0.84095 | no | 9.4078e-32 | 1.0701e-96 | 0 | -0.0019454 | yes | 7 | -0.0029862 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151014_M02641_0039_000000000-AFLFV | M02641 | 2015-10-14 | 151 | 151 | 93.1322 | no | 20043314 | low | 1036.9175 | OK | 0.024404 | no | 93.1321 | no | 0.86964 | 0.8 | no | 0.90574 | no | 0.83081 | no | 0.84095 | no | 1.4818e-31 | nan | 0 | -0.0019202 | yes | 8 | -0.0027745 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151015_M00766_0142_000000000-AJD8B | M00766 | 2015-10-15 | 151 | 151 | 88.82 | no | 16091719 | low | 856.7867 | low | 0.030984 | no | 88.9965 | no | 0.9206 | 0.8 | no | 0.95153 | no | 0.88855 | no | 0.67794 | no | 9.7576e-25 | nan | 0 | -0.0004536 | no | 1 | -0.0013313 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151021_M02641_0041_000000000-AGHAA | M02641 | 2015-10-21 | 75 | 75 | 93.3815 | no | 29221047 | low | 1189.4531 | OK | 0.018033 | no | 93.5119 | no | 0.96054 | 0.85 | no | 0.97325 | no | 0.9489 | no | 0.84095 | no | 0.001648 | 0.00019262 | 0 | -0.00043517 | no | 0 | -0.00032505 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151022_M00766_0146_000000000-AFN3B | M00766 | 2015-10-22 | 151 | 151 | 81.4565 | no | 19862739 | low | 1107.5448 | very high | 0.031521 | no | 80.9673 | no | 0.90849 | 0.8 | no | 0.93308 | no | 0.88275 | no | 0.95413 | no | 6.0372e-16 | 1.7055e-184 | 0 | -0.00050949 | yes | 0 | -0.00097367 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
151023_M00766_0147_000000000-AFN37 | M00766 | 2015-10-23 | 151 | 151 | 84.6509 | no | 20693511 | low | 1134.0694 | very high | 0.039199 | no | 85.3046 | no | 0.92707 | 0.8 | no | 0.95123 | no | 0.90247 | no | 0.95413 | no | 6.8289e-18 | nan | 0 | -0.00030046 | no | 0 | -0.00056736 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151028_M02641_0043_000000000-AFL74 | M02641 | 2015-10-28 | 151 | 151 | 96.8353 | no | 15543815 | low | 805.7896 | low | 0.035546 | no | 96.7917 | no | 0.96903 | 0.8 | no | 0.97972 | no | 0.95847 | no | 0.67794 | no | 8.1444e-08 | 1.4081e-14 | 0 | -0.00020911 | no | 0 | -0.0004479 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151030_M00766_0150_000000000-AFLDT | M00766 | 2015-10-30 | 151 | 151 | 89.1258 | no | 19498823 | low | 1043.7728 | OK | 0.022316 | no | 89.1735 | no | 0.94282 | 0.8 | no | 0.95917 | no | 0.92637 | no | 0.95413 | no | 6.2812e-11 | 1.3707e-27 | 0 | -0.00027197 | no | 0 | -0.00043246 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151030_M02641_0044_000000000-AJHLJ | M02641 | 2015-10-30 | 75 | 75 | 89.5684 | no | 33168068 | low | 1381.3797 | high | 0.015021 | no | 89.5367 | no | 0.94181 | 0.85 | no | 0.96027 | no | 0.92327 | no | 0.84095 | no | 1.9212e-09 | 1.7277e-37 | 0 | -0.00053639 | yes | 0 | -0.00046394 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151103_M02641_0046_000000000-AJ5Y7 | M02641 | 2015-11-03 | 151 | 151 | 96.1111 | no | 17192399 | low | 883.2398 | low | 0.021428 | no | 96.0917 | no | 0.96652 | 0.8 | no | 0.97476 | no | 0.95884 | no | 0.95413 | no | 0.00063959 | nan | 0 | -0.00027215 | no | 0 | -0.00034016 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151105_M02641_0047_000000000-AJF4E | M02641 | 2015-11-05 | 151 | 151 | 96.0812 | no | 17316625 | low | 888.5653 | low | 0.024517 | no | 96.1753 | no | 0.96613 | 0.8 | no | 0.97668 | no | 0.95675 | no | 0.84095 | no | 7.3533e-09 | 2.5766e-25 | 0 | -0.00020337 | no | 0 | -0.00027243 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151116_M02641_0050_000000000-AJ6L3 | M02641 | 2015-11-16 | 151 | 151 | 95.1303 | no | 13144452 | low | 672.6944 | low | 0.074365 | no | 94.8898 | no | 0.95687 | 0.8 | no | 0.96857 | no | 0.94432 | no | 0.95413 | no | 0.0016197 | 1.4817e-06 | 0 | -0.00029535 | no | 0 | -0.00053054 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151117_M00766_0155_000000000-AJ6LL | M00766 | 2015-11-17 | 151 | 151 | 87.0005 | no | 16509092 | low | 884.2256 | OK | 0.024563 | no | 86.7919 | no | 0.86377 | 0.8 | no | 0.9042 | no | 0.821 | no | 0.84095 | no | 8.628e-16 | nan | 0 | -0.0012235 | yes | 1 | -0.00063026 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151117_M02641_0051_000000000-AJF4A | M02641 | 2015-11-17 | 151 | 151 | 94.2241 | no | 17785498 | low | 905.3067 | low | 0.03067 | no | 94.2545 | no | 0.88313 | 0.8 | no | 0.91617 | no | 0.84833 | no | 0.84095 | no | 7.9592e-12 | nan | 0 | -0.0012532 | yes | 0 | -0.00072138 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151125_M02641_0053_000000000-AJJ3G | M02641 | 2015-11-25 | 151 | 151 | 96.6504 | no | 3965190 | low | 201.2481 | very low | 0.02812 | no | 96.9381 | no | 0.97423 | 0.8 | no | 0.98297 | no | 0.96916 | no | 0.67794 | no | 0.069974 | 0.0030109 | 0 | -0.00019201 | no | 0 | -0.00032396 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
151126_M00766_0159_000000000-AJD7L | M00766 | 2015-11-26 | 75 | 75 | 88.0477 | no | 24133209 | low | 1032.2091 | low | 0.017289 | no | 87.5838 | no | 0.95346 | 0.85 | no | 0.96296 | no | 0.94323 | no | 0.84095 | no | 0.093829 | 1.6996e-15 | 0 | -0.0004886 | no | 0 | -0.00047968 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151126_M02641_0054_000000000-AJD7T | M02641 | 2015-11-26 | 75 | 75 | 94.0217 | no | 24949166 | low | 1025.4829 | low | 0.029656 | no | 93.8083 | no | 0.96822 | 0.85 | no | 0.97532 | no | 0.96185 | no | 0.95413 | no | 0.1922 | nan | 0 | -0.00030117 | no | 0 | -0.00033189 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151127_M00766_0161_000000000-AJERP | M00766 | 2015-11-27 | 75 | 75 | 90.539 | no | 24073552 | low | 1002.8265 | low | 0.022959 | no | 90.0889 | no | 0.95458 | 0.85 | no | 0.96258 | no | 0.94625 | no | 0.95413 | no | 0.1243 | nan | 0 | -0.00057536 | yes | 0 | -0.00051622 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151201_M02641_0055_000000000-AJ6B2 | M02641 | 2015-12-01 | 151 | 151 | 93.147 | no | 20354768 | low | 1062.1404 | OK | 0.019929 | no | 93.6277 | no | 0.95729 | 0.8 | no | 0.96997 | no | 0.94485 | no | 0.95413 | no | 2.0042e-09 | nan | 0 | -0.00022088 | no | 0 | -0.0003646 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151203_M00766_0165_000000000-AJFAA | M00766 | 2015-12-03 | 151 | 151 | 86.1622 | no | 16734072 | low | 890.0657 | OK | 0.048269 | no | 84.2361 | no | 0.92678 | 0.8 | no | 0.94416 | no | 0.90681 | no | 0.95413 | no | 1.2134e-06 | nan | 0 | -0.0003637 | no | 0 | -0.00065096 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151208_M00766_0167_000000000-AJ5JR | M00766 | 2015-12-08 | 151 | 151 | 97.4409 | no | 5807543 | low | 293.2064 | very low | 0.036817 | no | 97.4486 | no | 0.96505 | 0.8 | no | 0.973 | no | 0.95786 | no | 0.95413 | no | 0.12937 | nan | 0 | -0.00036542 | no | 0 | -0.00043496 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
151210_M02641_0058_000000000-AJ5YJ | M02641 | 2015-12-10 | 151 | 151 | 91.7526 | no | 18225198 | low | 951.6353 | OK | 0.049608 | no | 91.4934 | no | 0.93902 | 0.8 | no | 0.95565 | no | 0.92066 | no | 0.95413 | no | 6.5736e-05 | nan | 0 | -0.00040902 | no | 0 | -0.00072147 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151216_M00766_0170_000000000-ALEAF | M00766 | 2015-12-16 | 151 | 151 | 93.7279 | no | 17003866 | low | 880.7103 | low | 0.036745 | no | 93.9642 | no | 0.94991 | 0.8 | no | 0.96257 | no | 0.937 | no | 0.95413 | no | 4.5633e-06 | nan | 0 | -0.00035203 | no | 0 | -0.00057532 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151217_M02641_0059_000000000-AJFAD | M02641 | 2015-12-17 | 151 | 151 | 58.3542 | yes | 4092271 | low | 204.0891 | very low | 0.28007 | yes | 55.6594 | yes | 0.92478 | 0.8 | no | 0.94271 | no | 0.91061 | no | 0.84095 | no | 0.016952 | 0.00029975 | 0 | -0.00041346 | no | 0 | -0.00076263 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff
151223_M00766_0172_000000000-AJRMN | M00766 | 2015-12-23 | 75 | 75 | 92.6132 | no | 23635202 | low | 977.0359 | low | 0.01686 | no | 92.2318 | no | 0.94189 | 0.85 | no | 0.9657 | no | 0.91522 | no | 0.50766 | no | 8.2817e-14 | nan | 0 | -0.00070693 | yes | 1 | -0.0018884 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151223_M02641_0060_000000000-AGHV2 | M02641 | 2015-12-23 | 75 | 75 | 87.711 | no | 26108005 | low | 1077.9707 | OK | 0.015686 | no | 87.8311 | no | 0.92979 | 0.85 | no | 0.94545 | no | 0.911 | no | 0.95413 | no | 1.0014e-06 | 6.7599e-86 | 0 | -0.00093076 | yes | 0 | -0.0011369 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151231_M02641_0061_000000000-AJD8L | M02641 | 2015-12-31 | 75 | 75 | 88.6234 | no | 30986761 | low | 1292.1541 | high | 0.043562 | no | 89.278 | no | 0.93779 | 0.85 | no | 0.95903 | no | 0.91501 | no | 0.95413 | no | 2.22e-08 | nan | 0 | -0.00055157 | yes | 0 | -0.00062885 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160107_M00766_0175_000000000-AH6C3 | M00766 | 2016-01-07 | 75 | 75 | 87.9117 | no | 26044869 | low | 1123.4126 | OK | 0.012327 | no | 86.8229 | no | 0.95101 | 0.85 | no | 0.96122 | no | 0.9395 | no | 0.95413 | no | 0.056283 | nan | 0 | -0.00056289 | yes | 0 | -0.00070024 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160108_M00766_0176_000000000-AJDAJ | M00766 | 2016-01-08 | 75 | 75 | 86.9378 | no | 24436074 | low | 1042.2918 | low | 0.023113 | no | 87.1996 | no | 0.79786 | 0.85 | yes | 0.93046 | no | 0.6482 | yes | 0.15463 | no | 1.2164e-131 | nan | 0 | -0.0018269 | yes | 36 | -0.0074037 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same
160114_M00766_0177_000000000-ALYTH | M00766 | 2016-01-14 | 75 | 75 | 87.5409 | no | 27732794 | low | 1180.5507 | OK | 0.016589 | no | 87.6219 | no | 0.94367 | 0.85 | no | 0.95391 | no | 0.93177 | no | 0.95413 | no | 0.043231 | nan | 0 | -0.00066231 | yes | 0 | -0.00077358 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160114_M02641_0065_000000000-ALRU0 | M02641 | 2016-01-14 | 75 | 75 | 91.58 | no | 29247440 | low | 1214.9636 | OK | 0.024503 | no | 91.5468 | no | 0.95131 | 0.85 | no | 0.96723 | no | 0.93592 | no | 0.84095 | no | 7.4878e-05 | nan | 0 | -0.00049951 | no | 0 | -0.00034584 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160128_M02641_0070_000000000-ALAGE | M02641 | 2016-01-28 | 75 | 75 | 91.826 | no | 24305600 | low | 1013.4527 | low | 0.0159 | no | 91.7655 | no | 0.96174 | 0.85 | no | 0.96994 | no | 0.95369 | no | 0.95413 | no | 0.073098 | nan | 0 | -0.0003153 | no | 0 | -0.00042581 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160129_M00766_0003_000000000-ALAH0 | M00766 | 2016-01-29 | 75 | 75 | 86.7633 | no | 32429499 | low | 1380.9019 | high | 0.017673 | no | 87.2761 | no | 0.93918 | 0.85 | no | 0.95507 | no | 0.92174 | no | 0.84095 | no | 1.4395e-05 | nan | 0 | -0.00058412 | yes | 0 | -0.00060062 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160129_M02641_0071_000000000-AMC8N | M02641 | 2016-01-29 | 200 | 200 | 91.7102 | no | 25390554 | low | 1054.7902 | low | 0.01824 | no | 91.6301 | no | 0.86183 | 0.7 | no | 0.88102 | no | 0.83939 | no | 0.67794 | no | 1.3683e-14 | 6.476e-15 | 1 | -0.0017129 | yes | 0 | -0.0020671 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160206_M00766_0005_000000000-ALYU1 | M00766 | 2016-02-06 | 75 | 75 | 84.6656 | no | 32152720 | low | 1381.1846 | high | 0.017535 | no | 83.6833 | no | 0.93615 | 0.85 | no | 0.9506 | no | 0.92023 | no | 0.84095 | no | 0.0023058 | 1.3476e-13 | 0 | -0.00053805 | yes | 0 | -0.00042151 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
160206_M02641_0074_000000000-AL604 | M02641 | 2016-02-06 | 151 | 151 | 88.8163 | no | 18930346 | low | 1014.3724 | OK | 0.015119 | no | 88.4642 | no | 0.92173 | 0.8 | no | 0.9483 | no | 0.89269 | no | 0.67794 | no | 2.5587e-23 | 1.4688e-132 | 0 | -0.00040622 | no | 0 | -0.00095862 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
160209_M00766_0007_000000000-AMERJ | M00766 | 2016-02-09 | 75 | 75 | 79.4588 | no | 34759702 | low | 1528.9595 | very high | 0.016677 | no | 80.1993 | no | 0.90703 | 0.85 | no | 0.93282 | no | 0.87982 | no | 0.84095 | no | 1.1867e-08 | 0 | 0 | -0.00081183 | yes | 0 | -0.00073619 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
160216_M00766_0009_000000000-AL73W | M00766 | 2016-02-16 | 151 | 151 | 92.117 | no | 18300588 | low | 961.5089 | OK | 0.025014 | no | 92.5591 | no | 0.94116 | 0.8 | no | 0.96156 | no | 0.91989 | no | 0.84095 | no | 1.5361e-15 | 3.496e-24 | 0 | -0.00029836 | no | 0 | -0.00066338 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
160219_M00766_0011_000000000-AL6GM | M00766 | 2016-02-19 | 131 | 131 | 5.6098 | yes | 661574 | low | 22.0869 | very low | 0.9623 | yes | 9.7324 | yes | 0.77871 | 0.8 | yes | 0.79992 | yes | 0.78598 | yes | 0.99584 | no | 1 | 0.73572 | 5 | -0.00056401 | yes | 5 | -0.00059171 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
160220_M02641_0081_000000000-AMETF | M02641 | 2016-02-20 | 200 | 200 | 92.9979 | no | 23074941 | low | 958.5024 | low | 0.027553 | no | 93.6783 | no | 0.87882 | 0.7 | no | 0.90362 | no | 0.85132 | no | 0.35842 | no | 6.4948e-26 | nan | 0 | -0.0012552 | yes | 0 | -0.0017992 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160223_M00766_0012_000000000-AL5YC | M00766 | 2016-02-23 | 151 | 151 | 93.2958 | no | 18111084 | low | 951.0894 | OK | 0.019086 | no | 93.3336 | no | 0.94516 | 0.8 | no | 0.96589 | no | 0.92393 | no | 0.84095 | no | 9.6635e-16 | 3.8603e-52 | 0 | -0.0002968 | no | 0 | -0.0005917 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
| 76 |  |  | NUM FAILS |  | 7 |  | 123 |  | 25 |  | 11 |  | 13 |  |  | 10 |  | 4 |  | 15 |  | 0 |  |  | 8 |  | 77 | 39 |  | 85 | 36 | 40 | 23 | 28 | 11 | 0 |  |  | 11 | 24
|  |  |  | % FAILS |  | 5.6911 |  | 100 |  | 20.3252 |  | 8.9431 |  | 10.5691 |  |  | 8.1301 |  | 3.252 |  | 12.1951 |  | 0 |  |  | 6.5041 |  | 62.6016 | 31.7073 |  | 69.1057 | 87 | 83 | 100 | 95
Exclude this one as useless
[32 empty] | Method 1 | Method 2 | Method 1 (no CD) | Method 2 (no CD)
[31 empty] | Overall | 86 | 82 | 99 | 94
[31 empty] | Pass | 36 | 40 | 23 | 28
Fail

## AnalyDataCD
RunID | InstrumentID | RunDate | NumberOfCyclesR1 | NumberofCyclesR2 | %ReadsPassingFilter | %ReadsPassingFilterOutsideThreshold? | NumberofReadsPassingFilter | ReadCountWithinThreshold? | MeanClusterDensityPassingFilter | MeanClusterDensityWithinThreshold | COVClusterDensity | COVClusterDensityExceedsThreshold? | %MedianClusterDensityPassingFilter | GapBetweenClusterDensitiesPassingFilterAndNotExceedsThreshold? | ProportionOverQ30 | IlluminaQ30Threshold | ProportionOverQ30ExceedsThreshold? | ProportionOverQ30Read1 | ProportionOverQ30ExceedsThresholdRead1 | ProportionOverQ30Read2 | ProportionOverQ30ExceedsThresholdRead2 | K-SPValue | K-SPValueSignificant? | K-SUnpackPvalue | M-WUnpackPvalue | NumberOfCyclesExceedingThresholdUnderQ30Read1 | SlopeOfLineOfBestFitRead1 | SlopeOfLineOfBestFitRead1ExceedsThreshold? | NumberOfCyclesExceedingThresholdUnderQ30Read2 | SlopeOfLineOfBestFitRead2 | SlopeOfLineOfBestFitRead2ExceedsThreshold? | Fail? | Fail2? | Fail? No CD | Fail2? No CD | R1vsR2PropQ30 | R1vsR2K-S | NumCycleR1Threshold | NumCycleR2Threshold | R1vsR2CyclesQ30 | R1vsR2bestfit
121009_M00766_0002_000000000-A1U6P | M00766 | 2012-10-09 | 151 | 151 | 45.385 | yes | 3138050 | low | 341.3139 | low | 0.34653 | yes | 0 | yes | 0.87575 | 0.8 | no | 0.87642 | no | 0.87507 | no | 0.99998 | no | 0.99997 | 0.80608 | 1 | 0 | no | 0 | 0 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
130206_M00766_0002_000000000-A23JM | M00766 | 2013-02-06 | 250 | 250 | 2.0076 | yes | 160923 | low | 10.2917 | low | 0.26427 | yes | 1.9741 | yes | 0.27654 | 0.75 | yes | 0.091188 | yes | 0.46189 | yes | 0.35842 | no | 0.0011204 | 0.012662 | 193 | -0.0010965 | yes | 65 | -0.0032351 | yes | fail | fail | fail | fail | same | pass | exceeded | exceeded | same | same
130618_M00766_0019_000000000-A4FEU | M00766 | 2013-06-18 | 251 | 251 | 93.703 | no | 10147864 | low | 556.7555 | low | 0.009214 | no | 93.9257 | no | 0.79545 | 0.75 | no | 0.90383 | no | 0.68024 | yes | 0.35842 | no | 1.6395e-83 | 9.0394e-306 | 0 | -0.0010825 | yes | 47 | -0.0037713 | yes | pass | fail | pass | fail | diff | pass | not exceeded | exceeded | diff | same
130624_M00766_0021_000000000-A53PT | M00766 | 2013-06-24 | 251 | 251 | 88.7381 | no | 17725761 | low | 940.8335 | OK | 0.011629 | no | 89.0017 | no | 0.84736 | 0.75 | no | 0.90275 | no | 0.78576 | no | 0.67794 | no | 1.0912e-52 | 1.2366e-63 | 0 | -0.0009554 | yes | 4 | -0.0019889 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
130712_M00766_0024_000000000-A5AVV | M00766 | 2013-07-12 | 151 | 151 | 82.4401 | no | 13180142 | low | 710.2954 | low | 0.09422 | yes | 82.1308 | no | 0.92195 | 0.8 | no | 0.92377 | no | 0.92292 | no | 0.99998 | no | 0.0096055 | 4.6381e-11 | 0 | -0.00059232 | yes | 0 | -0.00062141 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
130817_M00766_0037_000000000-A4EDD | M00766 | 2013-08-17 | 176 | 176 | 93.613 | no | 14493580 | low | 748.7721 | low | 0.020101 | no | 93.6127 | no | 0.94897 | 0.75 | no | 0.96312 | no | 0.93388 | no | 0.84095 | no | 4.6236e-05 | nan | 0 | -0.00039473 | no | 1 | -0.0006017 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
130913_M00766_0049_000000000-A5B0E | M00766 | 2013-09-13 | 251 | 251 | 89.8636 | no | 17433281 | low | 913.7303 | OK | 0.024696 | no | 89.1779 | no | 0.83393 | 0.75 | no | 0.88573 | no | 0.77648 | no | 0.84095 | no | 3.8371e-42 | 5.6819e-66 | 0 | -0.0011656 | yes | 4 | -0.0019759 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
130916_M00766_0050_000000000-A5M1U | M00766 | 2013-09-16 | 151 | 151 | 86.0618 | no | 22033351 | low | 1185.4173 | very high | 0.028468 | no | 86.0122 | no | 0.9138 | 0.8 | no | 0.93477 | no | 0.89056 | no | 0.84095 | no | 4.3159e-13 | 1.1035e-23 | 0 | -0.00070235 | yes | 0 | -0.0009045 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
130924_M00766_0052_000000000-A5BE6 | M00766 | 2013-09-24 | 251 | 251 | 67.1952 | yes | 14006895 | low | 805.1388 | high | 0.092176 | yes | 66.8443 | yes | 0.71517 | 0.75 | yes | 0.77069 | no | 0.66348 | yes | 0.99584 | no | 6.3637e-20 | nan | 16 | -0.0019789 | yes | 32 | -0.002186 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same
131007_M00766_0055_000000000-A59JT | M00766 | 2013-10-07 | 251 | 251 | 88.9869 | no | 7066429 | low | 464.0561 | low | 0.11395 | yes | 90.7266 | no | 0.86005 | 0.75 | no | 0.86612 | no | 0.85247 | no | 1 | no | 0.071573 | nan | 1 | -0.0014022 | yes | 1 | -0.0015446 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
131025_M00766_0059_000000000-A5P9A | M00766 | 2013-10-25 | 151 | 151 | 90.6341 | no | 17754371 | low | 930.3591 | OK | 0.030836 | no | 91.4036 | no | 0.92242 | 0.8 | no | 0.93549 | no | 0.90546 | no | 0.99584 | no | 1.9427e-05 | nan | 0 | -0.00094972 | yes | 0 | -0.0010274 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
140207_M00766_0022_000000000-A7PH8 | M00766 | 2014-02-07 | 251 | 251 | 88.3301 | no | 10785103 | low | 558.1003 | low | 0.053446 | no | 89.8657 | no | 0.85142 | 0.75 | no | 0.8548 | no | 0.84338 | no | 1 | no | 0.42818 | nan | 1 | -0.0014636 | yes | 1 | -0.0016776 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
140424_M00766_0033_000000000-A7BNA | M00766 | 2014-04-24 | 151 | 151 | 97.0072 | no | 5094671 | low | 251.0711 | very low | 0.231 | yes | 98.5469 | no | 0.93672 | 0.8 | no | 0.94645 | no | 0.92545 | no | 0.95413 | no | 0.083102 | 0.0055314 | 0 | -0.00096088 | yes | 0 | -0.0010359 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
140514_M00766_0036_000000000-A7BRK | M00766 | 2014-05-14 | 151 | 151 | 81.4684 | no | 20334208 | low | 1124.4827 | very high | 0.017101 | no | 81.2844 | no | 0.79665 | 0.8 | yes | 0.83677 | no | 0.75502 | yes | 0.99584 | no | 4.1492e-14 | nan | 0 | -0.0024166 | yes | 7 | -0.0018547 | yes | fail | fail | fail | fail | diff | pass | not exceeded | not exceeded | same | same
140612_M00766_0039_000000000-A7BNL | M00766 | 2014-06-12 | 251 | 251 | 91.373 | no | 20643048 | low | 1088.1751 | OK | 0.014529 | no | 90.3695 | no | 0.78766 | 0.75 | no | 0.82619 | no | 0.74426 | yes | 0.84095 | no | 1.1316e-41 | nan | 19 | -0.0022137 | yes | 39 | -0.0026621 | yes | pass | fail | pass | fail | diff | pass | not exceeded | exceeded | diff | same
140616_M00766_0040_000000000-A78V9 | M00766 | 2014-06-16 | 251 | 251 | 97.3311 | no | 10710545 | low | 534.5911 | low | 0.0065097 | no | 97.505 | no | 0.81498 | 0.75 | no | 0.8406 | no | 0.7849 | no | 0.95413 | no | 7.2918e-10 | nan | 1 | -0.0024062 | yes | 30 | -0.0029712 | yes | pass | pass | pass | pass | same | pass | not exceeded | exceeded | diff | same
140813_M00766_0047_000000000-A8PJL | M00766 | 2014-08-13 | 151 | 151 | 91.9352 | no | 17936069 | low | 943.905 | OK | 0.026213 | no | 91.9798 | no | 0.92465 | 0.8 | no | 0.94352 | no | 0.90463 | no | 0.95413 | no | 3.3998e-09 | 3.0143e-185 | 0 | -0.00068224 | yes | 0 | -0.00075951 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
140901_M00766_0048_000000000-AA63M | M00766 | 2014-09-01 | 151 | 151 | 94.0673 | no | 17640397 | low | 909.8289 | low | 0.02007 | no | 93.8944 | no | 0.93862 | 0.8 | no | 0.96072 | no | 0.91538 | no | 0.84095 | no | 4.2142e-16 | 6.6658e-224 | 0 | -0.00044115 | no | 0 | -0.00052041 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
141014_M00766_0053_000000000-A8R6M | M00766 | 2014-10-14 | 151 | 151 | 89.6387 | no | 19084469 | low | 1011.4321 | OK | 0.014006 | no | 90.0342 | no | 0.89321 | 0.8 | no | 0.91865 | no | 0.86561 | no | 0.95413 | no | 8.0153e-17 | 2.8333e-171 | 0 | -0.0010561 | yes | 0 | -0.0011248 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
141017_M02641_0022_000000000-AA66H | M02641 | 2014-10-17 | 151 | 151 | 79.2948 | no | 22430099 | low | 1210.1492 | very high | 0.017994 | no | 78.9705 | yes | 0.87741 | 0.8 | no | 0.91601 | no | 0.83648 | no | 0.99584 | no | 1.1016e-30 | nan | 0 | -0.00063613 | yes | 0 | -0.00060429 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
141024_M00766_0056_000000000-A8PC5 | M00766 | 2014-10-24 | 151 | 151 | 86.6935 | no | 22517755 | low | 1216.7323 | very high | 0.017442 | no | 86.7307 | no | 0.88873 | 0.8 | no | 0.92881 | no | 0.84641 | no | 0.84095 | no | 1.3784e-37 | 2.2642e-64 | 0 | -0.00069708 | yes | 0 | -0.00081119 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
141031_M00766_0058_000000000-AA3GN | M00766 | 2014-10-31 | 75 | 75 | 84.8049 | no | 22830617 | low | 1020.36 | OK | 0.066963 | no | 85.4852 | no | 0.94188 | 0.85 | no | 0.94357 | no | 0.93798 | no | 0.99584 | no | 0.16727 | nan | 0 | -0.00072598 | yes | 0 | -0.0003319 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
141107_M00766_0060_000000000-AA8PM | M00766 | 2014-11-07 | 151 | 151 | 90.0798 | no | 11262793 | low | 579.4485 | low | 0.034966 | no | 90.2121 | no | 0.96028 | 0.8 | no | 0.96678 | no | 0.9539 | no | 0.99584 | no | 0.99203 | 0.8978 | 0 | -0.00013806 | no | 0 | -0.00011747 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
141118_M00766_0061_000000000-A8P8J | M00766 | 2014-11-18 | 151 | 151 | 92.7322 | no | 19181201 | low | 1009.9074 | OK | 0.027961 | no | 92.312 | no | 0.93572 | 0.8 | no | 0.94773 | no | 0.92218 | no | 0.95413 | no | 1.3904e-05 | nan | 0 | -0.0006231 | yes | 0 | -0.00072415 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
141208_M00766_0063_000000000-A8P7C | M00766 | 2014-12-08 | 151 | 151 | 90.051 | no | 21795042 | low | 1153.0041 | high | 0.020941 | no | 89.941 | no | 0.89826 | 0.8 | no | 0.94464 | no | 0.85869 | no | 0.67794 | no | 1.8648e-37 | nan | 0 | -0.00060048 | yes | 3 | -0.00051121 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
141208_M02641_0026_000000000-A8R55 | M02641 | 2014-12-08 | 151 | 151 | 94.9912 | no | 13800416 | low | 717.5691 | low | 0.027596 | no | 95.1078 | no | 0.84608 | 0.8 | no | 0.94004 | no | 0.74763 | yes | 0.35842 | no | 7.7537e-147 | 6.3912e-157 | 0 | -0.00078235 | yes | 5 | -0.0011709 | yes | pass | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same
141212_M00766_0064_000000000-ACCEB | M00766 | 2014-12-12 | 151 | 151 | 81.1933 | no | 20826514 | low | 1127.546 | very high | 0.021875 | no | 81.0162 | no | 0.74058 | 0.8 | yes | 0.80816 | no | 0.67993 | yes | 0.99584 | no | 3.8905e-34 | nan | 0 | -0.0028592 | yes | 24 | -0.0015004 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same
141216_M00766_0065_000000000-ACCDT | M00766 | 2014-12-16 | 151 | 151 | 79.0858 | no | 20548879 | low | 1120.7265 | very high | 0.02252 | no | 79.0848 | yes | 0.77398 | 0.8 | yes | 0.84162 | no | 0.70405 | yes | 0.84095 | no | 1.4032e-49 | 7.9356e-94 | 0 | -0.0021868 | yes | 16 | -0.00079947 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same
141219_M00766_0066_000000000-ACCB1 | M00766 | 2014-12-19 | 151 | 151 | 84.1075 | no | 16554828 | low | 882.4454 | OK | 0.030188 | no | 83.544 | no | 0.71631 | 0.8 | yes | 0.77794 | yes | 0.66061 | yes | 0.99998 | no | 9.0168e-19 | 1.7886e-30 | 0 | -0.0034797 | yes | 18 | -0.0018779 | yes | fail | fail | fail | fail | same | pass | not exceeded | exceeded | diff | same
150120_M00766_0071_000000000-AA63K | M00766 | 2015-01-20 | 151 | 151 | 85.9694 | no | 14794255 | low | 793.1929 | low | 0.05599 | no | 87.4592 | no | 0.90779 | 0.8 | no | 0.93487 | no | 0.89285 | no | 0.95413 | no | 0.01591 | nan | 0 | -0.00060834 | yes | 3 | -0.00064423 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150127_M02641_0027_000000000-AA65J | M02641 | 2015-01-27 | 151 | 151 | 94.4613 | no | 15926233 | low | 820.6824 | low | 0.035198 | no | 94.2959 | no | 0.93014 | 0.8 | no | 0.95953 | no | 0.9008 | no | 0.67794 | no | 1.0379e-16 | 3.674e-28 | 0 | -0.0004646 | no | 0 | -0.00058277 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150203_M00766_0074_000000000-AAUMH | M00766 | 2015-02-03 | 151 | 151 | 81.6068 | no | 16649356 | low | 884.3212 | OK | 0.052428 | no | 81.2451 | no | 0.90342 | 0.8 | no | 0.92886 | no | 0.88527 | no | 0.95413 | no | 0.0090932 | 0.028862 | 0 | -0.00047881 | no | 3 | -0.00020141 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150203_M02641_0028_000000000-ACBYG | M02641 | 2015-02-03 | 151 | 151 | 89.643 | no | 18845203 | low | 988.6001 | OK | 0.014743 | no | 90.0779 | no | 0.82841 | 0.8 | no | 0.86795 | no | 0.78849 | yes | 0.99584 | no | 1.2512e-16 | 0 | 0 | -0.0022305 | yes | 3 | -0.0018175 | yes | pass | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same
150205_M00766_0075_000000000-ACC43 | M00766 | 2015-02-05 | 151 | 151 | 93.254 | no | 17676890 | low | 914.0169 | low | 0.038617 | no | 93.3074 | no | 0.8595 | 0.8 | no | 0.89452 | no | 0.83021 | no | 0.95413 | no | 5.2726e-15 | 3.8357e-55 | 0 | -0.0019394 | yes | 4 | -0.0017318 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150224_M00766_0078_000000000-ACMP5 | M00766 | 2015-02-24 | 75 | 75 | 82.4896 | no | 36080266 | low | 1546.85 | very high | 0.024283 | no | 82.282 | no | 0.93168 | 0.85 | no | 0.94868 | no | 0.91268 | no | 0.84095 | no | 1.1086e-05 | 0 | 0 | -0.00047772 | no | 0 | -0.00047646 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150309_M02641_0035_000000000-ACC3J | M02641 | 2015-03-09 | 151 | 151 | 88.8745 | no | 22186253 | low | 1157.1795 | very high | 0.017088 | no | 89.0446 | no | 0.79385 | 0.8 | yes | 0.83769 | no | 0.74848 | yes | 0.99584 | no | 2.3483e-15 | nan | 0 | -0.002745 | yes | 7 | -0.0021184 | yes | fail | fail | fail | fail | diff | pass | not exceeded | not exceeded | same | same
150311_M00766_0081_000000000-ACN0W | M00766 | 2015-03-11 | 75 | 75 | 92.0795 | no | 28700881 | low | 1177.6318 | OK | 0.016158 | no | 92.5321 | no | 0.96191 | 0.85 | no | 0.9725 | no | 0.95082 | no | 0.84095 | no | 0.010916 | nan | 0 | -0.00034381 | no | 0 | -0.00028909 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150311_M02641_0036_000000000-AD8VB | M02641 | 2015-03-11 | 75 | 75 | 94.0269 | no | 26337766 | low | 1071.9868 | low | 0.027519 | no | 93.3836 | no | 0.95837 | 0.85 | no | 0.9654 | no | 0.95151 | no | 0.95413 | no | 0.49156 | nan | 0 | -0.0006745 | yes | 0 | -0.00014209 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150402_M00766_0087_000000000-ACCAY | M00766 | 2015-04-02 | 151 | 151 | 88.6379 | no | 17224632 | low | 896.3347 | OK | 0.068342 | no | 91.2709 | no | 0.85166 | 0.8 | no | 0.88173 | no | 0.81998 | no | 0.95413 | no | 2.2205e-10 | nan | 0 | -0.0022033 | yes | 5 | -0.0021469 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150414_M00766_0090_000000000-AAU39 | M00766 | 2015-04-14 | 151 | 151 | 86.0407 | no | 14841878 | low | 771.5068 | low | 0.03988 | no | 87.1861 | no | 0.93627 | 0.8 | no | 0.95058 | no | 0.9214 | no | 0.95413 | no | 0.27827 | 0.21791 | 0 | -0.00025771 | no | 0 | -0.00046336 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150427_M02641_0042_000000000-AD8LB | M02641 | 2015-04-27 | 75 | 75 | 97.4766 | no | 4116334 | low | 164.4147 | very low | 0.057123 | no | 97.3182 | no | 0.95813 | 0.85 | no | 0.97251 | no | 0.95269 | no | 0.95413 | no | 1 | 0.43484 | 0 | -0.00032162 | no | 0 | -0.00011879 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
150430_M00766_0092_000000000-AD77P | M00766 | 2015-04-30 | 75 | 75 | 86.4782 | no | 33656164 | low | 1420.3657 | high | 0.029215 | no | 86.0606 | no | 0.94127 | 0.85 | no | 0.95639 | no | 0.92495 | no | 0.95413 | no | 2.1517e-05 | 1.5334e-07 | 0 | -0.00053145 | yes | 0 | -0.00060527 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150506_M00766_0094_000000000-AEVP8 | M00766 | 2015-05-06 | 151 | 151 | 89.4984 | no | 20688072 | low | 1090.2056 | high | 0.029888 | no | 89.7217 | no | 0.92458 | 0.8 | no | 0.94568 | no | 0.90199 | no | 0.84095 | no | 4.8432e-13 | nan | 0 | -0.00051215 | yes | 0 | -0.00090332 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150508_M00766_0096_000000000-AF9MW | M00766 | 2015-05-08 | 75 | 75 | 91.4565 | no | 29237117 | low | 1200.3221 | OK | 0.049362 | no | 92.6001 | no | 0.95684 | 0.85 | no | 0.97022 | no | 0.94333 | no | 0.84095 | no | 0.0058945 | nan | 0 | -0.00039203 | no | 0 | -0.00040344 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150508_M02641_0047_000000000-AF7NG | M02641 | 2015-05-08 | 75 | 75 | 89.2921 | no | 33452889 | low | 1376.6665 | high | 0.032357 | no | 90.4595 | no | 0.94491 | 0.85 | no | 0.96109 | no | 0.92874 | no | 0.84095 | no | 5.3189e-05 | nan | 0 | -0.00048487 | no | 0 | -0.00044495 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150511_M02641_0048_000000000-AEUGT | M02641 | 2015-05-11 | 250 | 250 | 25.8223 | yes | 3140165 | low | 125.9118 | very low | 0.11338 | yes | 14.4297 | yes | 0.53707 | 0.7 | yes | 0.38185 | yes | 0.6802 | yes | 0.35842 | no | 2.3353e-41 | 4.6079e-40 | 205 | -0.0023104 | yes | 25 | -0.0033967 | yes | fail | fail | fail | fail | same | pass | exceeded | not exceeded | diff | same
150516_M00766_0099_000000000-AF81F | M00766 | 2015-05-16 | 75 | 75 | 92.0597 | no | 29948630 | low | 1238.4292 | OK | 0.017269 | no | 91.8305 | no | 0.96082 | 0.85 | no | 0.97249 | no | 0.94905 | no | 0.84095 | no | 0.0018358 | 5.1652e-05 | 0 | -0.00036272 | no | 0 | -0.00037743 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150521_M02641_0051_000000000-AF9JH | M02641 | 2015-05-21 | 75 | 75 | 85.4273 | no | 20774323 | low | 847.1236 | low | 0.093549 | yes | 85.9442 | no | 0.92414 | 0.85 | no | 0.93171 | no | 0.91686 | no | 0.99584 | no | 0.99485 | 0.61684 | 0 | -0.00098979 | yes | 0 | -0.00048486 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff
150522_M00766_0103_000000000-AF9M3 | M00766 | 2015-05-22 | 75 | 75 | 88.2546 | no | 33587117 | low | 1405.6196 | high | 0.026305 | no | 88.4186 | no | 0.94466 | 0.85 | no | 0.95853 | no | 0.92962 | no | 0.84095 | no | 0.0015346 | nan | 0 | -0.00056188 | yes | 0 | -0.00050146 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150522_M02641_0052_000000000-AF82F | M02641 | 2015-05-22 | 75 | 75 | 77.7302 | no | 26557402 | low | 1103.8322 | high | 0.045966 | no | 76.9724 | yes | 0.90712 | 0.85 | no | 0.92279 | no | 0.89107 | no | 0.99584 | no | 0.060099 | 5.5163e-09 | 0 | -0.00079849 | yes | 0 | -0.00064071 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150601_M00766_0107_000000000-AF9N8 | M00766 | 2015-06-01 | 75 | 75 | 96.709 | no | 17833767 | low | 714.0694 | low | 0.036487 | no | 97.1935 | no | 0.97916 | 0.85 | no | 0.98619 | no | 0.97361 | no | 0.84095 | no | 0.47307 | nan | 0 | -0.00024948 | no | 0 | -0.00016826 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150602_M00766_0108_000000000-AFMB8 | M00766 | 2015-06-02 | 200 | 200 | 87.6065 | no | 33648503 | low | 1419.0921 | high | 0.017208 | no | 87.2266 | no | 0.8224 | 0.7 | no | 0.84731 | no | 0.7928 | no | 0.84095 | no | 3.1434e-29 | 0 | 19 | -0.0019091 | yes | 33 | -0.0023749 | yes | pass | pass | pass | pass | same | pass | not exceeded | exceeded | diff | same
150603_M02641_0056_000000000-AEY8R | M02641 | 2015-06-03 | 151 | 151 | 75.779 | no | 10182405 | low | 516.5408 | low | 0.052494 | no | 76.2721 | yes | 0.92128 | 0.8 | no | 0.93202 | no | 0.91084 | no | 0.99584 | no | 0.72835 | nan | 0 | -0.00018443 | no | 0 | -0.00018378 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150604_M00766_0109_000000000-AFJ62 | M00766 | 2015-06-04 | 75 | 75 | 90.2876 | no | 30654786 | low | 1283.025 | high | 0.019566 | no | 90.5316 | no | 0.94854 | 0.85 | no | 0.96577 | no | 0.93018 | no | 0.84095 | no | 6.5656e-06 | nan | 0 | -0.0004699 | no | 0 | -0.00054877 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150604_M02641_0057_000000000-AFL7N | M02641 | 2015-06-04 | 75 | 75 | 83.6851 | no | 20989024 | low | 846.0443 | low | 0.089565 | yes | 87.1228 | no | 0.92879 | 0.85 | no | 0.94412 | no | 0.91514 | no | 0.84095 | no | 0.077245 | nan | 0 | -0.00064783 | yes | 0 | -0.00048886 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff
150605_M00766_0110_000000000-AF80F | M00766 | 2015-06-05 | 75 | 75 | 83.8601 | no | 32100609 | low | 1379.9375 | high | 0.019606 | no | 84.3276 | no | 0.92414 | 0.85 | no | 0.94606 | no | 0.89846 | no | 0.84095 | no | 2.0225e-07 | nan | 0 | -0.00071066 | yes | 0 | -0.00077961 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150610_M00766_0111_000000000-AFMWF | M00766 | 2015-06-10 | 75 | 75 | 89.8899 | no | 32008549 | low | 1328.1731 | high | 0.015813 | no | 89.8802 | no | 0.94912 | 0.85 | no | 0.96373 | no | 0.93447 | no | 0.84095 | no | 0.0020072 | nan | 0 | -0.00052315 | yes | 0 | -0.00050342 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150611_M02641_0060_000000000-AFN2Y | M02641 | 2015-06-11 | 75 | 75 | 76.6898 | no | 23799396 | low | 947.2643 | OK | 0.054797 | no | 75.8247 | yes | 0.88682 | 0.85 | no | 0.90551 | no | 0.86738 | no | 0.99584 | no | 0.046537 | 2.6959e-66 | 0 | -0.001158 | yes | 0 | -0.00077535 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150615_M02641_0062_000000000-AFH7C | M02641 | 2015-06-15 | 75 | 75 | 76.7102 | no | 21607117 | low | 847.9618 | low | 0.051033 | no | 78.099 | yes | 0.89692 | 0.85 | no | 0.91029 | no | 0.88432 | no | 0.99584 | no | 0.37812 | 0.020859 | 0 | -0.00092944 | yes | 0 | -0.00061072 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150619_M02641_0063_000000000-AFN2W | M02641 | 2015-06-19 | 75 | 75 | 86.313 | no | 22890237 | low | 921.1523 | low | 0.031622 | no | 86.5093 | no | 0.93054 | 0.85 | no | 0.94446 | no | 0.91735 | no | 0.95413 | no | 0.013174 | nan | 0 | -0.00066355 | yes | 0 | -0.0006793 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150625_M02641_0065_000000000-AEY2A | M02641 | 2015-06-25 | 151 | 151 | 67.3086 | yes | 7123794 | low | 354.4358 | low | 0.027403 | no | 65.5091 | yes | 0.90035 | 0.8 | no | 0.9059 | no | 0.89467 | no | 0.99998 | no | 0.37708 | 0.0069801 | 0 | -0.00018719 | no | 0 | -0.00016475 | no | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
150703_M00766_0117_000000000-AFN2F | M00766 | 2015-07-03 | 75 | 75 | 89.9328 | no | 28300840 | low | 1173.6259 | OK | 0.036091 | no | 90.4158 | no | 0.94919 | 0.85 | no | 0.95814 | no | 0.93984 | no | 0.84095 | no | 0.81385 | nan | 0 | -0.00062428 | yes | 0 | -0.00041123 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150703_M02641_0001_000000000-AFMW2 | M02641 | 2015-07-03 | 75 | 75 | 92.3482 | no | 29421610 | low | 1208.9627 | OK | 0.019251 | no | 92.6367 | no | 0.95661 | 0.85 | no | 0.96958 | no | 0.94384 | no | 0.84095 | no | 0.001929 | 2.1798e-15 | 0 | -0.0004435 | no | 0 | -0.00035874 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150709_M02641_0004_000000000-AFMPC | M02641 | 2015-07-09 | 75 | 75 | 94.2018 | no | 27443669 | low | 1115.769 | low | 0.036555 | no | 95.4991 | no | 0.96502 | 0.85 | no | 0.9763 | no | 0.95448 | no | 0.84095 | no | 0.016984 | nan | 0 | -0.00039431 | no | 0 | -0.00024497 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150727_M00766_0119_000000000-AF3TM | M00766 | 2015-07-27 | 151 | 151 | 88.8616 | no | 15391937 | low | 825.3634 | low | 0.055595 | no | 88.285 | no | 0.9196 | 0.8 | no | 0.9501 | no | 0.88668 | no | 0.84095 | no | 1.2664e-16 | nan | 0 | -0.00038471 | no | 0 | -0.00086889 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150727_M02641_0007_000000000-AGEW7 | M02641 | 2015-07-27 | 230 | 230 | 91.5639 | no | 31830282 | low | 1309.0338 | high | 0.015312 | no | 91.8906 | no | 0.76141 | 0.7 | no | 0.80125 | no | 0.71555 | no | 0.35842 | no | 5.1559e-54 | 2.0554e-121 | 29 | -0.0026602 | yes | 56 | -0.0035322 | yes | pass | pass | pass | pass | same | pass | exceeded | exceeded | same | same
150807_M00766_0122_000000000-AFMW7 | M00766 | 2015-08-07 | 75 | 75 | 95.8884 | no | 19129362 | low | 771.607 | low | 0.031184 | no | 96.3208 | no | 0.9746 | 0.85 | no | 0.98266 | no | 0.96788 | no | 0.84095 | no | 0.16299 | nan | 0 | -0.00031347 | no | 0 | -0.0002468 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150813_M00766_0124_000000000-AFYEH | M00766 | 2015-08-13 | 75 | 75 | 94.3227 | no | 25145640 | low | 1025.335 | low | 0.021357 | no | 94.2549 | no | 0.96769 | 0.85 | no | 0.97825 | no | 0.95812 | no | 0.84095 | no | 0.0084968 | nan | 0 | -0.00033121 | no | 0 | -0.00034652 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150818_M00766_0125_000000000-AE8BP | M00766 | 2015-08-18 | 151 | 151 | 91.3886 | no | 15351123 | low | 804.6945 | low | 0.058805 | no | 90.8408 | no | 0.92459 | 0.8 | no | 0.94957 | no | 0.89589 | no | 0.84095 | no | 6.1155e-16 | 3.3622e-25 | 0 | -0.00052265 | yes | 0 | -0.0011087 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150826_M00766_0127_000000000-AGKLT | M00766 | 2015-08-26 | 200 | 200 | 93.1048 | no | 24799829 | low | 1017.4002 | low | 0.016826 | no | 93.586 | no | 0.88995 | 0.7 | no | 0.91037 | no | 0.86811 | no | 0.67794 | no | 4.0958e-20 | nan | 0 | -0.0013198 | yes | 0 | -0.0017885 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150904_M00766_0129_000000000-AFNA2 | M00766 | 2015-09-04 | 151 | 151 | 84.0182 | no | 18908531 | low | 1029.9385 | high | 0.024332 | no | 83.7698 | no | 0.81255 | 0.8 | no | 0.86788 | no | 0.75533 | yes | 0.95413 | no | 5.0649e-39 | nan | 0 | -0.0022068 | yes | 9 | -0.0021487 | yes | pass | fail | pass | fail | diff | pass | not exceeded | not exceeded | same | same
150910_M00766_0130_000000000-AGJGG | M00766 | 2015-09-10 | 75 | 75 | 88.355 | no | 24463097 | low | 1059.3697 | low | 0.0095777 | no | 88.2815 | no | 0.94136 | 0.85 | no | 0.95905 | no | 0.92225 | no | 0.95413 | no | 1.5264e-06 | 1.3678e-23 | 0 | -0.00056717 | yes | 0 | -0.00075898 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150916_M02641_0025_000000000-AFN33 | M02641 | 2015-09-16 | 151 | 151 | 91.2782 | no | 21293722 | low | 1112.3198 | high | 0.023575 | no | 90.9013 | no | 0.8777 | 0.8 | no | 0.90422 | no | 0.84924 | no | 0.95413 | no | 3.9616e-16 | nan | 0 | -0.0018497 | yes | 5 | -0.0021883 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150917_M02641_0026_000000000-AFN0A | M02641 | 2015-09-17 | 151 | 151 | 91.5269 | no | 20122988 | low | 1048.5468 | OK | 0.02524 | no | 91.3116 | no | 0.86192 | 0.8 | no | 0.89213 | no | 0.82928 | no | 0.84095 | no | 2.1049e-19 | nan | 1 | -0.0022155 | yes | 10 | -0.0026282 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150918_M00766_0133_000000000-AH985 | M00766 | 2015-09-18 | 200 | 200 | 91.3077 | no | 30062610 | low | 1253.19 | OK | 0.031977 | no | 91.5655 | no | 0.84979 | 0.7 | no | 0.88478 | no | 0.81133 | no | 0.67794 | no | 4.0326e-70 | nan | 11 | -0.0017023 | yes | 24 | -0.0021771 | yes | pass | pass | pass | pass | same | pass | not exceeded | exceeded | diff | same
150923_M00766_0134_000000000-AFN32 | M00766 | 2015-09-23 | 151 | 151 | 89.0495 | no | 20113067 | low | 1072.9998 | high | 0.033799 | no | 88.906 | no | 0.86242 | 0.8 | no | 0.89294 | no | 0.82952 | no | 0.95413 | no | 1.619e-18 | nan | 0 | -0.002005 | yes | 9 | -0.0025255 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150924_M02641_0027_000000000-AGJGM | M02641 | 2015-09-24 | 75 | 75 | 95.0401 | no | 24554073 | low | 997.0288 | low | 0.0085367 | no | 95.0918 | no | 0.97292 | 0.85 | no | 0.98196 | no | 0.96511 | no | 0.84095 | no | 0.019748 | nan | 0 | -0.00019383 | no | 0 | -0.00024233 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150925_M00766_0135_000000000-AFL75 | M00766 | 2015-09-25 | 151 | 151 | 97.708 | no | 9080878 | low | 460.5892 | low | 0.039145 | no | 97.9623 | no | 0.94702 | 0.8 | no | 0.95724 | no | 0.93649 | no | 0.95413 | no | 0.0051842 | nan | 0 | -0.00069048 | yes | 0 | -0.0008509 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
150929_M02641_0030_000000000-AFWJ3 | M02641 | 2015-09-29 | 151 | 151 | 95.2167 | no | 14824546 | low | 765.6339 | low | 0.037672 | no | 95.287 | no | 0.94667 | 0.8 | no | 0.96411 | no | 0.92657 | no | 0.84095 | no | 2.2694e-08 | nan | 0 | -0.00044841 | no | 0 | -0.00094835 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
150930_M00766_0138_000000000-AH4BR | M00766 | 2015-09-30 | 75 | 75 | 90.8515 | no | 29081266 | low | 1203.4983 | OK | 0.018375 | no | 90.911 | no | 0.93455 | 0.85 | no | 0.91651 | no | 0.95034 | no | 0.99584 | no | 5.9019e-07 | nan | 0 | -0.0013505 | yes | 0 | -0.00035323 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151005_M02641_0034_000000000-AH9ML | M02641 | 2015-10-05 | 200 | 200 | 91.5287 | no | 33902064 | low | 1389.8402 | high | 0.014042 | no | 91.6743 | no | 0.83382 | 0.7 | no | 0.86254 | no | 0.80114 | no | 0.50766 | no | 1.3625e-44 | nan | 23 | -0.0021037 | yes | 30 | -0.0027427 | yes | pass | pass | pass | pass | same | pass | exceeded | exceeded | same | same
151008_M02641_0035_000000000-AFL73 | M02641 | 2015-10-08 | 151 | 151 | 92.9471 | no | 20624078 | low | 1070.1471 | OK | 0.02054 | no | 93.2815 | no | 0.88534 | 0.8 | no | 0.91095 | no | 0.85745 | no | 0.95413 | no | 1.7378e-17 | nan | 0 | -0.0018668 | yes | 5 | -0.0023798 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151013_M02641_0038_000000000-AFLW0 | M02641 | 2015-10-13 | 151 | 151 | 92.2835 | no | 19868115 | low | 1036.6449 | OK | 0.015868 | no | 92.3144 | no | 0.85852 | 0.8 | no | 0.89862 | no | 0.81584 | no | 0.84095 | no | 9.4078e-32 | 1.0701e-96 | 0 | -0.0019454 | yes | 7 | -0.0029862 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151014_M02641_0039_000000000-AFLFV | M02641 | 2015-10-14 | 151 | 151 | 93.1322 | no | 20043314 | low | 1036.9175 | OK | 0.024404 | no | 93.1321 | no | 0.86964 | 0.8 | no | 0.90574 | no | 0.83081 | no | 0.84095 | no | 1.4818e-31 | nan | 0 | -0.0019202 | yes | 8 | -0.0027745 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151015_M00766_0142_000000000-AJD8B | M00766 | 2015-10-15 | 151 | 151 | 88.82 | no | 16091719 | low | 856.7867 | low | 0.030984 | no | 88.9965 | no | 0.9206 | 0.8 | no | 0.95153 | no | 0.88855 | no | 0.67794 | no | 9.7576e-25 | nan | 0 | -0.0004536 | no | 1 | -0.0013313 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151021_M02641_0041_000000000-AGHAA | M02641 | 2015-10-21 | 75 | 75 | 93.3815 | no | 29221047 | low | 1189.4531 | OK | 0.018033 | no | 93.5119 | no | 0.96054 | 0.85 | no | 0.97325 | no | 0.9489 | no | 0.84095 | no | 0.001648 | 0.00019262 | 0 | -0.00043517 | no | 0 | -0.00032505 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151022_M00766_0146_000000000-AFN3B | M00766 | 2015-10-22 | 151 | 151 | 81.4565 | no | 19862739 | low | 1107.5448 | very high | 0.031521 | no | 80.9673 | no | 0.90849 | 0.8 | no | 0.93308 | no | 0.88275 | no | 0.95413 | no | 6.0372e-16 | 1.7055e-184 | 0 | -0.00050949 | yes | 0 | -0.00097367 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
151023_M00766_0147_000000000-AFN37 | M00766 | 2015-10-23 | 151 | 151 | 84.6509 | no | 20693511 | low | 1134.0694 | very high | 0.039199 | no | 85.3046 | no | 0.92707 | 0.8 | no | 0.95123 | no | 0.90247 | no | 0.95413 | no | 6.8289e-18 | nan | 0 | -0.00030046 | no | 0 | -0.00056736 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151028_M02641_0043_000000000-AFL74 | M02641 | 2015-10-28 | 151 | 151 | 96.8353 | no | 15543815 | low | 805.7896 | low | 0.035546 | no | 96.7917 | no | 0.96903 | 0.8 | no | 0.97972 | no | 0.95847 | no | 0.67794 | no | 8.1444e-08 | 1.4081e-14 | 0 | -0.00020911 | no | 0 | -0.0004479 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151030_M00766_0150_000000000-AFLDT | M00766 | 2015-10-30 | 151 | 151 | 89.1258 | no | 19498823 | low | 1043.7728 | OK | 0.022316 | no | 89.1735 | no | 0.94282 | 0.8 | no | 0.95917 | no | 0.92637 | no | 0.95413 | no | 6.2812e-11 | 1.3707e-27 | 0 | -0.00027197 | no | 0 | -0.00043246 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151030_M02641_0044_000000000-AJHLJ | M02641 | 2015-10-30 | 75 | 75 | 89.5684 | no | 33168068 | low | 1381.3797 | high | 0.015021 | no | 89.5367 | no | 0.94181 | 0.85 | no | 0.96027 | no | 0.92327 | no | 0.84095 | no | 1.9212e-09 | 1.7277e-37 | 0 | -0.00053639 | yes | 0 | -0.00046394 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151103_M02641_0046_000000000-AJ5Y7 | M02641 | 2015-11-03 | 151 | 151 | 96.1111 | no | 17192399 | low | 883.2398 | low | 0.021428 | no | 96.0917 | no | 0.96652 | 0.8 | no | 0.97476 | no | 0.95884 | no | 0.95413 | no | 0.00063959 | nan | 0 | -0.00027215 | no | 0 | -0.00034016 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151105_M02641_0047_000000000-AJF4E | M02641 | 2015-11-05 | 151 | 151 | 96.0812 | no | 17316625 | low | 888.5653 | low | 0.024517 | no | 96.1753 | no | 0.96613 | 0.8 | no | 0.97668 | no | 0.95675 | no | 0.84095 | no | 7.3533e-09 | 2.5766e-25 | 0 | -0.00020337 | no | 0 | -0.00027243 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151116_M02641_0050_000000000-AJ6L3 | M02641 | 2015-11-16 | 151 | 151 | 95.1303 | no | 13144452 | low | 672.6944 | low | 0.074365 | no | 94.8898 | no | 0.95687 | 0.8 | no | 0.96857 | no | 0.94432 | no | 0.95413 | no | 0.0016197 | 1.4817e-06 | 0 | -0.00029535 | no | 0 | -0.00053054 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151117_M00766_0155_000000000-AJ6LL | M00766 | 2015-11-17 | 151 | 151 | 87.0005 | no | 16509092 | low | 884.2256 | OK | 0.024563 | no | 86.7919 | no | 0.86377 | 0.8 | no | 0.9042 | no | 0.821 | no | 0.84095 | no | 8.628e-16 | nan | 0 | -0.0012235 | yes | 1 | -0.00063026 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151117_M02641_0051_000000000-AJF4A | M02641 | 2015-11-17 | 151 | 151 | 94.2241 | no | 17785498 | low | 905.3067 | low | 0.03067 | no | 94.2545 | no | 0.88313 | 0.8 | no | 0.91617 | no | 0.84833 | no | 0.84095 | no | 7.9592e-12 | nan | 0 | -0.0012532 | yes | 0 | -0.00072138 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151125_M02641_0053_000000000-AJJ3G | M02641 | 2015-11-25 | 151 | 151 | 96.6504 | no | 3965190 | low | 201.2481 | very low | 0.02812 | no | 96.9381 | no | 0.97423 | 0.8 | no | 0.98297 | no | 0.96916 | no | 0.67794 | no | 0.069974 | 0.0030109 | 0 | -0.00019201 | no | 0 | -0.00032396 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
151126_M00766_0159_000000000-AJD7L | M00766 | 2015-11-26 | 75 | 75 | 88.0477 | no | 24133209 | low | 1032.2091 | low | 0.017289 | no | 87.5838 | no | 0.95346 | 0.85 | no | 0.96296 | no | 0.94323 | no | 0.84095 | no | 0.093829 | 1.6996e-15 | 0 | -0.0004886 | no | 0 | -0.00047968 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151126_M02641_0054_000000000-AJD7T | M02641 | 2015-11-26 | 75 | 75 | 94.0217 | no | 24949166 | low | 1025.4829 | low | 0.029656 | no | 93.8083 | no | 0.96822 | 0.85 | no | 0.97532 | no | 0.96185 | no | 0.95413 | no | 0.1922 | nan | 0 | -0.00030117 | no | 0 | -0.00033189 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151127_M00766_0161_000000000-AJERP | M00766 | 2015-11-27 | 75 | 75 | 90.539 | no | 24073552 | low | 1002.8265 | low | 0.022959 | no | 90.0889 | no | 0.95458 | 0.85 | no | 0.96258 | no | 0.94625 | no | 0.95413 | no | 0.1243 | nan | 0 | -0.00057536 | yes | 0 | -0.00051622 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151201_M02641_0055_000000000-AJ6B2 | M02641 | 2015-12-01 | 151 | 151 | 93.147 | no | 20354768 | low | 1062.1404 | OK | 0.019929 | no | 93.6277 | no | 0.95729 | 0.8 | no | 0.96997 | no | 0.94485 | no | 0.95413 | no | 2.0042e-09 | nan | 0 | -0.00022088 | no | 0 | -0.0003646 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151203_M00766_0165_000000000-AJFAA | M00766 | 2015-12-03 | 151 | 151 | 86.1622 | no | 16734072 | low | 890.0657 | OK | 0.048269 | no | 84.2361 | no | 0.92678 | 0.8 | no | 0.94416 | no | 0.90681 | no | 0.95413 | no | 1.2134e-06 | nan | 0 | -0.0003637 | no | 0 | -0.00065096 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151208_M00766_0167_000000000-AJ5JR | M00766 | 2015-12-08 | 151 | 151 | 97.4409 | no | 5807543 | low | 293.2064 | very low | 0.036817 | no | 97.4486 | no | 0.96505 | 0.8 | no | 0.973 | no | 0.95786 | no | 0.95413 | no | 0.12937 | nan | 0 | -0.00036542 | no | 0 | -0.00043496 | no | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
151210_M02641_0058_000000000-AJ5YJ | M02641 | 2015-12-10 | 151 | 151 | 91.7526 | no | 18225198 | low | 951.6353 | OK | 0.049608 | no | 91.4934 | no | 0.93902 | 0.8 | no | 0.95565 | no | 0.92066 | no | 0.95413 | no | 6.5736e-05 | nan | 0 | -0.00040902 | no | 0 | -0.00072147 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151216_M00766_0170_000000000-ALEAF | M00766 | 2015-12-16 | 151 | 151 | 93.7279 | no | 17003866 | low | 880.7103 | low | 0.036745 | no | 93.9642 | no | 0.94991 | 0.8 | no | 0.96257 | no | 0.937 | no | 0.95413 | no | 4.5633e-06 | nan | 0 | -0.00035203 | no | 0 | -0.00057532 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
151217_M02641_0059_000000000-AJFAD | M02641 | 2015-12-17 | 151 | 151 | 58.3542 | yes | 4092271 | low | 204.0891 | very low | 0.28007 | yes | 55.6594 | yes | 0.92478 | 0.8 | no | 0.94271 | no | 0.91061 | no | 0.84095 | no | 0.016952 | 0.00029975 | 0 | -0.00041346 | no | 0 | -0.00076263 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | diff
151223_M00766_0172_000000000-AJRMN | M00766 | 2015-12-23 | 75 | 75 | 92.6132 | no | 23635202 | low | 977.0359 | low | 0.01686 | no | 92.2318 | no | 0.94189 | 0.85 | no | 0.9657 | no | 0.91522 | no | 0.50766 | no | 8.2817e-14 | nan | 0 | -0.00070693 | yes | 1 | -0.0018884 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151223_M02641_0060_000000000-AGHV2 | M02641 | 2015-12-23 | 75 | 75 | 87.711 | no | 26108005 | low | 1077.9707 | OK | 0.015686 | no | 87.8311 | no | 0.92979 | 0.85 | no | 0.94545 | no | 0.911 | no | 0.95413 | no | 1.0014e-06 | 6.7599e-86 | 0 | -0.00093076 | yes | 0 | -0.0011369 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
151231_M02641_0061_000000000-AJD8L | M02641 | 2015-12-31 | 75 | 75 | 88.6234 | no | 30986761 | low | 1292.1541 | high | 0.043562 | no | 89.278 | no | 0.93779 | 0.85 | no | 0.95903 | no | 0.91501 | no | 0.95413 | no | 2.22e-08 | nan | 0 | -0.00055157 | yes | 0 | -0.00062885 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160107_M00766_0175_000000000-AH6C3 | M00766 | 2016-01-07 | 75 | 75 | 87.9117 | no | 26044869 | low | 1123.4126 | OK | 0.012327 | no | 86.8229 | no | 0.95101 | 0.85 | no | 0.96122 | no | 0.9395 | no | 0.95413 | no | 0.056283 | nan | 0 | -0.00056289 | yes | 0 | -0.00070024 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160108_M00766_0176_000000000-AJDAJ | M00766 | 2016-01-08 | 75 | 75 | 86.9378 | no | 24436074 | low | 1042.2918 | low | 0.023113 | no | 87.1996 | no | 0.79786 | 0.85 | yes | 0.93046 | no | 0.6482 | yes | 0.15463 | no | 1.2164e-131 | nan | 0 | -0.0018269 | yes | 36 | -0.0074037 | yes | fail | fail | fail | fail | diff | pass | not exceeded | exceeded | diff | same
160114_M00766_0177_000000000-ALYTH | M00766 | 2016-01-14 | 75 | 75 | 87.5409 | no | 27732794 | low | 1180.5507 | OK | 0.016589 | no | 87.6219 | no | 0.94367 | 0.85 | no | 0.95391 | no | 0.93177 | no | 0.95413 | no | 0.043231 | nan | 0 | -0.00066231 | yes | 0 | -0.00077358 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160114_M02641_0065_000000000-ALRU0 | M02641 | 2016-01-14 | 75 | 75 | 91.58 | no | 29247440 | low | 1214.9636 | OK | 0.024503 | no | 91.5468 | no | 0.95131 | 0.85 | no | 0.96723 | no | 0.93592 | no | 0.84095 | no | 7.4878e-05 | nan | 0 | -0.00049951 | no | 0 | -0.00034584 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160128_M02641_0070_000000000-ALAGE | M02641 | 2016-01-28 | 75 | 75 | 91.826 | no | 24305600 | low | 1013.4527 | low | 0.0159 | no | 91.7655 | no | 0.96174 | 0.85 | no | 0.96994 | no | 0.95369 | no | 0.95413 | no | 0.073098 | nan | 0 | -0.0003153 | no | 0 | -0.00042581 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160129_M00766_0003_000000000-ALAH0 | M00766 | 2016-01-29 | 75 | 75 | 86.7633 | no | 32429499 | low | 1380.9019 | high | 0.017673 | no | 87.2761 | no | 0.93918 | 0.85 | no | 0.95507 | no | 0.92174 | no | 0.84095 | no | 1.4395e-05 | nan | 0 | -0.00058412 | yes | 0 | -0.00060062 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160129_M02641_0071_000000000-AMC8N | M02641 | 2016-01-29 | 200 | 200 | 91.7102 | no | 25390554 | low | 1054.7902 | low | 0.01824 | no | 91.6301 | no | 0.86183 | 0.7 | no | 0.88102 | no | 0.83939 | no | 0.67794 | no | 1.3683e-14 | 6.476e-15 | 1 | -0.0017129 | yes | 0 | -0.0020671 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160206_M00766_0005_000000000-ALYU1 | M00766 | 2016-02-06 | 75 | 75 | 84.6656 | no | 32152720 | low | 1381.1846 | high | 0.017535 | no | 83.6833 | no | 0.93615 | 0.85 | no | 0.9506 | no | 0.92023 | no | 0.84095 | no | 0.0023058 | 1.3476e-13 | 0 | -0.00053805 | yes | 0 | -0.00042151 | no | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
160206_M02641_0074_000000000-AL604 | M02641 | 2016-02-06 | 151 | 151 | 88.8163 | no | 18930346 | low | 1014.3724 | OK | 0.015119 | no | 88.4642 | no | 0.92173 | 0.8 | no | 0.9483 | no | 0.89269 | no | 0.67794 | no | 2.5587e-23 | 1.4688e-132 | 0 | -0.00040622 | no | 0 | -0.00095862 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
160209_M00766_0007_000000000-AMERJ | M00766 | 2016-02-09 | 75 | 75 | 79.4588 | no | 34759702 | low | 1528.9595 | very high | 0.016677 | no | 80.1993 | no | 0.90703 | 0.85 | no | 0.93282 | no | 0.87982 | no | 0.84095 | no | 1.1867e-08 | 0 | 0 | -0.00081183 | yes | 0 | -0.00073619 | yes | fail | fail | pass | pass | same | pass | not exceeded | not exceeded | same | same
160216_M00766_0009_000000000-AL73W | M00766 | 2016-02-16 | 151 | 151 | 92.117 | no | 18300588 | low | 961.5089 | OK | 0.025014 | no | 92.5591 | no | 0.94116 | 0.8 | no | 0.96156 | no | 0.91989 | no | 0.84095 | no | 1.5361e-15 | 3.496e-24 | 0 | -0.00029836 | no | 0 | -0.00066338 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
160219_M00766_0011_000000000-AL6GM | M00766 | 2016-02-19 | 131 | 131 | 5.6098 | yes | 661574 | low | 22.0869 | low | 0.9623 | yes | 9.7324 | yes | 0.77871 | 0.8 | yes | 0.79992 | yes | 0.78598 | yes | 0.99584 | no | 1 | 0.73572 | 5 | -0.00056401 | yes | 5 | -0.00059171 | yes | fail | fail | fail | fail | same | pass | not exceeded | not exceeded | same | same
160220_M02641_0081_000000000-AMETF | M02641 | 2016-02-20 | 200 | 200 | 92.9979 | no | 23074941 | low | 958.5024 | low | 0.027553 | no | 93.6783 | no | 0.87882 | 0.7 | no | 0.90362 | no | 0.85132 | no | 0.35842 | no | 6.4948e-26 | nan | 0 | -0.0012552 | yes | 0 | -0.0017992 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | same
160223_M00766_0012_000000000-AL5YC | M00766 | 2016-02-23 | 151 | 151 | 93.2958 | no | 18111084 | low | 951.0894 | OK | 0.019086 | no | 93.3336 | no | 0.94516 | 0.8 | no | 0.96589 | no | 0.92393 | no | 0.84095 | no | 9.6635e-16 | 3.8603e-52 | 0 | -0.0002968 | no | 0 | -0.0005917 | yes | pass | pass | pass | pass | same | pass | not exceeded | not exceeded | same | diff
| 76 |  |  | NUM FAILS |  | 7 |  | 123 |  | 17 |  | 11 |  | 13 |  |  | 10 |  | 4 |  | 15 |  | 0 |  |  | 8 |  | 77 | 39 |  | 85 | 32 | 37 | 23 | 28 | 11 | 0 |  |  | 11 | 24
|  |  |  | % FAILS |  | 5.6911 |  | 100 |  | 13.8211 |  | 8.9431 |  | 10.5691 |  |  | 8.1301 |  | 3.252 |  | 12.1951 |  | 0 |  |  | 6.5041 |  | 62.6016 | 31.7073 |  | 69.1057 | 91 | 86 | 100 | 95
Exclude this one as useless
[32 empty] | Method 1 | Method 2 | Method 1 (no CD) | Method 2 (no CD)
[31 empty] | Overall | 90 | 85 | 99 | 94
[31 empty] | Pass | 32 | 37 | 23 | 28
Fail

## BadRunsEye
RunID | InstrumentID | RunDate | Pass/Fail | Indexing Comments | General Comments
121009_M00766_0002_000000000-A1U6P | M00766 | 2012-10-09 | Fail | No info | Large difference in cluster density passing filter and cluster density; Smearing in heatmap
130206_M00766_0002_000000000-A23JM | M00766 | 2013-02-06 | Fail |  | Cluster density low; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
130618_M00766_0019_000000000-A4FEU | M00766 | 2013-06-18 | Fail |  | Cluster density low; Low >Q30; Smearing in heatmap
130624_M00766_0021_000000000-A53PT | M00766 | 2013-06-24 | Pass |  | Some cycles smearing in heatmap
130712_M00766_0024_000000000-A5AVV | M00766 | 2013-07-12 | Borderline Pass | One sample a little low | Some cycles smearing in heatmap; Reasonable difference in cluster density passing filter and cluster density
130817_M00766_0037_000000000-A4EDD | M00766 | 2013-08-17 | Pass | Some samples low; Some samples high
130913_M00766_0049_000000000-A5B0E | M00766 | 2013-09-13 | Pass |  | Some cycles smearing in heatmap
130916_M00766_0050_000000000-A5M1U | M00766 | 2013-09-16 | Pass | One sample low
130924_M00766_0052_000000000-A5BE6 | M00766 | 2013-09-24 | Fail | Two samples big range | Cluster density a bit high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
131007_M00766_0055_000000000-A59JT | M00766 | 2013-10-07 | Borderline Pass | Two samples big range | Cluster density a bit low; Reasonable difference in cluster density passing filter and cluster density; Some smearing in heatmap
131025_M00766_0059_000000000-A5P9A | M00766 | 2013-10-25 | Pass | Some samples very low; Some samples very high
140207_M00766_0022_000000000-A7PH8 | M00766 | 2014-02-07 | Pass | Two samples big range | Some smearing in heatmap towards end of reads (small) amount
140424_M00766_0033_000000000-A7BNA | M00766 | 2014-04-24 | Borderline Fail | One sample high | Very low cluster density
140514_M00766_0036_000000000-A7BRK | M00766 | 2014-05-14 | Fail | Two samples high; Two samples low | Cluster density high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
140612_M00766_0039_000000000-A7BNL | M00766 | 2014-06-12 | Fail |  | Low >Q30; Smearing in heatmap
140616_M00766_0040_000000000-A78V9 | M00766 | 2014-06-16 | Borderline Pass |  | Cluster density a little low; Some smearing in heatmap
140813_M00766_0047_000000000-A8PJL | M00766 | 2014-08-13 | Pass | One sample a little high
140901_M00766_0048_000000000-AA63M | M00766 | 2014-09-01 | Pass | One sample a little high; One sample a little low
141014_M00766_0053_000000000-A8R6M | M00766 | 2014-10-14 | Pass |  | Some small amount of smearing towards the ends of reads
141017_M02641_0022_000000000-AA66H | M02641 | 2014-10-17 | Fail |  | Cluster density high; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
141024_M00766_0056_000000000-A8PC5 | M00766 | 2014-10-24 | Borderline Fail |  | Cluster density high; Reasonable difference in cluster density passing filter and cluster density; Some smearing in heatmap
141031_M00766_0058_000000000-AA3GN | M00766 | 2014-10-31 | Pass |  | Cluster density could be a bit low
141107_M00766_0060_000000000-AA8PM | M00766 | 2014-11-07 | Borderline Pass | Some samples high; Some samples low | Cluster density could be a bit low
141118_M00766_0061_000000000-A8P8J | M00766 | 2014-11-18 | Pass
141208_M00766_0063_000000000-A8P7C | M00766 | 2014-12-08 | Borderline Pass | Two samples high | Cluster density a little high; Some smearing in heatmap towards end of reads (small) amount
141208_M02641_0026_000000000-A8R55 | M02641 | 2014-12-08 | Borderline Pass | One sample a little high; One sample a little low | Cluster density a little low; Smearing in heatmap particularly towards the end of read 2
141212_M00766_0064_000000000-ACCEB | M00766 | 2014-12-12 | Fail | Some samples high; Some samples low; NTC high | Cluster density high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
141216_M00766_0065_000000000-ACCDT | M00766 | 2014-12-16 | Fail | One sample low; One sample very high | Cluster density high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
141219_M00766_0066_000000000-ACCB1 | M00766 | 2014-12-19 | Fail | Some samples a little low; One sample very high | Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
150120_M00766_0071_000000000-AA63K | M00766 | 2015-01-20 | Pass | One sample very high | Some small amount of smearing towards the ends of reads
150127_M02641_0027_000000000-AA65J | M02641 | 2015-01-27 | Pass | One sample low
150203_M00766_0074_000000000-AAUMH | M00766 | 2015-02-03 | Pass | One sample very low; Some samples high | Reasonable difference in cluster density passing filter and cluster density
150203_M02641_0028_000000000-ACBYG | M02641 | 2015-02-03 | Borderline Pass | One sample very low; One sample very high | Reasonable difference in cluster density passing filter and cluster density; Some smearing in heatmap towards the ends of reads
150205_M00766_0075_000000000-ACC43 | M00766 | 2015-02-05 | Pass | Some samples very low; One sample high | Some smearing in heatmap towards the ends of reads
150224_M00766_0078_000000000-ACMP5 | M00766 | 2015-02-24 | Fail |  | High cluster density; Large difference in cluster density passing filter and cluster density
150309_M02641_0035_000000000-ACC3J | M02641 | 2015-03-09 | Fail | One sample low | Cluster density a little high; Low >Q30; Reasonable difference between cluster density passing filter and cluster density; Smearing in heatmap
150311_M00766_0081_000000000-ACN0W | M00766 | 2015-03-11 | Pass | Some samples low; Some samples high
150311_M02641_0036_000000000-AD8VB | M02641 | 2015-03-11 | Pass | One sample high | Cluster density a little low
150402_M00766_0087_000000000-ACCAY | M00766 | 2015-04-02 | Borderline Pass | Some samples high | Some smearing in heatmap towards end of reads
150414_M00766_0090_000000000-AAU39 | M00766 | 2015-04-14 | Pass | One sample very low
150427_M02641_0042_000000000-AD8LB | M02641 | 2015-04-27 | Borderline Fail | One sample very low; One sample very high | Very low cluster density
150430_M00766_0092_000000000-AD77P | M00766 | 2015-04-30 | Borderline Pass | Some samples high | Cluster density a little high; Reasonable difference between cluster density passing filter and cluster density
150506_M00766_0094_000000000-AEVP8 | M00766 | 2015-05-06 | Pass |  | Cluster density a little on the high side
150508_M00766_0096_000000000-AF9MW | M00766 | 2015-05-08 | Pass | One sample a bit low; Some samples high
150508_M02641_0047_000000000-AF7NG | M02641 | 2015-05-08 | Pass | Some samples high
150511_M02641_0048_000000000-AEUGT | M02641 | 2015-05-11 | Fail |  | Cluster density low; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
150516_M00766_0099_000000000-AF81F | M00766 | 2015-05-16 | Pass | One sample a bit low; Some samples high
150521_M02641_0051_000000000-AF9JH | M02641 | 2015-05-21 | Borderline Pass | One sample very low; One sample very high | Cluster density a little low
150522_M00766_0103_000000000-AF9M3 | M00766 | 2015-05-22 | Pass | One sample high
150522_M02641_0052_000000000-AF82F | M02641 | 2015-05-22 | Pass | Two samples high | Reasonable difference between cluster density passing filter and cluster density
150601_M00766_0107_000000000-AF9N8 | M00766 | 2015-06-01 | Borderline Pass | One sample very low; One sample very high | Cluster density quite low
150602_M00766_0108_000000000-AFMB8 | M00766 | 2015-06-02 | Borderline Pass | One sample very high | Cluster density a little high; Reasonable difference between cluster density passing filter and cluster density; Some smearing in heatmap (small)
150603_M02641_0056_000000000-AEY8R | M02641 | 2015-06-03 | Borderline Pass |  | Cluster density a little low; Reasonable difference between cluster density passing filter and cluster density
150604_M00766_0109_000000000-AFJ62 | M00766 | 2015-06-04 | Pass | Some samples high
150604_M02641_0057_000000000-AFL7N | M02641 | 2015-06-04 | Borderline Pass | Some samples low; Some samples high | Cluster density a little low; Slightly bigger than expected difference between cluster density passing filter and cluster density
150605_M00766_0110_000000000-AF80F | M00766 | 2015-06-05 | Borderline Pass | One sample very high | Cluster density a little high
150610_M00766_0111_000000000-AFMWF | M00766 | 2015-06-10 | Pass | One sample very high
150611_M02641_0060_000000000-AFN2Y | M02641 | 2015-06-11 | Borderline Pass | Some samples high | Reasonable difference between cluster density passing filter and cluster density; Quite spread box plots
150615_M02641_0062_000000000-AFH7C | M02641 | 2015-06-15 | Borderline Pass | One sample high | Cluster density a little low; Reasonable difference between cluster density passing filter and cluster density; Quite spread box plots
150619_M02641_0063_000000000-AFN2W | M02641 | 2015-06-19 | Pass | One sample low; One sample high | Cluster density a little low
150625_M02641_0065_000000000-AEY2A | M02641 | 2015-06-25 | Borderline Fail |  | Cluster density a little low; Large difference between cluster density passing filter and cluster density
150703_M00766_0117_000000000-AFN2F | M00766 | 2015-07-03 | Pass | One sample a little high
150703_M02641_0001_000000000-AFMW2 | M02641 | 2015-07-03 | Pass | One sample very high; One sample very low
150709_M02641_0004_000000000-AFMPC | M02641 | 2015-07-09 | Pass | One sample low; Some samples high | Cluster density slightly low
150727_M00766_0119_000000000-AF3TM | M00766 | 2015-07-27 | Pass | Cluster of samples high; Cluster of samples low
150727_M02641_0007_000000000-AGEW7 | M02641 | 2015-07-27 | Fail | One sample a little high; One sample low | Low >Q30; Smearing in heatmap towards ends of reads
150807_M00766_0122_000000000-AFMW7 | M00766 | 2015-08-07 | Borderline Pass |  | Low cluster density
150813_M00766_0124_000000000-AFYEH | M00766 | 2015-08-13 | Pass |  | Cluster density slightly on the low side
150818_M00766_0125_000000000-AE8BP | M00766 | 2015-08-18 | Pass | One sample a little high
150826_M00766_0127_000000000-AGKLT | M00766 | 2015-08-26 | Pass
150904_M00766_0129_000000000-AFNA2 | M00766 | 2015-09-04 | Borderline Fail | One sample high; Some samples low | High cluster density; Slightly bigger than expected difference between cluster density passing filter and cluster density; Smearing in heatmap towards ends of reads
150910_M00766_0130_000000000-AGJGG | M00766 | 2015-09-10 | Pass |  | Slightly bigger than expected difference between cluster density passing filter and cluster density
150916_M02641_0025_000000000-AFN33 | M02641 | 2015-09-16 | Borderline Pass | One sample a little low | Slightly high cluster density; Small amount of smearing towards the ends of reads
150917_M02641_0026_000000000-AFN0A | M02641 | 2015-09-17 | Pass | One sample low | Small amount of smearing towards the ends of reads
150918_M00766_0133_000000000-AH985 | M00766 | 2015-09-18 | Pass
150923_M00766_0134_000000000-AFN32 | M00766 | 2015-09-23 | Pass | One sample very high; One sample very low
150924_M02641_0027_000000000-AGJGM | M02641 | 2015-09-24 | Pass | One sample a little high | Cluster density slightly on the low side
150925_M00766_0135_000000000-AFL75 | M00766 | 2015-09-25 | Borderline Pass | Some samples high; One sample very low | Low cluster density
150929_M02641_0030_000000000-AFWJ3 | M02641 | 2015-09-29 | Pass | Some samples high; Some samples very low
150930_M00766_0138_000000000-AH4BR | M00766 | 2015-09-30 | Borderline Pass | One sample high | Some cycles smearing in heatmap (read 1)
151005_M02641_0034_000000000-AH9ML | M02641 | 2015-10-05 | Borderline Pass | One sample a little high | Cluster density a little high; Some smearing in heatmap towards end of reads (small) amount
151008_M02641_0035_000000000-AFL73 | M02641 | 2015-10-08 | Pass | Some samples very high; Most samples low
151013_M02641_0038_000000000-AFLW0 | M02641 | 2015-10-13 | Pass
151014_M02641_0039_000000000-AFLFV | M02641 | 2015-10-14 | Pass | One sample very low
151015_M00766_0142_000000000-AJD8B | M00766 | 2015-10-15 | Pass | One sample a little high
151021_M02641_0041_000000000-AGHAA | M02641 | 2015-10-21 | Pass | One sample high
151022_M00766_0146_000000000-AFN3B | M00766 | 2015-10-22 | Borderline Pass |  | Cluster density a little high; Reasonable difference between cluster density passing filter and cluster density
151023_M00766_0147_000000000-AFN37 | M00766 | 2015-10-23 | Borderline Pass | One sample high; One sample low | Cluster density a little high
151028_M02641_0043_000000000-AFL74 | M02641 | 2015-10-28 | Pass
151030_M00766_0150_000000000-AFLDT | M00766 | 2015-10-30 | Pass | Some samples low; Some samples high
151030_M02641_0044_000000000-AJHLJ | M02641 | 2015-10-30 | Pass | One sample high; Some samples low
151103_M02641_0046_000000000-AJ5Y7 | M02641 | 2015-11-03 | Pass | One sample high
151105_M02641_0047_000000000-AJF4E | M02641 | 2015-11-05 | Pass | One sample very high; All others low
151116_M02641_0050_000000000-AJ6L3 | M02641 | 2015-11-16 | Pass |  | Cluster density slightly low
151117_M00766_0155_000000000-AJ6LL | M00766 | 2015-11-17 | Borderline Pass | Some samples high; Some samples low | Some smearing towards the ends of reads (small amount)
151117_M02641_0051_000000000-AJF4A | M02641 | 2015-11-17 | Pass | Some samples high; Three samples very low | Low cluster density
151125_M02641_0053_000000000-AJJ3G | M02641 | 2015-11-25 | Borderline Fail | One sample very high
151126_M00766_0159_000000000-AJD7L | M00766 | 2015-11-26 | Pass
151126_M02641_0054_000000000-AJD7T | M02641 | 2015-11-26 | Pass | One sample a little high | Cluster density slightly low
151127_M00766_0161_000000000-AJERP | M00766 | 2015-11-27 | Pass | One sample very low; Some samples high | Cluster density slightly low
151201_M02641_0055_000000000-AJ6B2 | M02641 | 2015-12-01 | Pass | One sample very high
151203_M00766_0165_000000000-AJFAA | M00766 | 2015-12-03 | Pass | Some samples low; Some samples high
151208_M00766_0167_000000000-AJ5JR | M00766 | 2015-12-08 | Borderline Fail | Some samples low; Some samples high | Low cluster density
151210_M02641_0058_000000000-AJ5YJ | M02641 | 2015-12-10 | Pass
151216_M00766_0170_000000000-ALEAF | M00766 | 2015-12-16 | Pass | One sample high; One sample very low
151217_M02641_0059_000000000-AJFAD | M02641 | 2015-12-17 | Fail | Range | Cluster density low; Large difference between cluster density passing filter and cluster density
151223_M00766_0172_000000000-AJRMN | M00766 | 2015-12-23 | Pass | One sample high | Cluster density slightly low
151223_M02641_0060_000000000-AGHV2 | M02641 | 2015-12-23 | Pass | One sample low; One sample high
151231_M02641_0061_000000000-AJD8L | M02641 | 2015-12-31 | Pass | One sample high
160107_M00766_0175_000000000-AH6C3 | M00766 | 2016-01-07 | Pass | One sample high
160108_M00766_0176_000000000-AJDAJ | M00766 | 2016-01-08 | Fail | One sample low; One sample high | Low >Q30; Smearing in heatmap
160114_M00766_0177_000000000-ALYTH | M00766 | 2016-01-14 | Pass | One sample low; Some samples high | Slightly bigger than expected difference between cluster density passing filter and cluster density
160114_M02641_0065_000000000-ALRU0 | M02641 | 2016-01-14 | Pass | Two samples high
160128_M02641_0070_000000000-ALAGE | M02641 | 2016-01-28 | Pass | One sample high | Cluster density slightly low
160129_M00766_0003_000000000-ALAH0 | M00766 | 2016-01-29 | Pass | One sample high
160129_M02641_0071_000000000-AMC8N | M02641 | 2016-01-29 | Pass | One sample low
160206_M00766_0005_000000000-ALYU1 | M00766 | 2016-02-06 | Borderline Pass | Two samples high | Cluster density slightly high; Slightly bigger than expected difference between cluster density passing filter and cluster density
160206_M02641_0074_000000000-AL604 | M02641 | 2016-02-06 | Pass | Some samples very high; Some samples very low
160209_M00766_0007_000000000-AMERJ | M00766 | 2016-02-09 | Borderline Fail | Two samples high | Cluster density high; >Q30 lower than expected; Bigger than expected difference between cluster density passing filter and cluster density
160216_M00766_0009_000000000-AL73W | M00766 | 2016-02-16 | Pass | One sample very high; One sample a little low
160219_M00766_0011_000000000-AL6GM | M00766 | 2016-02-19 | Fail | No info | Cluster density aa little low; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
160220_M02641_0081_000000000-AMETF | M02641 | 2016-02-20 | Pass | One sample very low | Bit of smearing towards the ends of reads (small)
160223_M00766_0012_000000000-AL5YC | M00766 | 2016-02-23 | Pass | One sample very high; Most samples low
123

## FailsPassesAuto
Pass Method 1 | Fail Method 1 | Pass Method 2 | Fail Method 2
0 | 121009_M00766_0002_000000000-A1U6P | 0 | 121009_M00766_0002_000000000-A1U6P | INDEX($A$2:$A$123,SMALL(IF($A$2:$A$123<>0,ROW($A$2:$A$123),ROW(2:2)),2))
0 | 130206_M00766_0002_000000000-A23JM | 0 | 130206_M00766_0002_000000000-A23JM
0 | 130618_M00766_0019_000000000-A4FEU | 0 | 130618_M00766_0019_000000000-A4FEU
130624_M00766_0021_000000000-A53PT | 0 | 130624_M00766_0021_000000000-A53PT | 0
0 | 130712_M00766_0024_000000000-A5AVV | 0 | 130712_M00766_0024_000000000-A5AVV
130817_M00766_0037_000000000-A4EDD | 0 | 130817_M00766_0037_000000000-A4EDD | 0
130913_M00766_0049_000000000-A5B0E | 0 | 130913_M00766_0049_000000000-A5B0E | 0
0 | 130916_M00766_0050_000000000-A5M1U | 0 | 130916_M00766_0050_000000000-A5M1U
0 | 130924_M00766_0052_000000000-A5BE6 | 0 | 130924_M00766_0052_000000000-A5BE6
0 | 131007_M00766_0055_000000000-A59JT | 0 | 131007_M00766_0055_000000000-A59JT
131025_M00766_0059_000000000-A5P9A | 0 | 131025_M00766_0059_000000000-A5P9A | 0
140207_M00766_0022_000000000-A7PH8 | 0 | 140207_M00766_0022_000000000-A7PH8 | 0
0 | 140424_M00766_0033_000000000-A7BNA | 0 | 140424_M00766_0033_000000000-A7BNA
0 | 140514_M00766_0036_000000000-A7BRK | 0 | 140514_M00766_0036_000000000-A7BRK
140612_M00766_0039_000000000-A7BNL | 0 | 0 | 140612_M00766_0039_000000000-A7BNL
0 | 140616_M00766_0040_000000000-A78V9 | 0 | 140616_M00766_0040_000000000-A78V9
140813_M00766_0047_000000000-A8PJL | 0 | 140813_M00766_0047_000000000-A8PJL | 0
140901_M00766_0048_000000000-AA63M | 0 | 140901_M00766_0048_000000000-AA63M | 0
141014_M00766_0053_000000000-A8R6M | 0 | 141014_M00766_0053_000000000-A8R6M | 0
0 | 141017_M02641_0022_000000000-AA66H | 0 | 141017_M02641_0022_000000000-AA66H
0 | 141024_M00766_0056_000000000-A8PC5 | 0 | 141024_M00766_0056_000000000-A8PC5
141031_M00766_0058_000000000-AA3GN | 0 | 141031_M00766_0058_000000000-AA3GN | 0
141107_M00766_0060_000000000-AA8PM | 0 | 141107_M00766_0060_000000000-AA8PM | 0
141118_M00766_0061_000000000-A8P8J | 0 | 141118_M00766_0061_000000000-A8P8J | 0
141208_M00766_0063_000000000-A8P7C | 0 | 141208_M00766_0063_000000000-A8P7C | 0
141208_M02641_0026_000000000-A8R55 | 0 | 0 | 141208_M02641_0026_000000000-A8R55
0 | 141212_M00766_0064_000000000-ACCEB | 0 | 141212_M00766_0064_000000000-ACCEB
0 | 141216_M00766_0065_000000000-ACCDT | 0 | 141216_M00766_0065_000000000-ACCDT
0 | 141219_M00766_0066_000000000-ACCB1 | 0 | 141219_M00766_0066_000000000-ACCB1
150120_M00766_0071_000000000-AA63K | 0 | 150120_M00766_0071_000000000-AA63K | 0
150127_M02641_0027_000000000-AA65J | 0 | 150127_M02641_0027_000000000-AA65J | 0
150203_M00766_0074_000000000-AAUMH | 0 | 150203_M00766_0074_000000000-AAUMH | 0
150203_M02641_0028_000000000-ACBYG | 0 | 0 | 150203_M02641_0028_000000000-ACBYG
150205_M00766_0075_000000000-ACC43 | 0 | 150205_M00766_0075_000000000-ACC43 | 0
0 | 150224_M00766_0078_000000000-ACMP5 | 0 | 150224_M00766_0078_000000000-ACMP5
0 | 150309_M02641_0035_000000000-ACC3J | 0 | 150309_M02641_0035_000000000-ACC3J
150311_M00766_0081_000000000-ACN0W | 0 | 150311_M00766_0081_000000000-ACN0W | 0
150311_M02641_0036_000000000-AD8VB | 0 | 150311_M02641_0036_000000000-AD8VB | 0
150402_M00766_0087_000000000-ACCAY | 0 | 150402_M00766_0087_000000000-ACCAY | 0
150414_M00766_0090_000000000-AAU39 | 0 | 150414_M00766_0090_000000000-AAU39 | 0
0 | 150427_M02641_0042_000000000-AD8LB | 0 | 150427_M02641_0042_000000000-AD8LB
150430_M00766_0092_000000000-AD77P | 0 | 150430_M00766_0092_000000000-AD77P | 0
150506_M00766_0094_000000000-AEVP8 | 0 | 150506_M00766_0094_000000000-AEVP8 | 0
150508_M00766_0096_000000000-AF9MW | 0 | 150508_M00766_0096_000000000-AF9MW | 0
150508_M02641_0047_000000000-AF7NG | 0 | 150508_M02641_0047_000000000-AF7NG | 0
0 | 150511_M02641_0048_000000000-AEUGT | 0 | 150511_M02641_0048_000000000-AEUGT
150516_M00766_0099_000000000-AF81F | 0 | 150516_M00766_0099_000000000-AF81F | 0
0 | 150521_M02641_0051_000000000-AF9JH | 0 | 150521_M02641_0051_000000000-AF9JH
150522_M00766_0103_000000000-AF9M3 | 0 | 150522_M00766_0103_000000000-AF9M3 | 0
0 | 150522_M02641_0052_000000000-AF82F | 0 | 150522_M02641_0052_000000000-AF82F
0 | 150601_M00766_0107_000000000-AF9N8 | 0 | 150601_M00766_0107_000000000-AF9N8
150602_M00766_0108_000000000-AFMB8 | 0 | 150602_M00766_0108_000000000-AFMB8 | 0
0 | 150603_M02641_0056_000000000-AEY8R | 0 | 150603_M02641_0056_000000000-AEY8R
150604_M00766_0109_000000000-AFJ62 | 0 | 150604_M00766_0109_000000000-AFJ62 | 0
0 | 150604_M02641_0057_000000000-AFL7N | 0 | 150604_M02641_0057_000000000-AFL7N
150605_M00766_0110_000000000-AF80F | 0 | 150605_M00766_0110_000000000-AF80F | 0
150610_M00766_0111_000000000-AFMWF | 0 | 150610_M00766_0111_000000000-AFMWF | 0
0 | 150611_M02641_0060_000000000-AFN2Y | 0 | 150611_M02641_0060_000000000-AFN2Y
0 | 150615_M02641_0062_000000000-AFH7C | 0 | 150615_M02641_0062_000000000-AFH7C
150619_M02641_0063_000000000-AFN2W | 0 | 150619_M02641_0063_000000000-AFN2W | 0
0 | 150625_M02641_0065_000000000-AEY2A | 0 | 150625_M02641_0065_000000000-AEY2A
150703_M00766_0117_000000000-AFN2F | 0 | 150703_M00766_0117_000000000-AFN2F | 0
150703_M02641_0001_000000000-AFMW2 | 0 | 150703_M02641_0001_000000000-AFMW2 | 0
150709_M02641_0004_000000000-AFMPC | 0 | 150709_M02641_0004_000000000-AFMPC | 0
150727_M00766_0119_000000000-AF3TM | 0 | 150727_M00766_0119_000000000-AF3TM | 0
150727_M02641_0007_000000000-AGEW7 | 0 | 150727_M02641_0007_000000000-AGEW7 | 0
150807_M00766_0122_000000000-AFMW7 | 0 | 150807_M00766_0122_000000000-AFMW7 | 0
150813_M00766_0124_000000000-AFYEH | 0 | 150813_M00766_0124_000000000-AFYEH | 0
150818_M00766_0125_000000000-AE8BP | 0 | 150818_M00766_0125_000000000-AE8BP | 0
150826_M00766_0127_000000000-AGKLT | 0 | 150826_M00766_0127_000000000-AGKLT | 0
150904_M00766_0129_000000000-AFNA2 | 0 | 0 | 150904_M00766_0129_000000000-AFNA2
150910_M00766_0130_000000000-AGJGG | 0 | 150910_M00766_0130_000000000-AGJGG | 0
150916_M02641_0025_000000000-AFN33 | 0 | 150916_M02641_0025_000000000-AFN33 | 0
150917_M02641_0026_000000000-AFN0A | 0 | 150917_M02641_0026_000000000-AFN0A | 0
150918_M00766_0133_000000000-AH985 | 0 | 150918_M00766_0133_000000000-AH985 | 0
150923_M00766_0134_000000000-AFN32 | 0 | 150923_M00766_0134_000000000-AFN32 | 0
150924_M02641_0027_000000000-AGJGM | 0 | 150924_M02641_0027_000000000-AGJGM | 0
0 | 150925_M00766_0135_000000000-AFL75 | 0 | 150925_M00766_0135_000000000-AFL75
150929_M02641_0030_000000000-AFWJ3 | 0 | 150929_M02641_0030_000000000-AFWJ3 | 0
150930_M00766_0138_000000000-AH4BR | 0 | 150930_M00766_0138_000000000-AH4BR | 0
151005_M02641_0034_000000000-AH9ML | 0 | 151005_M02641_0034_000000000-AH9ML | 0
151008_M02641_0035_000000000-AFL73 | 0 | 151008_M02641_0035_000000000-AFL73 | 0
151013_M02641_0038_000000000-AFLW0 | 0 | 151013_M02641_0038_000000000-AFLW0 | 0
151014_M02641_0039_000000000-AFLFV | 0 | 151014_M02641_0039_000000000-AFLFV | 0
151015_M00766_0142_000000000-AJD8B | 0 | 151015_M00766_0142_000000000-AJD8B | 0
151021_M02641_0041_000000000-AGHAA | 0 | 151021_M02641_0041_000000000-AGHAA | 0
0 | 151022_M00766_0146_000000000-AFN3B | 0 | 151022_M00766_0146_000000000-AFN3B
0 | 151023_M00766_0147_000000000-AFN37 | 0 | 151023_M00766_0147_000000000-AFN37
151028_M02641_0043_000000000-AFL74 | 0 | 151028_M02641_0043_000000000-AFL74 | 0
151030_M00766_0150_000000000-AFLDT | 0 | 151030_M00766_0150_000000000-AFLDT | 0
151030_M02641_0044_000000000-AJHLJ | 0 | 151030_M02641_0044_000000000-AJHLJ | 0
151103_M02641_0046_000000000-AJ5Y7 | 0 | 151103_M02641_0046_000000000-AJ5Y7 | 0
151105_M02641_0047_000000000-AJF4E | 0 | 151105_M02641_0047_000000000-AJF4E | 0
151116_M02641_0050_000000000-AJ6L3 | 0 | 151116_M02641_0050_000000000-AJ6L3 | 0
151117_M00766_0155_000000000-AJ6LL | 0 | 151117_M00766_0155_000000000-AJ6LL | 0
151117_M02641_0051_000000000-AJF4A | 0 | 151117_M02641_0051_000000000-AJF4A | 0
0 | 151125_M02641_0053_000000000-AJJ3G | 0 | 151125_M02641_0053_000000000-AJJ3G
151126_M00766_0159_000000000-AJD7L | 0 | 151126_M00766_0159_000000000-AJD7L | 0
151126_M02641_0054_000000000-AJD7T | 0 | 151126_M02641_0054_000000000-AJD7T | 0
151127_M00766_0161_000000000-AJERP | 0 | 151127_M00766_0161_000000000-AJERP | 0
151201_M02641_0055_000000000-AJ6B2 | 0 | 151201_M02641_0055_000000000-AJ6B2 | 0
151203_M00766_0165_000000000-AJFAA | 0 | 151203_M00766_0165_000000000-AJFAA | 0
0 | 151208_M00766_0167_000000000-AJ5JR | 0 | 151208_M00766_0167_000000000-AJ5JR
151210_M02641_0058_000000000-AJ5YJ | 0 | 151210_M02641_0058_000000000-AJ5YJ | 0
151216_M00766_0170_000000000-ALEAF | 0 | 151216_M00766_0170_000000000-ALEAF | 0
0 | 151217_M02641_0059_000000000-AJFAD | 0 | 151217_M02641_0059_000000000-AJFAD
151223_M00766_0172_000000000-AJRMN | 0 | 151223_M00766_0172_000000000-AJRMN | 0
151223_M02641_0060_000000000-AGHV2 | 0 | 151223_M02641_0060_000000000-AGHV2 | 0
151231_M02641_0061_000000000-AJD8L | 0 | 151231_M02641_0061_000000000-AJD8L | 0
160107_M00766_0175_000000000-AH6C3 | 0 | 160107_M00766_0175_000000000-AH6C3 | 0
0 | 160108_M00766_0176_000000000-AJDAJ | 0 | 160108_M00766_0176_000000000-AJDAJ
160114_M00766_0177_000000000-ALYTH | 0 | 160114_M00766_0177_000000000-ALYTH | 0
160114_M02641_0065_000000000-ALRU0 | 0 | 160114_M02641_0065_000000000-ALRU0 | 0
160128_M02641_0070_000000000-ALAGE | 0 | 160128_M02641_0070_000000000-ALAGE | 0
160129_M00766_0003_000000000-ALAH0 | 0 | 160129_M00766_0003_000000000-ALAH0 | 0
160129_M02641_0071_000000000-AMC8N | 0 | 160129_M02641_0071_000000000-AMC8N | 0
160206_M00766_0005_000000000-ALYU1 | 0 | 160206_M00766_0005_000000000-ALYU1 | 0
160206_M02641_0074_000000000-AL604 | 0 | 160206_M02641_0074_000000000-AL604 | 0
0 | 160209_M00766_0007_000000000-AMERJ | 0 | 160209_M00766_0007_000000000-AMERJ
160216_M00766_0009_000000000-AL73W | 0 | 160216_M00766_0009_000000000-AL73W | 0
0 | 160219_M00766_0011_000000000-AL6GM | 0 | 160219_M00766_0011_000000000-AL6GM
160220_M02641_0081_000000000-AMETF | 0 | 160220_M02641_0081_000000000-AMETF | 0
160223_M00766_0012_000000000-AL5YC | 0 | 160223_M00766_0012_000000000-AL5YC | 0

## FailsPassesManual
Pass | Fail | Borderline Pass | Borderline Fail | Passes | Fails
0 | 121009_M00766_0002_000000000-A1U6P | 0 | 0 | 0 | 121009_M00766_0002_000000000-A1U6P
0 | 130206_M00766_0002_000000000-A23JM | 0 | 0 | 0 | 130206_M00766_0002_000000000-A23JM
0 | 130618_M00766_0019_000000000-A4FEU | 0 | 0 | 0 | 130618_M00766_0019_000000000-A4FEU
130624_M00766_0021_000000000-A53PT | 0 | 0 | 0 | 130624_M00766_0021_000000000-A53PT | 0
0 | 0 | 130712_M00766_0024_000000000-A5AVV | 0 | 130712_M00766_0024_000000000-A5AVV | 0
130817_M00766_0037_000000000-A4EDD | 0 | 0 | 0 | 130817_M00766_0037_000000000-A4EDD | 0
130913_M00766_0049_000000000-A5B0E | 0 | 0 | 0 | 130913_M00766_0049_000000000-A5B0E | 0
130916_M00766_0050_000000000-A5M1U | 0 | 0 | 0 | 130916_M00766_0050_000000000-A5M1U | 0
0 | 130924_M00766_0052_000000000-A5BE6 | 0 | 0 | 0 | 130924_M00766_0052_000000000-A5BE6
0 | 0 | 131007_M00766_0055_000000000-A59JT | 0 | 131007_M00766_0055_000000000-A59JT | 0
131025_M00766_0059_000000000-A5P9A | 0 | 0 | 0 | 131025_M00766_0059_000000000-A5P9A | 0
140207_M00766_0022_000000000-A7PH8 | 0 | 0 | 0 | 140207_M00766_0022_000000000-A7PH8 | 0
0 | 0 | 0 | 140424_M00766_0033_000000000-A7BNA | 0 | 140424_M00766_0033_000000000-A7BNA
0 | 140514_M00766_0036_000000000-A7BRK | 0 | 0 | 0 | 140514_M00766_0036_000000000-A7BRK
0 | 140612_M00766_0039_000000000-A7BNL | 0 | 0 | 0 | 140612_M00766_0039_000000000-A7BNL
0 | 0 | 140616_M00766_0040_000000000-A78V9 | 0 | 140616_M00766_0040_000000000-A78V9 | 0
140813_M00766_0047_000000000-A8PJL | 0 | 0 | 0 | 140813_M00766_0047_000000000-A8PJL | 0
140901_M00766_0048_000000000-AA63M | 0 | 0 | 0 | 140901_M00766_0048_000000000-AA63M | 0
141014_M00766_0053_000000000-A8R6M | 0 | 0 | 0 | 141014_M00766_0053_000000000-A8R6M | 0
0 | 141017_M02641_0022_000000000-AA66H | 0 | 0 | 0 | 141017_M02641_0022_000000000-AA66H
0 | 0 | 0 | 141024_M00766_0056_000000000-A8PC5 | 0 | 141024_M00766_0056_000000000-A8PC5
141031_M00766_0058_000000000-AA3GN | 0 | 0 | 0 | 141031_M00766_0058_000000000-AA3GN | 0
0 | 0 | 141107_M00766_0060_000000000-AA8PM | 0 | 141107_M00766_0060_000000000-AA8PM | 0
141118_M00766_0061_000000000-A8P8J | 0 | 0 | 0 | 141118_M00766_0061_000000000-A8P8J | 0
0 | 0 | 141208_M00766_0063_000000000-A8P7C | 0 | 141208_M00766_0063_000000000-A8P7C | 0
0 | 0 | 141208_M02641_0026_000000000-A8R55 | 0 | 141208_M02641_0026_000000000-A8R55 | 0
0 | 141212_M00766_0064_000000000-ACCEB | 0 | 0 | 0 | 141212_M00766_0064_000000000-ACCEB
0 | 141216_M00766_0065_000000000-ACCDT | 0 | 0 | 0 | 141216_M00766_0065_000000000-ACCDT
0 | 141219_M00766_0066_000000000-ACCB1 | 0 | 0 | 0 | 141219_M00766_0066_000000000-ACCB1
150120_M00766_0071_000000000-AA63K | 0 | 0 | 0 | 150120_M00766_0071_000000000-AA63K | 0
150127_M02641_0027_000000000-AA65J | 0 | 0 | 0 | 150127_M02641_0027_000000000-AA65J | 0
150203_M00766_0074_000000000-AAUMH | 0 | 0 | 0 | 150203_M00766_0074_000000000-AAUMH | 0
0 | 0 | 150203_M02641_0028_000000000-ACBYG | 0 | 150203_M02641_0028_000000000-ACBYG | 0
150205_M00766_0075_000000000-ACC43 | 0 | 0 | 0 | 150205_M00766_0075_000000000-ACC43 | 0
0 | 150224_M00766_0078_000000000-ACMP5 | 0 | 0 | 0 | 150224_M00766_0078_000000000-ACMP5
0 | 150309_M02641_0035_000000000-ACC3J | 0 | 0 | 0 | 150309_M02641_0035_000000000-ACC3J
150311_M00766_0081_000000000-ACN0W | 0 | 0 | 0 | 150311_M00766_0081_000000000-ACN0W | 0
150311_M02641_0036_000000000-AD8VB | 0 | 0 | 0 | 150311_M02641_0036_000000000-AD8VB | 0
0 | 0 | 150402_M00766_0087_000000000-ACCAY | 0 | 150402_M00766_0087_000000000-ACCAY | 0
150414_M00766_0090_000000000-AAU39 | 0 | 0 | 0 | 150414_M00766_0090_000000000-AAU39 | 0
0 | 0 | 0 | 150427_M02641_0042_000000000-AD8LB | 0 | 150427_M02641_0042_000000000-AD8LB
0 | 0 | 150430_M00766_0092_000000000-AD77P | 0 | 150430_M00766_0092_000000000-AD77P | 0
150506_M00766_0094_000000000-AEVP8 | 0 | 0 | 0 | 150506_M00766_0094_000000000-AEVP8 | 0
150508_M00766_0096_000000000-AF9MW | 0 | 0 | 0 | 150508_M00766_0096_000000000-AF9MW | 0
150508_M02641_0047_000000000-AF7NG | 0 | 0 | 0 | 150508_M02641_0047_000000000-AF7NG | 0
0 | 150511_M02641_0048_000000000-AEUGT | 0 | 0 | 0 | 150511_M02641_0048_000000000-AEUGT
150516_M00766_0099_000000000-AF81F | 0 | 0 | 0 | 150516_M00766_0099_000000000-AF81F | 0
0 | 0 | 150521_M02641_0051_000000000-AF9JH | 0 | 150521_M02641_0051_000000000-AF9JH | 0
150522_M00766_0103_000000000-AF9M3 | 0 | 0 | 0 | 150522_M00766_0103_000000000-AF9M3 | 0
150522_M02641_0052_000000000-AF82F | 0 | 0 | 0 | 150522_M02641_0052_000000000-AF82F | 0
0 | 0 | 150601_M00766_0107_000000000-AF9N8 | 0 | 150601_M00766_0107_000000000-AF9N8 | 0
0 | 0 | 150602_M00766_0108_000000000-AFMB8 | 0 | 150602_M00766_0108_000000000-AFMB8 | 0
0 | 0 | 150603_M02641_0056_000000000-AEY8R | 0 | 150603_M02641_0056_000000000-AEY8R | 0
150604_M00766_0109_000000000-AFJ62 | 0 | 0 | 0 | 150604_M00766_0109_000000000-AFJ62 | 0
0 | 0 | 150604_M02641_0057_000000000-AFL7N | 0 | 150604_M02641_0057_000000000-AFL7N | 0
0 | 0 | 150605_M00766_0110_000000000-AF80F | 0 | 150605_M00766_0110_000000000-AF80F | 0
150610_M00766_0111_000000000-AFMWF | 0 | 0 | 0 | 150610_M00766_0111_000000000-AFMWF | 0
0 | 0 | 150611_M02641_0060_000000000-AFN2Y | 0 | 150611_M02641_0060_000000000-AFN2Y | 0
0 | 0 | 150615_M02641_0062_000000000-AFH7C | 0 | 150615_M02641_0062_000000000-AFH7C | 0
150619_M02641_0063_000000000-AFN2W | 0 | 0 | 0 | 150619_M02641_0063_000000000-AFN2W | 0
0 | 0 | 0 | 150625_M02641_0065_000000000-AEY2A | 0 | 150625_M02641_0065_000000000-AEY2A
150703_M00766_0117_000000000-AFN2F | 0 | 0 | 0 | 150703_M00766_0117_000000000-AFN2F | 0
150703_M02641_0001_000000000-AFMW2 | 0 | 0 | 0 | 150703_M02641_0001_000000000-AFMW2 | 0
150709_M02641_0004_000000000-AFMPC | 0 | 0 | 0 | 150709_M02641_0004_000000000-AFMPC | 0
150727_M00766_0119_000000000-AF3TM | 0 | 0 | 0 | 150727_M00766_0119_000000000-AF3TM | 0
0 | 150727_M02641_0007_000000000-AGEW7 | 0 | 0 | 0 | 150727_M02641_0007_000000000-AGEW7
0 | 0 | 150807_M00766_0122_000000000-AFMW7 | 0 | 150807_M00766_0122_000000000-AFMW7 | 0
150813_M00766_0124_000000000-AFYEH | 0 | 0 | 0 | 150813_M00766_0124_000000000-AFYEH | 0
150818_M00766_0125_000000000-AE8BP | 0 | 0 | 0 | 150818_M00766_0125_000000000-AE8BP | 0
150826_M00766_0127_000000000-AGKLT | 0 | 0 | 0 | 150826_M00766_0127_000000000-AGKLT | 0
0 | 0 | 0 | 150904_M00766_0129_000000000-AFNA2 | 0 | 150904_M00766_0129_000000000-AFNA2
150910_M00766_0130_000000000-AGJGG | 0 | 0 | 0 | 150910_M00766_0130_000000000-AGJGG | 0
0 | 0 | 150916_M02641_0025_000000000-AFN33 | 0 | 150916_M02641_0025_000000000-AFN33 | 0
150917_M02641_0026_000000000-AFN0A | 0 | 0 | 0 | 150917_M02641_0026_000000000-AFN0A | 0
150918_M00766_0133_000000000-AH985 | 0 | 0 | 0 | 150918_M00766_0133_000000000-AH985 | 0
150923_M00766_0134_000000000-AFN32 | 0 | 0 | 0 | 150923_M00766_0134_000000000-AFN32 | 0
150924_M02641_0027_000000000-AGJGM | 0 | 0 | 0 | 150924_M02641_0027_000000000-AGJGM | 0
0 | 0 | 150925_M00766_0135_000000000-AFL75 | 0 | 150925_M00766_0135_000000000-AFL75 | 0
150929_M02641_0030_000000000-AFWJ3 | 0 | 0 | 0 | 150929_M02641_0030_000000000-AFWJ3 | 0
0 | 0 | 150930_M00766_0138_000000000-AH4BR | 0 | 150930_M00766_0138_000000000-AH4BR | 0
0 | 0 | 151005_M02641_0034_000000000-AH9ML | 0 | 151005_M02641_0034_000000000-AH9ML | 0
151008_M02641_0035_000000000-AFL73 | 0 | 0 | 0 | 151008_M02641_0035_000000000-AFL73 | 0
151013_M02641_0038_000000000-AFLW0 | 0 | 0 | 0 | 151013_M02641_0038_000000000-AFLW0 | 0
151014_M02641_0039_000000000-AFLFV | 0 | 0 | 0 | 151014_M02641_0039_000000000-AFLFV | 0
151015_M00766_0142_000000000-AJD8B | 0 | 0 | 0 | 151015_M00766_0142_000000000-AJD8B | 0
151021_M02641_0041_000000000-AGHAA | 0 | 0 | 0 | 151021_M02641_0041_000000000-AGHAA | 0
0 | 0 | 151022_M00766_0146_000000000-AFN3B | 0 | 151022_M00766_0146_000000000-AFN3B | 0
0 | 0 | 151023_M00766_0147_000000000-AFN37 | 0 | 151023_M00766_0147_000000000-AFN37 | 0
151028_M02641_0043_000000000-AFL74 | 0 | 0 | 0 | 151028_M02641_0043_000000000-AFL74 | 0
151030_M00766_0150_000000000-AFLDT | 0 | 0 | 0 | 151030_M00766_0150_000000000-AFLDT | 0
151030_M02641_0044_000000000-AJHLJ | 0 | 0 | 0 | 151030_M02641_0044_000000000-AJHLJ | 0
151103_M02641_0046_000000000-AJ5Y7 | 0 | 0 | 0 | 151103_M02641_0046_000000000-AJ5Y7 | 0
151105_M02641_0047_000000000-AJF4E | 0 | 0 | 0 | 151105_M02641_0047_000000000-AJF4E | 0
151116_M02641_0050_000000000-AJ6L3 | 0 | 0 | 0 | 151116_M02641_0050_000000000-AJ6L3 | 0
0 | 0 | 151117_M00766_0155_000000000-AJ6LL | 0 | 151117_M00766_0155_000000000-AJ6LL | 0
151117_M02641_0051_000000000-AJF4A | 0 | 0 | 0 | 151117_M02641_0051_000000000-AJF4A | 0
0 | 0 | 0 | 151125_M02641_0053_000000000-AJJ3G | 0 | 151125_M02641_0053_000000000-AJJ3G
151126_M00766_0159_000000000-AJD7L | 0 | 0 | 0 | 151126_M00766_0159_000000000-AJD7L | 0
151126_M02641_0054_000000000-AJD7T | 0 | 0 | 0 | 151126_M02641_0054_000000000-AJD7T | 0
151127_M00766_0161_000000000-AJERP | 0 | 0 | 0 | 151127_M00766_0161_000000000-AJERP | 0
151201_M02641_0055_000000000-AJ6B2 | 0 | 0 | 0 | 151201_M02641_0055_000000000-AJ6B2 | 0
151203_M00766_0165_000000000-AJFAA | 0 | 0 | 0 | 151203_M00766_0165_000000000-AJFAA | 0
0 | 0 | 0 | 151208_M00766_0167_000000000-AJ5JR | 0 | 151208_M00766_0167_000000000-AJ5JR
151210_M02641_0058_000000000-AJ5YJ | 0 | 0 | 0 | 151210_M02641_0058_000000000-AJ5YJ | 0
151216_M00766_0170_000000000-ALEAF | 0 | 0 | 0 | 151216_M00766_0170_000000000-ALEAF | 0
0 | 151217_M02641_0059_000000000-AJFAD | 0 | 0 | 0 | 151217_M02641_0059_000000000-AJFAD
151223_M00766_0172_000000000-AJRMN | 0 | 0 | 0 | 151223_M00766_0172_000000000-AJRMN | 0
151223_M02641_0060_000000000-AGHV2 | 0 | 0 | 0 | 151223_M02641_0060_000000000-AGHV2 | 0
151231_M02641_0061_000000000-AJD8L | 0 | 0 | 0 | 151231_M02641_0061_000000000-AJD8L | 0
160107_M00766_0175_000000000-AH6C3 | 0 | 0 | 0 | 160107_M00766_0175_000000000-AH6C3 | 0
0 | 160108_M00766_0176_000000000-AJDAJ | 0 | 0 | 0 | 160108_M00766_0176_000000000-AJDAJ
160114_M00766_0177_000000000-ALYTH | 0 | 0 | 0 | 160114_M00766_0177_000000000-ALYTH | 0
160114_M02641_0065_000000000-ALRU0 | 0 | 0 | 0 | 160114_M02641_0065_000000000-ALRU0 | 0
160128_M02641_0070_000000000-ALAGE | 0 | 0 | 0 | 160128_M02641_0070_000000000-ALAGE | 0
160129_M00766_0003_000000000-ALAH0 | 0 | 0 | 0 | 160129_M00766_0003_000000000-ALAH0 | 0
160129_M02641_0071_000000000-AMC8N | 0 | 0 | 0 | 160129_M02641_0071_000000000-AMC8N | 0
0 | 0 | 160206_M00766_0005_000000000-ALYU1 | 0 | 160206_M00766_0005_000000000-ALYU1 | 0
160206_M02641_0074_000000000-AL604 | 0 | 0 | 0 | 160206_M02641_0074_000000000-AL604 | 0
0 | 0 | 0 | 160209_M00766_0007_000000000-AMERJ | 0 | 160209_M00766_0007_000000000-AMERJ
160216_M00766_0009_000000000-AL73W | 0 | 0 | 0 | 160216_M00766_0009_000000000-AL73W | 0
0 | 160219_M00766_0011_000000000-AL6GM | 0 | 0 | 0 | 160219_M00766_0011_000000000-AL6GM
160220_M02641_0081_000000000-AMETF | 0 | 0 | 0 | 160220_M02641_0081_000000000-AMETF | 0
160223_M00766_0012_000000000-AL5YC | 0 | 0 | 0 | 160223_M00766_0012_000000000-AL5YC | 0

## CompareManual1Auto
Intentionally blank as this method did not incorporate individual reads information

## CompareManual2Auto
Pass Method 2 (Automated) | Passes | same/diff | run name for lookup | borderline? | auto wrong (to pass)? | false +ves | false -ves | true +ves | true -ves | Fail Method 2 (Automated) | Fails | same/diff | run name for lookup | borderline? | auto wrong? (to fail) | what test(s) triggered?
0 | 0 | same |  |  |  | no | no | yes | no | 121009_M00766_0002_000000000-A1U6P | 121009_M00766_0002_000000000-A1U6P | same
0 | 0 | same |  |  |  | no | no | yes | no | 130206_M00766_0002_000000000-A23JM | 130206_M00766_0002_000000000-A23JM | same
0 | 0 | same |  |  |  | no | no | yes | no | 130618_M00766_0019_000000000-A4FEU | 130618_M00766_0019_000000000-A4FEU | same
130624_M00766_0021_000000000-A53PT | 130624_M00766_0021_000000000-A53PT | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered | Under Q30 threshold triggered | Under Q30 threshold triggered R2
0 | 130712_M00766_0024_000000000-A5AVV | diff | 130712_M00766_0024_000000000-A5AVV | yes | no | yes | no | no | no | 130712_M00766_0024_000000000-A5AVV | 0 | diff | 130712_M00766_0024_000000000-A5AVV | no | yes | Under Q30 threshold triggered R2 | Read Slope R2 triggered (new test)
130817_M00766_0037_000000000-A4EDD | 130817_M00766_0037_000000000-A4EDD | same |  |  |  | no | no | no | yes | 0 | 0 | same
130913_M00766_0049_000000000-A5B0E | 130913_M00766_0049_000000000-A5B0E | same |  |  |  | no | no | no | yes | 0 | 0 | same
130916_M00766_0050_000000000-A5M1U | 130916_M00766_0050_000000000-A5M1U | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Under Q30 threshold triggered R2
0 | 0 | same |  |  |  | no | no | yes | no | 130924_M00766_0052_000000000-A5BE6 | 130924_M00766_0052_000000000-A5BE6 | same |  |  |  | Under Q30 threshold triggered R2
0 | 131007_M00766_0055_000000000-A59JT | diff | 131007_M00766_0055_000000000-A59JT | yes | no | yes | no | no | no | 131007_M00766_0055_000000000-A59JT | 0 | diff | 131007_M00766_0055_000000000-A59JT |  | yes
131025_M00766_0059_000000000-A5P9A | 131025_M00766_0059_000000000-A5P9A | same |  |  |  | no | no | no | yes | 0 | 0 | same
140207_M00766_0022_000000000-A7PH8 | 140207_M00766_0022_000000000-A7PH8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 140424_M00766_0033_000000000-A7BNA | 140424_M00766_0033_000000000-A7BNA | same |  |  |  | Under Q30 threshold triggered R2
0 | 0 | same |  |  |  | no | no | yes | no | 140514_M00766_0036_000000000-A7BRK | 140514_M00766_0036_000000000-A7BRK | same |  |  |  | COV cluster density triggered
0 | 0 | same |  |  |  | no | no | yes | no | 140612_M00766_0039_000000000-A7BNL | 140612_M00766_0039_000000000-A7BNL | same
140616_M00766_0040_000000000-A78V9 | 140616_M00766_0040_000000000-A78V9 | same |  |  |  | no | no | no | yes | 0 | 0 | same
140813_M00766_0047_000000000-A8PJL | 140813_M00766_0047_000000000-A8PJL | same |  |  |  | no | no | no | yes | 0 | 0 | same
140901_M00766_0048_000000000-AA63M | 140901_M00766_0048_000000000-AA63M | same |  |  |  | no | no | no | yes | 0 | 0 | same
141014_M00766_0053_000000000-A8R6M | 141014_M00766_0053_000000000-A8R6M | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered | Under Q30 threshold triggered R2
0 | 0 | same |  |  |  | no | no | yes | no | 141017_M02641_0022_000000000-AA66H | 141017_M02641_0022_000000000-AA66H | same
141024_M00766_0056_000000000-A8PC5 | 0 | diff | 141024_M00766_0056_000000000-A8PC5 | no | yes | no | yes | no | no | 0 | 141024_M00766_0056_000000000-A8PC5 | diff | 141024_M00766_0056_000000000-A8PC5 |  | no | % reads passing filter triggered | Cluster density pf gap triggered
141031_M00766_0058_000000000-AA3GN | 141031_M00766_0058_000000000-AA3GN | same |  |  |  | no | no | no | yes | 0 | 0 | same
141107_M00766_0060_000000000-AA8PM | 141107_M00766_0060_000000000-AA8PM | same |  |  |  | no | no | no | yes | 0 | 0 | same
141118_M00766_0061_000000000-A8P8J | 141118_M00766_0061_000000000-A8P8J | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Cluster density pf gap triggered | Under Q30 threshold triggered | Under Q30 threshold triggered R1 | Under Q30 threshold triggered R2
141208_M00766_0063_000000000-A8P7C | 141208_M00766_0063_000000000-A8P7C | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Under Q30 threshold triggered R2
0 | 141208_M02641_0026_000000000-A8R55 | diff | 141208_M02641_0026_000000000-A8R55 | yes | no | yes | no | no | no | 141208_M02641_0026_000000000-A8R55 | 0 | diff | 141208_M02641_0026_000000000-A8R55 |  | yes | % reads passing filter triggered | COV cluster density triggered | Cluster density pf gap triggered | Under Q30 threshold triggered | Under Q30 threshold triggered R1 | Under Q30 threshold triggered R2
0 | 0 | same |  |  |  | no | no | yes | no | 141212_M00766_0064_000000000-ACCEB | 141212_M00766_0064_000000000-ACCEB | same
0 | 0 | same |  |  |  | no | no | yes | no | 141216_M00766_0065_000000000-ACCDT | 141216_M00766_0065_000000000-ACCDT | same |  |  |  | COV cluster density triggered
0 | 0 | same |  |  |  | no | no | yes | no | 141219_M00766_0066_000000000-ACCB1 | 141219_M00766_0066_000000000-ACCB1 | same
150120_M00766_0071_000000000-AA63K | 150120_M00766_0071_000000000-AA63K | same |  |  |  | no | no | no | yes | 0 | 0 | same
150127_M02641_0027_000000000-AA65J | 150127_M02641_0027_000000000-AA65J | same |  |  |  | no | no | no | yes | 0 | 0 | same
150203_M00766_0074_000000000-AAUMH | 150203_M00766_0074_000000000-AAUMH | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150203_M02641_0028_000000000-ACBYG | diff | 150203_M02641_0028_000000000-ACBYG | yes | no | yes | no | no | no | 150203_M02641_0028_000000000-ACBYG | 0 | diff | 150203_M02641_0028_000000000-ACBYG |  | yes
150205_M00766_0075_000000000-ACC43 | 150205_M00766_0075_000000000-ACC43 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150224_M00766_0078_000000000-ACMP5 | 0 | diff | 150224_M00766_0078_000000000-ACMP5 | no | yes | no | yes | no | no | 0 | 150224_M00766_0078_000000000-ACMP5 | diff | 150224_M00766_0078_000000000-ACMP5 |  | no
0 | 0 | same |  |  |  | no | no | yes | no | 150309_M02641_0035_000000000-ACC3J | 150309_M02641_0035_000000000-ACC3J | same |  |  |  | Under Q30 threshold triggered R2
150311_M00766_0081_000000000-ACN0W | 150311_M00766_0081_000000000-ACN0W | same |  |  |  | no | no | no | yes | 0 | 0 | same
150311_M02641_0036_000000000-AD8VB | 150311_M02641_0036_000000000-AD8VB | same |  |  |  | no | no | no | yes | 0 | 0 | same
150402_M00766_0087_000000000-ACCAY | 150402_M00766_0087_000000000-ACCAY | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
150414_M00766_0090_000000000-AAU39 | 150414_M00766_0090_000000000-AAU39 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150427_M02641_0042_000000000-AD8LB | 0 | diff | 150427_M02641_0042_000000000-AD8LB | no | yes | no | yes | no | no | 0 | 150427_M02641_0042_000000000-AD8LB | diff | 150427_M02641_0042_000000000-AD8LB |  | no
150430_M00766_0092_000000000-AD77P | 150430_M00766_0092_000000000-AD77P | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Cluster density pf gap triggered
150506_M00766_0094_000000000-AEVP8 | 150506_M00766_0094_000000000-AEVP8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150508_M00766_0096_000000000-AF9MW | 150508_M00766_0096_000000000-AF9MW | same |  |  |  | no | no | no | yes | 0 | 0 | same
150508_M02641_0047_000000000-AF7NG | 150508_M02641_0047_000000000-AF7NG | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150511_M02641_0048_000000000-AEUGT | 150511_M02641_0048_000000000-AEUGT | same
150516_M00766_0099_000000000-AF81F | 150516_M00766_0099_000000000-AF81F | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150521_M02641_0051_000000000-AF9JH | diff | 150521_M02641_0051_000000000-AF9JH |  | no | yes | no | no | no | 150521_M02641_0051_000000000-AF9JH | 0 | diff | 150521_M02641_0051_000000000-AF9JH |  | yes | % reads passing filter triggered | COV cluster density triggered | Cluster density pf gap triggered | Under Q30 threshold triggered | Under Q30 threshold triggered R1 | Under Q30 threshold triggered R2
150522_M00766_0103_000000000-AF9M3 | 150522_M00766_0103_000000000-AF9M3 | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  | yes
0 | 150522_M02641_0052_000000000-AF82F | diff | 150522_M02641_0052_000000000-AF82F |  | no | yes | no | no | no | 150522_M02641_0052_000000000-AF82F | 0 | diff | 150522_M02641_0052_000000000-AF82F |  | yes
150601_M00766_0107_000000000-AF9N8 | 150601_M00766_0107_000000000-AF9N8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150602_M00766_0108_000000000-AFMB8 | 150602_M00766_0108_000000000-AFMB8 | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Cluster density pf gap triggered
0 | 150603_M02641_0056_000000000-AEY8R | diff | 150612_M00766_0112_000000000-AFMW5 |  | no | yes | no | no | no | 150603_M02641_0056_000000000-AEY8R | 0 | diff | 150603_M02641_0056_000000000-AEY8R |  | yes
150604_M00766_0109_000000000-AFJ62 | 150604_M00766_0109_000000000-AFJ62 | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
0 | 150604_M02641_0057_000000000-AFL7N | diff | 150604_M02641_0057_000000000-AFL7N |  | no | yes | no | no | no | 150604_M02641_0057_000000000-AFL7N | 0 | diff | 150604_M02641_0057_000000000-AFL7N |  | yes | Under Q30 threshold triggered | Under Q30 threshold triggered R2
150605_M00766_0110_000000000-AF80F | 150605_M00766_0110_000000000-AF80F | same |  |  |  | no | no | no | yes | 0 | 0 | same
150610_M00766_0111_000000000-AFMWF | 150610_M00766_0111_000000000-AFMWF | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150611_M02641_0060_000000000-AFN2Y | diff | 150611_M02641_0060_000000000-AFN2Y |  | no | yes | no | no | no | 150611_M02641_0060_000000000-AFN2Y | 0 | diff | 150611_M02641_0060_000000000-AFN2Y |  | yes | Under Q30 threshold triggered R2
0 | 150615_M02641_0062_000000000-AFH7C | diff | 150615_M02641_0062_000000000-AFH7C |  | no | yes | no | no | no | 150615_M02641_0062_000000000-AFH7C | 0 | diff | 150615_M02641_0062_000000000-AFH7C |  | yes
150619_M02641_0063_000000000-AFN2W | 150619_M02641_0063_000000000-AFN2W | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150625_M02641_0065_000000000-AEY2A | 150625_M02641_0065_000000000-AEY2A | same |  |  |  | COV cluster density triggered
150703_M00766_0117_000000000-AFN2F | 150703_M00766_0117_000000000-AFN2F | same |  |  |  | no | no | no | yes | 0 | 0 | same
150703_M02641_0001_000000000-AFMW2 | 150703_M02641_0001_000000000-AFMW2 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150709_M02641_0004_000000000-AFMPC | 150709_M02641_0004_000000000-AFMPC | same |  |  |  | no | no | no | yes | 0 | 0 | same
150727_M00766_0119_000000000-AF3TM | 150727_M00766_0119_000000000-AF3TM | same |  |  |  | no | no | no | yes | 0 | 0 | same
150727_M02641_0007_000000000-AGEW7 | 0 | diff | 150727_M02641_0007_000000000-AGEW7 |  | yes | no | yes | no | no | 0 | 150727_M02641_0007_000000000-AGEW7 | diff | 150727_M02641_0007_000000000-AGEW7 |  | no
150807_M00766_0122_000000000-AFMW7 | 150807_M00766_0122_000000000-AFMW7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150813_M00766_0124_000000000-AFYEH | 150813_M00766_0124_000000000-AFYEH | same |  |  |  | no | no | no | yes | 0 | 0 | same
150818_M00766_0125_000000000-AE8BP | 150818_M00766_0125_000000000-AE8BP | same |  |  |  | no | no | no | yes | 0 | 0 | same
150826_M00766_0127_000000000-AGKLT | 150826_M00766_0127_000000000-AGKLT | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150904_M00766_0129_000000000-AFNA2 | 150904_M00766_0129_000000000-AFNA2 | same |  |  |  | Under Q30 threshold triggered R2
150910_M00766_0130_000000000-AGJGG | 150910_M00766_0130_000000000-AGJGG | same |  |  |  | no | no | no | yes | 0 | 0 | same
150916_M02641_0025_000000000-AFN33 | 150916_M02641_0025_000000000-AFN33 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150917_M02641_0026_000000000-AFN0A | 150917_M02641_0026_000000000-AFN0A | same |  |  |  | no | no | no | yes | 0 | 0 | same
150918_M00766_0133_000000000-AH985 | 150918_M00766_0133_000000000-AH985 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150923_M00766_0134_000000000-AFN32 | 150923_M00766_0134_000000000-AFN32 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150924_M02641_0027_000000000-AGJGM | 150924_M02641_0027_000000000-AGJGM | same |  |  |  | no | no | no | yes | 0 | 0 | same
150925_M00766_0135_000000000-AFL75 | 150925_M00766_0135_000000000-AFL75 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150929_M02641_0030_000000000-AFWJ3 | 150929_M02641_0030_000000000-AFWJ3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150930_M00766_0138_000000000-AH4BR | 150930_M00766_0138_000000000-AH4BR | same |  |  |  | no | no | no | yes | 0 | 0 | same
151005_M02641_0034_000000000-AH9ML | 151005_M02641_0034_000000000-AH9ML | same |  |  |  | no | no | no | yes | 0 | 0 | same
151008_M02641_0035_000000000-AFL73 | 151008_M02641_0035_000000000-AFL73 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151013_M02641_0038_000000000-AFLW0 | 151013_M02641_0038_000000000-AFLW0 | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Read Slope R2 triggered (new test)
151014_M02641_0039_000000000-AFLFV | 151014_M02641_0039_000000000-AFLFV | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Read Slope R2 triggered (new test)
151015_M00766_0142_000000000-AJD8B | 151015_M00766_0142_000000000-AJD8B | same |  |  |  | no | no | no | yes | 0 | 0 | same
151021_M02641_0041_000000000-AGHAA | 151021_M02641_0041_000000000-AGHAA | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Cluster density pf gap triggered
151022_M00766_0146_000000000-AFN3B | 151022_M00766_0146_000000000-AFN3B | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Under Q30 threshold triggered | Under Q30 threshold triggered R2
151023_M00766_0147_000000000-AFN37 | 151023_M00766_0147_000000000-AFN37 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151028_M02641_0043_000000000-AFL74 | 151028_M02641_0043_000000000-AFL74 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151030_M00766_0150_000000000-AFLDT | 151030_M00766_0150_000000000-AFLDT | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | COV cluster density triggered | Cluster density pf gap triggered
151030_M02641_0044_000000000-AJHLJ | 151030_M02641_0044_000000000-AJHLJ | same |  |  |  | no | no | no | yes | 0 | 0 | same
151103_M02641_0046_000000000-AJ5Y7 | 151103_M02641_0046_000000000-AJ5Y7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151105_M02641_0047_000000000-AJF4E | 151105_M02641_0047_000000000-AJF4E | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | COV cluster density triggered
151116_M02641_0050_000000000-AJ6L3 | 151116_M02641_0050_000000000-AJ6L3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151117_M00766_0155_000000000-AJ6LL | 151117_M00766_0155_000000000-AJ6LL | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
151117_M02641_0051_000000000-AJF4A | 151117_M02641_0051_000000000-AJF4A | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
151125_M02641_0053_000000000-AJJ3G | 0 | diff | 151125_M02641_0053_000000000-AJJ3G |  | yes | no | yes | no | no | 0 | 151125_M02641_0053_000000000-AJJ3G | diff | 151125_M02641_0053_000000000-AJJ3G |  | no
151126_M00766_0159_000000000-AJD7L | 151126_M00766_0159_000000000-AJD7L | same |  |  |  | no | no | no | yes | 0 | 0 | same
151126_M02641_0054_000000000-AJD7T | 151126_M02641_0054_000000000-AJD7T | same |  |  |  | no | no | no | yes | 0 | 0 | same
151127_M00766_0161_000000000-AJERP | 151127_M00766_0161_000000000-AJERP | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Under Q30 threshold triggered | Under Q30 threshold triggered R1 | Under Q30 threshold triggered R2
151201_M02641_0055_000000000-AJ6B2 | 151201_M02641_0055_000000000-AJ6B2 | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Under Q30 threshold triggered | Under Q30 threshold triggered R1 | Under Q30 threshold triggered R2
151203_M00766_0165_000000000-AJFAA | 151203_M00766_0165_000000000-AJFAA | same |  |  |  | no | no | no | yes | 0 | 0 | same
151208_M00766_0167_000000000-AJ5JR | 0 | diff | 151208_M00766_0167_000000000-AJ5JR |  | yes | no | yes | no | no | 0 | 151208_M00766_0167_000000000-AJ5JR | diff | 151208_M00766_0167_000000000-AJ5JR |  | no
151210_M02641_0058_000000000-AJ5YJ | 151210_M02641_0058_000000000-AJ5YJ | same |  |  |  | no | no | no | yes | 0 | 0 | same
151216_M00766_0170_000000000-ALEAF | 151216_M00766_0170_000000000-ALEAF | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 151217_M02641_0059_000000000-AJFAD | 151217_M02641_0059_000000000-AJFAD | same |  |  |  | COV cluster density triggered | Cluster density pf gap triggered
151223_M00766_0172_000000000-AJRMN | 151223_M00766_0172_000000000-AJRMN | same |  |  |  | no | no | no | yes | 0 | 0 | same
151223_M02641_0060_000000000-AGHV2 | 151223_M02641_0060_000000000-AGHV2 | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | COV cluster density triggered | Under Q30 threshold triggered R2
151231_M02641_0061_000000000-AJD8L | 151231_M02641_0061_000000000-AJD8L | same |  |  |  | no | no | no | yes | 0 | 0 | same
160107_M00766_0175_000000000-AH6C3 | 160107_M00766_0175_000000000-AH6C3 | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | % reads passing filter triggered | COV cluster density triggered | Cluster density pf gap triggered
0 | 0 | same |  |  |  | no | no | yes | no | 160108_M00766_0176_000000000-AJDAJ | 160108_M00766_0176_000000000-AJDAJ | same
160114_M00766_0177_000000000-ALYTH | 160114_M00766_0177_000000000-ALYTH | same |  |  |  | no | no | no | yes | 0 | 0 | same
160114_M02641_0065_000000000-ALRU0 | 160114_M02641_0065_000000000-ALRU0 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160128_M02641_0070_000000000-ALAGE | 160128_M02641_0070_000000000-ALAGE | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | Cluster density pf gap triggered
160129_M00766_0003_000000000-ALAH0 | 160129_M00766_0003_000000000-ALAH0 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160129_M02641_0071_000000000-AMC8N | 160129_M02641_0071_000000000-AMC8N | same |  |  |  | no | no | no | yes | 0 | 0 | same |  |  |  | % reads passing filter triggered | Cluster density pf gap triggered
160206_M00766_0005_000000000-ALYU1 | 160206_M00766_0005_000000000-ALYU1 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160206_M02641_0074_000000000-AL604 | 160206_M02641_0074_000000000-AL604 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160209_M00766_0007_000000000-AMERJ | 0 | diff | 160209_M00766_0007_000000000-AMERJ |  | yes | no | yes | no | no | 0 | 160209_M00766_0007_000000000-AMERJ | diff | 160209_M00766_0007_000000000-AMERJ |  | no
160216_M00766_0009_000000000-AL73W | 160216_M00766_0009_000000000-AL73W | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 160219_M00766_0011_000000000-AL6GM | 160219_M00766_0011_000000000-AL6GM | same
160220_M02641_0081_000000000-AMETF | 160220_M02641_0081_000000000-AMETF | same |  |  |  | no | no | no | yes | 0 | 0 | same
160223_M00766_0012_000000000-AL5YC | 160223_M00766_0012_000000000-AL5YC | same |  |  |  | no | no | no | yes | 0 | 0 | same
[6 empty] | 10 | 7 | 18 | 88
| same pass | 106 | [8 empty] | same fail | 106
| diff pass | 17 | [8 empty] | diff fail | 17
| false passes by automated | 7 | [8 empty] | false fails by automated | 11
| missed passes by automated | 10 | [8 empty] | missed fails by automated | 7
[10 empty] | concordant | 106
[10 empty] | not concordant | 17
[9 empty] | Failing | false +ves | 10
[10 empty] | false -ves | 7
[10 empty] | true +ves | 18
[10 empty] | true -ves | 88
[10 empty] | total | 123
[12 empty] | 95% CI | 95% CI
[10 empty] | sensitivity (true +ve rate) | 72 | 50.61 | 87.93
[10 empty] | specificity (true -ve rate) | 89.7959 | 82.03 | 95
[6 empty] | Removed false manual fail |  |  | Failing | false +ves | 10
[10 empty] | false -ves | 6
[10 empty] | true +ves | 18
[10 empty] | true -ves | 89
[10 empty] | total | 123
[10 empty] | sensitivity (true +ve rate) | 75 | 53.29 | 90.23
[10 empty] | specificity (true -ve rate) | 89.899 | 82.21 | 95.05

## EyeAutoFails
RunID | InstrumentID | RunDate | Pass/Fail | Indexing Comments | General Comments
| M00766 | 2012-10-09 | Fail | No info | Large difference in cluster density passing filter and cluster density; Smearing in heatmap
| M00766 | 2013-02-06 | Fail |  | Cluster density low; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
| M00766 | 2013-06-18 | Fail |  | Cluster density low; Low >Q30; Smearing in heatmap
| M00766 | 2013-06-24 | Pass |  | Some cycles smearing in heatmap
130712_M00766_0024_000000000-A5AVV | M00766 | 2013-07-12 | Borderline Pass | One sample a little low | Some cycles smearing in heatmap; Reasonable difference in cluster density passing filter and cluster density
| M00766 | 2013-08-17 | Pass | Some samples low; Some samples high
| M00766 | 2013-09-13 | Pass |  | Some cycles smearing in heatmap
| M00766 | 2013-09-16 | Pass | One sample low
| M00766 | 2013-09-24 | Fail | Two samples big range | Cluster density a bit high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
131007_M00766_0055_000000000-A59JT | M00766 | 2013-10-07 | Borderline Pass | Two samples big range | Cluster density a bit low; Reasonable difference in cluster density passing filter and cluster density; Some smearing in heatmap
| M00766 | 2013-10-25 | Pass | Some samples very low; Some samples very high
| M00766 | 2014-02-07 | Pass | Two samples big range | Some smearing in heatmap towards end of reads (small) amount
| M00766 | 2014-04-24 | Borderline Fail | One sample high | Very low cluster density
| M00766 | 2014-05-14 | Fail | Two samples high; Two samples low | Cluster density high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
| M00766 | 2014-06-12 | Fail |  | Low >Q30; Smearing in heatmap
| M00766 | 2014-06-16 | Borderline Pass |  | Cluster density a little low; Some smearing in heatmap
| M00766 | 2014-08-13 | Pass | One sample a little high
| M00766 | 2014-09-01 | Pass | One sample a little high; One sample a little low
| M00766 | 2014-10-14 | Pass |  | Some small amount of smearing towards the ends of reads
| M02641 | 2014-10-17 | Fail |  | Cluster density high; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
141024_M00766_0056_000000000-A8PC5 | M00766 | 2014-10-24 | Borderline Fail |  | Cluster density high; Reasonable difference in cluster density passing filter and cluster density; Some smearing in heatmap
| M00766 | 2014-10-31 | Pass |  | Cluster density could be a bit low
| M00766 | 2014-11-07 | Borderline Pass | Some samples high; Some samples low | Cluster density could be a bit low
| M00766 | 2014-11-18 | Pass
| M00766 | 2014-12-08 | Borderline Pass | Two samples high | Cluster density a little high; Some smearing in heatmap towards end of reads (small) amount
141208_M02641_0026_000000000-A8R55 | M02641 | 2014-12-08 | Borderline Pass | One sample a little high; One sample a little low | Cluster density a little low; Smearing in heatmap particularly towards the end of read 2
| M00766 | 2014-12-12 | Fail | Some samples high; Some samples low; NTC high | Cluster density high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
| M00766 | 2014-12-16 | Fail | One sample low; One sample very high | Cluster density high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
| M00766 | 2014-12-19 | Fail | Some samples a little low; One sample very high | Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
| M00766 | 2015-01-20 | Pass | One sample very high | Some small amount of smearing towards the ends of reads
| M02641 | 2015-01-27 | Pass | One sample low
| M00766 | 2015-02-03 | Pass | One sample very low; Some samples high | Reasonable difference in cluster density passing filter and cluster density
150203_M02641_0028_000000000-ACBYG | M02641 | 2015-02-03 | Borderline Pass | One sample very low; One sample very high | Reasonable difference in cluster density passing filter and cluster density; Some smearing in heatmap towards the ends of reads
| M00766 | 2015-02-05 | Pass | Some samples very low; One sample high | Some smearing in heatmap towards the ends of reads
150224_M00766_0078_000000000-ACMP5 | M00766 | 2015-02-24 | Fail |  | High cluster density; Large difference in cluster density passing filter and cluster density
| M02641 | 2015-03-09 | Fail | One sample low | Cluster density a little high; Low >Q30; Reasonable difference between cluster density passing filter and cluster density; Smearing in heatmap
| M00766 | 2015-03-11 | Pass | Some samples low; Some samples high
| M02641 | 2015-03-11 | Pass | One sample high | Cluster density a little low
| M00766 | 2015-04-02 | Borderline Pass | Some samples high | Some smearing in heatmap towards end of reads
| M00766 | 2015-04-14 | Pass | One sample very low
150427_M02641_0042_000000000-AD8LB | M02641 | 2015-04-27 | Borderline Fail | One sample very low; One sample very high | Very low cluster density
| M00766 | 2015-04-30 | Borderline Pass | Some samples high | Cluster density a little high; Reasonable difference between cluster density passing filter and cluster density
| M00766 | 2015-05-06 | Pass |  | Cluster density a little on the high side
| M00766 | 2015-05-08 | Pass | One sample a bit low; Some samples high
| M02641 | 2015-05-08 | Pass | Some samples high
| M02641 | 2015-05-11 | Fail |  | Cluster density low; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
| M00766 | 2015-05-16 | Pass | One sample a bit low; Some samples high
150521_M02641_0051_000000000-AF9JH | M02641 | 2015-05-21 | Borderline Pass | One sample very low; One sample very high | Cluster density a little low
| M00766 | 2015-05-22 | Pass | One sample high
150522_M02641_0052_000000000-AF82F | M02641 | 2015-05-22 | Pass | Two samples high | Reasonable difference between cluster density passing filter and cluster density
| M00766 | 2015-06-01 | Borderline Pass | One sample very low; One sample very high | Cluster density quite low
| M00766 | 2015-06-02 | Borderline Pass | One sample very high | Cluster density a little high; Reasonable difference between cluster density passing filter and cluster density; Some smearing in heatmap (small)
150612_M00766_0112_000000000-AFMW5 | M02641 | 2015-06-03 | Borderline Pass |  | Cluster density a little low; Reasonable difference between cluster density passing filter and cluster density
| M00766 | 2015-06-04 | Pass | Some samples high
150604_M02641_0057_000000000-AFL7N | M02641 | 2015-06-04 | Borderline Pass | Some samples low; Some samples high | Cluster density a little low; Slightly bigger than expected difference between cluster density passing filter and cluster density
| M00766 | 2015-06-05 | Borderline Pass | One sample very high | Cluster density a little high
| M00766 | 2015-06-10 | Pass | One sample very high
150611_M02641_0060_000000000-AFN2Y | M02641 | 2015-06-11 | Borderline Pass | Some samples high | Reasonable difference between cluster density passing filter and cluster density; Quite spread box plots
150615_M02641_0062_000000000-AFH7C | M02641 | 2015-06-15 | Borderline Pass | One sample high | Cluster density a little low; Reasonable difference between cluster density passing filter and cluster density; Quite spread box plots
| M02641 | 2015-06-19 | Pass | One sample low; One sample high | Cluster density a little low
| M02641 | 2015-06-25 | Borderline Fail |  | Cluster density a little low; Large difference between cluster density passing filter and cluster density
| M00766 | 2015-07-03 | Pass | One sample a little high
| M02641 | 2015-07-03 | Pass | One sample very high; One sample very low
| M02641 | 2015-07-09 | Pass | One sample low; Some samples high | Cluster density slightly low
| M00766 | 2015-07-27 | Pass | Cluster of samples high; Cluster of samples low
150727_M02641_0007_000000000-AGEW7 | M02641 | 2015-07-27 | Fail | One sample a little high; One sample low | Low >Q30; Smearing in heatmap towards ends of reads
| M00766 | 2015-08-07 | Borderline Pass |  | Low cluster density
| M00766 | 2015-08-13 | Pass |  | Cluster density slightly on the low side
| M00766 | 2015-08-18 | Pass | One sample a little high
| M00766 | 2015-08-26 | Pass
| M00766 | 2015-09-04 | Borderline Fail | One sample high; Some samples low | High cluster density; Slightly bigger than expected difference between cluster density passing filter and cluster density; Smearing in heatmap towards ends of reads
| M00766 | 2015-09-10 | Pass |  | Slightly bigger than expected difference between cluster density passing filter and cluster density
| M02641 | 2015-09-16 | Borderline Pass | One sample a little low | Slightly high cluster density; Small amount of smearing towards the ends of reads
| M02641 | 2015-09-17 | Pass | One sample low | Small amount of smearing towards the ends of reads
| M00766 | 2015-09-18 | Pass
| M00766 | 2015-09-23 | Pass | One sample very high; One sample very low
| M02641 | 2015-09-24 | Pass | One sample a little high | Cluster density slightly on the low side
| M00766 | 2015-09-25 | Borderline Pass | Some samples high; One sample very low | Low cluster density
| M02641 | 2015-09-29 | Pass | Some samples high; Some samples very low
| M00766 | 2015-09-30 | Borderline Pass | One sample high | Some cycles smearing in heatmap (read 1)
| M02641 | 2015-10-05 | Borderline Pass | One sample a little high | Cluster density a little high; Some smearing in heatmap towards end of reads (small) amount
| M02641 | 2015-10-08 | Pass | Some samples very high; Most samples low
| M02641 | 2015-10-13 | Pass
| M02641 | 2015-10-14 | Pass | One sample very low
| M00766 | 2015-10-15 | Pass | One sample a little high
| M02641 | 2015-10-21 | Pass | One sample high
| M00766 | 2015-10-22 | Borderline Pass |  | Cluster density a little high; Reasonable difference between cluster density passing filter and cluster density
| M00766 | 2015-10-23 | Borderline Pass | One sample high; One sample low | Cluster density a little high
| M02641 | 2015-10-28 | Pass
| M00766 | 2015-10-30 | Pass | Some samples low; Some samples high
| M02641 | 2015-10-30 | Pass | One sample high; Some samples low
| M02641 | 2015-11-03 | Pass | One sample high
| M02641 | 2015-11-05 | Pass | One sample very high; All others low
| M02641 | 2015-11-16 | Pass |  | Cluster density slightly low
| M00766 | 2015-11-17 | Borderline Pass | Some samples high; Some samples low | Some smearing towards the ends of reads (small amount)
| M02641 | 2015-11-17 | Pass | Some samples high; Three samples very low | Low cluster density
151125_M02641_0053_000000000-AJJ3G | M02641 | 2015-11-25 | Borderline Fail | One sample very high
| M00766 | 2015-11-26 | Pass
| M02641 | 2015-11-26 | Pass | One sample a little high | Cluster density slightly low
| M00766 | 2015-11-27 | Pass | One sample very low; Some samples high | Cluster density slightly low
| M02641 | 2015-12-01 | Pass | One sample very high
| M00766 | 2015-12-03 | Pass | Some samples low; Some samples high
151208_M00766_0167_000000000-AJ5JR | M00766 | 2015-12-08 | Borderline Fail | Some samples low; Some samples high | Low cluster density
| M02641 | 2015-12-10 | Pass
| M00766 | 2015-12-16 | Pass | One sample high; One sample very low
| M02641 | 2015-12-17 | Fail | Range | Cluster density low; Large difference between cluster density passing filter and cluster density
| M00766 | 2015-12-23 | Pass | One sample high | Cluster density slightly low
| M02641 | 2015-12-23 | Pass | One sample low; One sample high
| M02641 | 2015-12-31 | Pass | One sample high
| M00766 | 2016-01-07 | Pass | One sample high
| M00766 | 2016-01-08 | Fail | One sample low; One sample high | Low >Q30; Smearing in heatmap
| M00766 | 2016-01-14 | Pass | One sample low; Some samples high | Slightly bigger than expected difference between cluster density passing filter and cluster density
| M02641 | 2016-01-14 | Pass | Two samples high
| M02641 | 2016-01-28 | Pass | One sample high | Cluster density slightly low
| M00766 | 2016-01-29 | Pass | One sample high
| M02641 | 2016-01-29 | Pass | One sample low
| M00766 | 2016-02-06 | Borderline Pass | Two samples high | Cluster density slightly high; Slightly bigger than expected difference between cluster density passing filter and cluster density
| M02641 | 2016-02-06 | Pass | Some samples very high; Some samples very low
160209_M00766_0007_000000000-AMERJ | M00766 | 2016-02-09 | Borderline Fail | Two samples high | Cluster density high; >Q30 lower than expected; Bigger than expected difference between cluster density passing filter and cluster density
| M00766 | 2016-02-16 | Pass | One sample very high; One sample a little low
| M00766 | 2016-02-19 | Fail | No info | Cluster density aa little low; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
| M02641 | 2016-02-20 | Pass | One sample very low | Bit of smearing towards the ends of reads (small)
| M00766 | 2016-02-23 | Pass | One sample very high; Most samples low
123

## CompareManual2AutoCD
Pass Method 2 (Automated) | Passes | same/diff | run name for lookup | borderline? | auto wrong (to pass)? | false +ves | false -ves | true +ves | true -ves | Fail Method 2 (Automated) | Fails | same/diff | run name for lookup | borderline? | auto wrong? (to fail) | what test(s) triggered?
0 | 0 | same |  |  |  | no | no | yes | no | 121009_M00766_0002_000000000-A1U6P | 121009_M00766_0002_000000000-A1U6P | same
0 | 0 | same |  |  |  | no | no | yes | no | 130206_M00766_0002_000000000-A23JM | 130206_M00766_0002_000000000-A23JM | same
0 | 0 | same |  |  |  | no | no | yes | no | 130618_M00766_0019_000000000-A4FEU | 130618_M00766_0019_000000000-A4FEU | same
130624_M00766_0021_000000000-A53PT | 130624_M00766_0021_000000000-A53PT | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 130712_M00766_0024_000000000-A5AVV | diff | 130712_M00766_0024_000000000-A5AVV | yes | no | yes | no | no | no | 130712_M00766_0024_000000000-A5AVV | 0 | diff | 130712_M00766_0024_000000000-A5AVV | no | yes
130817_M00766_0037_000000000-A4EDD | 130817_M00766_0037_000000000-A4EDD | same |  |  |  | no | no | no | yes | 0 | 0 | same
130913_M00766_0049_000000000-A5B0E | 130913_M00766_0049_000000000-A5B0E | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 130916_M00766_0050_000000000-A5M1U | diff | 130916_M00766_0050_000000000-A5M1U |  | no | yes | no | no | no | 130916_M00766_0050_000000000-A5M1U | 0 | diff | 130916_M00766_0050_000000000-A5M1U |  | yes
0 | 0 | same |  |  |  | no | no | yes | no | 130924_M00766_0052_000000000-A5BE6 | 130924_M00766_0052_000000000-A5BE6 | same
0 | 131007_M00766_0055_000000000-A59JT | diff | 131007_M00766_0055_000000000-A59JT |  | no | yes | no | no | no | 131007_M00766_0055_000000000-A59JT | 0 | diff | 131007_M00766_0055_000000000-A59JT |  | yes
131025_M00766_0059_000000000-A5P9A | 131025_M00766_0059_000000000-A5P9A | same |  |  |  | no | no | no | yes | 0 | 0 | same
140207_M00766_0022_000000000-A7PH8 | 140207_M00766_0022_000000000-A7PH8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 140424_M00766_0033_000000000-A7BNA | 140424_M00766_0033_000000000-A7BNA | same
0 | 0 | same |  |  |  | no | no | yes | no | 140514_M00766_0036_000000000-A7BRK | 140514_M00766_0036_000000000-A7BRK | same
0 | 0 | same |  |  |  | no | no | yes | no | 140612_M00766_0039_000000000-A7BNL | 140612_M00766_0039_000000000-A7BNL | same
0 | 140616_M00766_0040_000000000-A78V9 | diff | 140616_M00766_0040_000000000-A78V9 |  | no | yes | no | no | no | 140616_M00766_0040_000000000-A78V9 | 0 | diff | 140616_M00766_0040_000000000-A78V9 |  | yes
140813_M00766_0047_000000000-A8PJL | 140813_M00766_0047_000000000-A8PJL | same |  |  |  | no | no | no | yes | 0 | 0 | same
140901_M00766_0048_000000000-AA63M | 140901_M00766_0048_000000000-AA63M | same |  |  |  | no | no | no | yes | 0 | 0 | same
141014_M00766_0053_000000000-A8R6M | 141014_M00766_0053_000000000-A8R6M | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 141017_M02641_0022_000000000-AA66H | 141017_M02641_0022_000000000-AA66H | same
0 | 0 | same |  |  |  | no | no | yes | no | 141024_M00766_0056_000000000-A8PC5 | 141024_M00766_0056_000000000-A8PC5 | same
141031_M00766_0058_000000000-AA3GN | 141031_M00766_0058_000000000-AA3GN | same |  |  |  | no | no | no | yes | 0 | 0 | same
141107_M00766_0060_000000000-AA8PM | 141107_M00766_0060_000000000-AA8PM | same |  |  |  | no | no | no | yes | 0 | 0 | same
141118_M00766_0061_000000000-A8P8J | 141118_M00766_0061_000000000-A8P8J | same |  |  |  | no | no | no | yes | 0 | 0 | same
141208_M00766_0063_000000000-A8P7C | 141208_M00766_0063_000000000-A8P7C | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 141208_M02641_0026_000000000-A8R55 | diff | 141208_M02641_0026_000000000-A8R55 |  | no | yes | no | no | no | 141208_M02641_0026_000000000-A8R55 | 0 | diff | 141208_M02641_0026_000000000-A8R55 |  | yes
0 | 0 | same |  |  |  | no | no | yes | no | 141212_M00766_0064_000000000-ACCEB | 141212_M00766_0064_000000000-ACCEB | same
0 | 0 | same |  |  |  | no | no | yes | no | 141216_M00766_0065_000000000-ACCDT | 141216_M00766_0065_000000000-ACCDT | same
0 | 0 | same |  |  |  | no | no | yes | no | 141219_M00766_0066_000000000-ACCB1 | 141219_M00766_0066_000000000-ACCB1 | same
150120_M00766_0071_000000000-AA63K | 150120_M00766_0071_000000000-AA63K | same |  |  |  | no | no | no | yes | 0 | 0 | same
150127_M02641_0027_000000000-AA65J | 150127_M02641_0027_000000000-AA65J | same |  |  |  | no | no | no | yes | 0 | 0 | same
150203_M00766_0074_000000000-AAUMH | 150203_M00766_0074_000000000-AAUMH | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150203_M02641_0028_000000000-ACBYG | diff | 150203_M02641_0028_000000000-ACBYG |  | no | yes | no | no | no | 150203_M02641_0028_000000000-ACBYG | 0 | diff | 150203_M02641_0028_000000000-ACBYG |  | yes
150205_M00766_0075_000000000-ACC43 | 150205_M00766_0075_000000000-ACC43 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150224_M00766_0078_000000000-ACMP5 | 150224_M00766_0078_000000000-ACMP5 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150309_M02641_0035_000000000-ACC3J | 150309_M02641_0035_000000000-ACC3J | same
150311_M00766_0081_000000000-ACN0W | 150311_M00766_0081_000000000-ACN0W | same |  |  |  | no | no | no | yes | 0 | 0 | same
150311_M02641_0036_000000000-AD8VB | 150311_M02641_0036_000000000-AD8VB | same |  |  |  | no | no | no | yes | 0 | 0 | same
150402_M00766_0087_000000000-ACCAY | 150402_M00766_0087_000000000-ACCAY | same |  |  |  | no | no | no | yes | 0 | 0 | same
150414_M00766_0090_000000000-AAU39 | 150414_M00766_0090_000000000-AAU39 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150427_M02641_0042_000000000-AD8LB | 150427_M02641_0042_000000000-AD8LB | same
150430_M00766_0092_000000000-AD77P | 150430_M00766_0092_000000000-AD77P | same |  |  |  | no | no | no | yes | 0 | 0 | same
150506_M00766_0094_000000000-AEVP8 | 150506_M00766_0094_000000000-AEVP8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150508_M00766_0096_000000000-AF9MW | 150508_M00766_0096_000000000-AF9MW | same |  |  |  | no | no | no | yes | 0 | 0 | same
150508_M02641_0047_000000000-AF7NG | 150508_M02641_0047_000000000-AF7NG | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150511_M02641_0048_000000000-AEUGT | 150511_M02641_0048_000000000-AEUGT | same
150516_M00766_0099_000000000-AF81F | 150516_M00766_0099_000000000-AF81F | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150521_M02641_0051_000000000-AF9JH | diff | 150521_M02641_0051_000000000-AF9JH |  | no | yes | no | no | no | 150521_M02641_0051_000000000-AF9JH | 0 | diff | 150521_M02641_0051_000000000-AF9JH |  | yes
150522_M00766_0103_000000000-AF9M3 | 150522_M00766_0103_000000000-AF9M3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150522_M02641_0052_000000000-AF82F | diff | 150522_M02641_0052_000000000-AF82F |  | no | yes | no | no | no | 150522_M02641_0052_000000000-AF82F | 0 | diff | 150522_M02641_0052_000000000-AF82F |  | yes
0 | 150601_M00766_0107_000000000-AF9N8 | diff | 150601_M00766_0107_000000000-AF9N8 |  | no | yes | no | no | no | 150601_M00766_0107_000000000-AF9N8 | 0 | diff | 150601_M00766_0107_000000000-AF9N8 |  | yes
150602_M00766_0108_000000000-AFMB8 | 150602_M00766_0108_000000000-AFMB8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150603_M02641_0056_000000000-AEY8R | diff | 150603_M02641_0056_000000000-AEY8R |  | no | yes | no | no | no | 150603_M02641_0056_000000000-AEY8R | 0 | diff | 150603_M02641_0056_000000000-AEY8R |  | yes
150604_M00766_0109_000000000-AFJ62 | 150604_M00766_0109_000000000-AFJ62 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150604_M02641_0057_000000000-AFL7N | diff | 150604_M02641_0057_000000000-AFL7N |  | no | yes | no | no | no | 150604_M02641_0057_000000000-AFL7N | 0 | diff | 150604_M02641_0057_000000000-AFL7N |  | yes
150605_M00766_0110_000000000-AF80F | 150605_M00766_0110_000000000-AF80F | same |  |  |  | no | no | no | yes | 0 | 0 | same
150610_M00766_0111_000000000-AFMWF | 150610_M00766_0111_000000000-AFMWF | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150611_M02641_0060_000000000-AFN2Y | diff | 150611_M02641_0060_000000000-AFN2Y |  | no | yes | no | no | no | 150611_M02641_0060_000000000-AFN2Y | 0 | diff | 150611_M02641_0060_000000000-AFN2Y |  | yes
0 | 150615_M02641_0062_000000000-AFH7C | diff | 150615_M02641_0062_000000000-AFH7C |  | no | yes | no | no | no | 150615_M02641_0062_000000000-AFH7C | 0 | diff | 150615_M02641_0062_000000000-AFH7C |  | yes
150619_M02641_0063_000000000-AFN2W | 150619_M02641_0063_000000000-AFN2W | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150625_M02641_0065_000000000-AEY2A | 150625_M02641_0065_000000000-AEY2A | same
150703_M00766_0117_000000000-AFN2F | 150703_M00766_0117_000000000-AFN2F | same |  |  |  | no | no | no | yes | 0 | 0 | same
150703_M02641_0001_000000000-AFMW2 | 150703_M02641_0001_000000000-AFMW2 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150709_M02641_0004_000000000-AFMPC | 150709_M02641_0004_000000000-AFMPC | same |  |  |  | no | no | no | yes | 0 | 0 | same
150727_M00766_0119_000000000-AF3TM | 150727_M00766_0119_000000000-AF3TM | same |  |  |  | no | no | no | yes | 0 | 0 | same
150727_M02641_0007_000000000-AGEW7 | 0 | diff | 150727_M02641_0007_000000000-AGEW7 |  | yes | no | yes | no | no | 0 | 150727_M02641_0007_000000000-AGEW7 | diff | 150727_M02641_0007_000000000-AGEW7 |  | no
150807_M00766_0122_000000000-AFMW7 | 150807_M00766_0122_000000000-AFMW7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150813_M00766_0124_000000000-AFYEH | 150813_M00766_0124_000000000-AFYEH | same |  |  |  | no | no | no | yes | 0 | 0 | same
150818_M00766_0125_000000000-AE8BP | 150818_M00766_0125_000000000-AE8BP | same |  |  |  | no | no | no | yes | 0 | 0 | same
150826_M00766_0127_000000000-AGKLT | 150826_M00766_0127_000000000-AGKLT | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150904_M00766_0129_000000000-AFNA2 | 150904_M00766_0129_000000000-AFNA2 | same
150910_M00766_0130_000000000-AGJGG | 150910_M00766_0130_000000000-AGJGG | same |  |  |  | no | no | no | yes | 0 | 0 | same
150916_M02641_0025_000000000-AFN33 | 150916_M02641_0025_000000000-AFN33 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150917_M02641_0026_000000000-AFN0A | 150917_M02641_0026_000000000-AFN0A | same |  |  |  | no | no | no | yes | 0 | 0 | same
150918_M00766_0133_000000000-AH985 | 150918_M00766_0133_000000000-AH985 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150923_M00766_0134_000000000-AFN32 | 150923_M00766_0134_000000000-AFN32 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150924_M02641_0027_000000000-AGJGM | 150924_M02641_0027_000000000-AGJGM | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150925_M00766_0135_000000000-AFL75 | diff | 150925_M00766_0135_000000000-AFL75 |  | no | yes | no | no | no | 150925_M00766_0135_000000000-AFL75 | 0 | diff | 150925_M00766_0135_000000000-AFL75 |  | yes
150929_M02641_0030_000000000-AFWJ3 | 150929_M02641_0030_000000000-AFWJ3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150930_M00766_0138_000000000-AH4BR | 150930_M00766_0138_000000000-AH4BR | same |  |  |  | no | no | no | yes | 0 | 0 | same
151005_M02641_0034_000000000-AH9ML | 151005_M02641_0034_000000000-AH9ML | same |  |  |  | no | no | no | yes | 0 | 0 | same
151008_M02641_0035_000000000-AFL73 | 151008_M02641_0035_000000000-AFL73 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151013_M02641_0038_000000000-AFLW0 | 151013_M02641_0038_000000000-AFLW0 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151014_M02641_0039_000000000-AFLFV | 151014_M02641_0039_000000000-AFLFV | same |  |  |  | no | no | no | yes | 0 | 0 | same
151015_M00766_0142_000000000-AJD8B | 151015_M00766_0142_000000000-AJD8B | same |  |  |  | no | no | no | yes | 0 | 0 | same
151021_M02641_0041_000000000-AGHAA | 151021_M02641_0041_000000000-AGHAA | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 151022_M00766_0146_000000000-AFN3B | diff | 151022_M00766_0146_000000000-AFN3B |  | no | yes | no | no | no | 151022_M00766_0146_000000000-AFN3B | 0 | diff | 151022_M00766_0146_000000000-AFN3B |  | yes
0 | 151023_M00766_0147_000000000-AFN37 | diff | 151023_M00766_0147_000000000-AFN37 |  | no | yes | no | no | no | 151023_M00766_0147_000000000-AFN37 | 0 | diff | 151023_M00766_0147_000000000-AFN37 |  | yes
151028_M02641_0043_000000000-AFL74 | 151028_M02641_0043_000000000-AFL74 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151030_M00766_0150_000000000-AFLDT | 151030_M00766_0150_000000000-AFLDT | same |  |  |  | no | no | no | yes | 0 | 0 | same
151030_M02641_0044_000000000-AJHLJ | 151030_M02641_0044_000000000-AJHLJ | same |  |  |  | no | no | no | yes | 0 | 0 | same
151103_M02641_0046_000000000-AJ5Y7 | 151103_M02641_0046_000000000-AJ5Y7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151105_M02641_0047_000000000-AJF4E | 151105_M02641_0047_000000000-AJF4E | same |  |  |  | no | no | no | yes | 0 | 0 | same
151116_M02641_0050_000000000-AJ6L3 | 151116_M02641_0050_000000000-AJ6L3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151117_M00766_0155_000000000-AJ6LL | 151117_M00766_0155_000000000-AJ6LL | same |  |  |  | no | no | no | yes | 0 | 0 | same
151117_M02641_0051_000000000-AJF4A | 151117_M02641_0051_000000000-AJF4A | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 151125_M02641_0053_000000000-AJJ3G | 151125_M02641_0053_000000000-AJJ3G | same
151126_M00766_0159_000000000-AJD7L | 151126_M00766_0159_000000000-AJD7L | same |  |  |  | no | no | no | yes | 0 | 0 | same
151126_M02641_0054_000000000-AJD7T | 151126_M02641_0054_000000000-AJD7T | same |  |  |  | no | no | no | yes | 0 | 0 | same
151127_M00766_0161_000000000-AJERP | 151127_M00766_0161_000000000-AJERP | same |  |  |  | no | no | no | yes | 0 | 0 | same
151201_M02641_0055_000000000-AJ6B2 | 151201_M02641_0055_000000000-AJ6B2 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151203_M00766_0165_000000000-AJFAA | 151203_M00766_0165_000000000-AJFAA | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 151208_M00766_0167_000000000-AJ5JR | 151208_M00766_0167_000000000-AJ5JR | same
151210_M02641_0058_000000000-AJ5YJ | 151210_M02641_0058_000000000-AJ5YJ | same |  |  |  | no | no | no | yes | 0 | 0 | same
151216_M00766_0170_000000000-ALEAF | 151216_M00766_0170_000000000-ALEAF | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 151217_M02641_0059_000000000-AJFAD | 151217_M02641_0059_000000000-AJFAD | same
151223_M00766_0172_000000000-AJRMN | 151223_M00766_0172_000000000-AJRMN | same |  |  |  | no | no | no | yes | 0 | 0 | same
151223_M02641_0060_000000000-AGHV2 | 151223_M02641_0060_000000000-AGHV2 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151231_M02641_0061_000000000-AJD8L | 151231_M02641_0061_000000000-AJD8L | same |  |  |  | no | no | no | yes | 0 | 0 | same
160107_M00766_0175_000000000-AH6C3 | 160107_M00766_0175_000000000-AH6C3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 160108_M00766_0176_000000000-AJDAJ | 160108_M00766_0176_000000000-AJDAJ | same
160114_M00766_0177_000000000-ALYTH | 160114_M00766_0177_000000000-ALYTH | same |  |  |  | no | no | no | yes | 0 | 0 | same
160114_M02641_0065_000000000-ALRU0 | 160114_M02641_0065_000000000-ALRU0 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160128_M02641_0070_000000000-ALAGE | 160128_M02641_0070_000000000-ALAGE | same |  |  |  | no | no | no | yes | 0 | 0 | same
160129_M00766_0003_000000000-ALAH0 | 160129_M00766_0003_000000000-ALAH0 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160129_M02641_0071_000000000-AMC8N | 160129_M02641_0071_000000000-AMC8N | same |  |  |  | no | no | no | yes | 0 | 0 | same
160206_M00766_0005_000000000-ALYU1 | 160206_M00766_0005_000000000-ALYU1 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160206_M02641_0074_000000000-AL604 | 160206_M02641_0074_000000000-AL604 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 160209_M00766_0007_000000000-AMERJ | 160209_M00766_0007_000000000-AMERJ | same
160216_M00766_0009_000000000-AL73W | 160216_M00766_0009_000000000-AL73W | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 160219_M00766_0011_000000000-AL6GM | 160219_M00766_0011_000000000-AL6GM | same
160220_M02641_0081_000000000-AMETF | 160220_M02641_0081_000000000-AMETF | same |  |  |  | no | no | no | yes | 0 | 0 | same
160223_M00766_0012_000000000-AL5YC | 160223_M00766_0012_000000000-AL5YC | same |  |  |  | no | no | no | yes | 0 | 0 | same
[6 empty] | 16 | 1 | 24 | 82
| same pass | 106 | [8 empty] | same fail | 106
| diff pass | 17 | [8 empty] | diff fail | 17
| false passes by automated | 1 | [8 empty] | false fails by automated | 16
| missed passes by automated | 16 | [8 empty] | missed fails by automated | 1
[10 empty] | concordant | 106
[10 empty] | not concordant | 17
[9 empty] | Failing | false +ves | 16
[10 empty] | false -ves | 1
[10 empty] | true +ves | 24
[10 empty] | true -ves | 82
[10 empty] | total | 123
95% CI
[10 empty] | sensitivity (true +ve rate) | 96 | 79.65 | 99.9
[10 empty] | specificity (true -ve rate) | 83.6735 | 74.84 | 90.37
[6 empty] | Removed false manual fail |  |  | Failing | false +ves | 16
[10 empty] | false -ves | 0
[10 empty] | true +ves | 24
[10 empty] | true -ves | 83
[10 empty] | total | 123
[10 empty] | sensitivity (true +ve rate) | 100 | 85.75 | 100
[10 empty] | specificity (true -ve rate) | 83.8384 | 75.09 | 90.47

## EyeAutoFailsWithCD
Old Incorrects without CD | RunID | InstrumentID | RunDate | Pass/Fail | Indexing Comments | General Comments
|  | M00766 | 2012-10-09 | Fail | No info | Large difference in cluster density passing filter and cluster density; Smearing in heatmap
|  | M00766 | 2013-02-06 | Fail |  | Cluster density low; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
|  | M00766 | 2013-06-18 | Fail |  | Cluster density low; Low >Q30; Smearing in heatmap
|  | M00766 | 2013-06-24 | Pass |  | Some cycles smearing in heatmap
130712_M00766_0024_000000000-A5AVV | 130712_M00766_0024_000000000-A5AVV | M00766 | 2013-07-12 | Borderline Pass | One sample a little low | Some cycles smearing in heatmap; Reasonable difference in cluster density passing filter and cluster density
|  | M00766 | 2013-08-17 | Pass | Some samples low; Some samples high
|  | M00766 | 2013-09-13 | Pass |  | Some cycles smearing in heatmap
| 130916_M00766_0050_000000000-A5M1U | M00766 | 2013-09-16 | Pass | One sample low
|  | M00766 | 2013-09-24 | Fail | Two samples big range | Cluster density a bit high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
131007_M00766_0055_000000000-A59JT | 131007_M00766_0055_000000000-A59JT | M00766 | 2013-10-07 | Borderline Pass | Two samples big range | Cluster density a bit low; Reasonable difference in cluster density passing filter and cluster density; Some smearing in heatmap
|  | M00766 | 2013-10-25 | Pass | Some samples very low; Some samples very high
|  | M00766 | 2014-02-07 | Pass | Two samples big range | Some smearing in heatmap towards end of reads (small) amount
|  | M00766 | 2014-04-24 | Borderline Fail | One sample high | Very low cluster density
|  | M00766 | 2014-05-14 | Fail | Two samples high; Two samples low | Cluster density high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
|  | M00766 | 2014-06-12 | Fail |  | Low >Q30; Smearing in heatmap
| 140616_M00766_0040_000000000-A78V9 | M00766 | 2014-06-16 | Borderline Pass |  | Cluster density a little low; Some smearing in heatmap
|  | M00766 | 2014-08-13 | Pass | One sample a little high
|  | M00766 | 2014-09-01 | Pass | One sample a little high; One sample a little low
|  | M00766 | 2014-10-14 | Pass |  | Some small amount of smearing towards the ends of reads
|  | M02641 | 2014-10-17 | Fail |  | Cluster density high; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
141024_M00766_0056_000000000-A8PC5 |  | M00766 | 2014-10-24 | Borderline Fail |  | Cluster density high; Reasonable difference in cluster density passing filter and cluster density; Some smearing in heatmap
|  | M00766 | 2014-10-31 | Pass |  | Cluster density could be a bit low
|  | M00766 | 2014-11-07 | Borderline Pass | Some samples high; Some samples low | Cluster density could be a bit low
|  | M00766 | 2014-11-18 | Pass
|  | M00766 | 2014-12-08 | Borderline Pass | Two samples high | Cluster density a little high; Some smearing in heatmap towards end of reads (small) amount
141208_M02641_0026_000000000-A8R55 | 141208_M02641_0026_000000000-A8R55 | M02641 | 2014-12-08 | Borderline Pass | One sample a little high; One sample a little low | Cluster density a little low; Smearing in heatmap particularly towards the end of read 2
|  | M00766 | 2014-12-12 | Fail | Some samples high; Some samples low; NTC high | Cluster density high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
|  | M00766 | 2014-12-16 | Fail | One sample low; One sample very high | Cluster density high; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
|  | M00766 | 2014-12-19 | Fail | Some samples a little low; One sample very high | Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
|  | M00766 | 2015-01-20 | Pass | One sample very high | Some small amount of smearing towards the ends of reads
|  | M02641 | 2015-01-27 | Pass | One sample low
|  | M00766 | 2015-02-03 | Pass | One sample very low; Some samples high | Reasonable difference in cluster density passing filter and cluster density
150203_M02641_0028_000000000-ACBYG | 150203_M02641_0028_000000000-ACBYG | M02641 | 2015-02-03 | Borderline Pass | One sample very low; One sample very high | Reasonable difference in cluster density passing filter and cluster density; Some smearing in heatmap towards the ends of reads
|  | M00766 | 2015-02-05 | Pass | Some samples very low; One sample high | Some smearing in heatmap towards the ends of reads
150224_M00766_0078_000000000-ACMP5 |  | M00766 | 2015-02-24 | Fail |  | High cluster density; Large difference in cluster density passing filter and cluster density
|  | M02641 | 2015-03-09 | Fail | One sample low | Cluster density a little high; Low >Q30; Reasonable difference between cluster density passing filter and cluster density; Smearing in heatmap
|  | M00766 | 2015-03-11 | Pass | Some samples low; Some samples high
|  | M02641 | 2015-03-11 | Pass | One sample high | Cluster density a little low
|  | M00766 | 2015-04-02 | Borderline Pass | Some samples high | Some smearing in heatmap towards end of reads
|  | M00766 | 2015-04-14 | Pass | One sample very low
150427_M02641_0042_000000000-AD8LB |  | M02641 | 2015-04-27 | Borderline Fail | One sample very low; One sample very high | Very low cluster density
|  | M00766 | 2015-04-30 | Borderline Pass | Some samples high | Cluster density a little high; Reasonable difference between cluster density passing filter and cluster density
|  | M00766 | 2015-05-06 | Pass |  | Cluster density a little on the high side
|  | M00766 | 2015-05-08 | Pass | One sample a bit low; Some samples high
|  | M02641 | 2015-05-08 | Pass | Some samples high
|  | M02641 | 2015-05-11 | Fail |  | Cluster density low; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
|  | M00766 | 2015-05-16 | Pass | One sample a bit low; Some samples high
150521_M02641_0051_000000000-AF9JH | 150521_M02641_0051_000000000-AF9JH | M02641 | 2015-05-21 | Borderline Pass | One sample very low; One sample very high | Cluster density a little low
|  | M00766 | 2015-05-22 | Pass | One sample high
150522_M02641_0052_000000000-AF82F | 150522_M02641_0052_000000000-AF82F | M02641 | 2015-05-22 | Pass | Two samples high | Reasonable difference between cluster density passing filter and cluster density
| 150601_M00766_0107_000000000-AF9N8 | M00766 | 2015-06-01 | Borderline Pass | One sample very low; One sample very high | Cluster density quite low
|  | M00766 | 2015-06-02 | Borderline Pass | One sample very high | Cluster density a little high; Reasonable difference between cluster density passing filter and cluster density; Some smearing in heatmap (small)
150612_M00766_0112_000000000-AFMW5 | 150603_M02641_0056_000000000-AEY8R | M02641 | 2015-06-03 | Borderline Pass |  | Cluster density a little low; Reasonable difference between cluster density passing filter and cluster density
|  | M00766 | 2015-06-04 | Pass | Some samples high
150604_M02641_0057_000000000-AFL7N | 150604_M02641_0057_000000000-AFL7N | M02641 | 2015-06-04 | Borderline Pass | Some samples low; Some samples high | Cluster density a little low; Slightly bigger than expected difference between cluster density passing filter and cluster density
|  | M00766 | 2015-06-05 | Borderline Pass | One sample very high | Cluster density a little high
|  | M00766 | 2015-06-10 | Pass | One sample very high
150611_M02641_0060_000000000-AFN2Y | 150611_M02641_0060_000000000-AFN2Y | M02641 | 2015-06-11 | Borderline Pass | Some samples high | Reasonable difference between cluster density passing filter and cluster density; Quite spread box plots
150615_M02641_0062_000000000-AFH7C | 150615_M02641_0062_000000000-AFH7C | M02641 | 2015-06-15 | Borderline Pass | One sample high | Cluster density a little low; Reasonable difference between cluster density passing filter and cluster density; Quite spread box plots
|  | M02641 | 2015-06-19 | Pass | One sample low; One sample high | Cluster density a little low
|  | M02641 | 2015-06-25 | Borderline Fail |  | Cluster density a little low; Large difference between cluster density passing filter and cluster density
|  | M00766 | 2015-07-03 | Pass | One sample a little high
|  | M02641 | 2015-07-03 | Pass | One sample very high; One sample very low
|  | M02641 | 2015-07-09 | Pass | One sample low; Some samples high | Cluster density slightly low
|  | M00766 | 2015-07-27 | Pass | Cluster of samples high; Cluster of samples low
150727_M02641_0007_000000000-AGEW7 | 150727_M02641_0007_000000000-AGEW7 | M02641 | 2015-07-27 | Fail | One sample a little high; One sample low | Low >Q30; Smearing in heatmap towards ends of reads
|  | M00766 | 2015-08-07 | Borderline Pass |  | Low cluster density
|  | M00766 | 2015-08-13 | Pass |  | Cluster density slightly on the low side
|  | M00766 | 2015-08-18 | Pass | One sample a little high
|  | M00766 | 2015-08-26 | Pass
|  | M00766 | 2015-09-04 | Borderline Fail | One sample high; Some samples low | High cluster density; Slightly bigger than expected difference between cluster density passing filter and cluster density; Smearing in heatmap towards ends of reads
|  | M00766 | 2015-09-10 | Pass |  | Slightly bigger than expected difference between cluster density passing filter and cluster density
|  | M02641 | 2015-09-16 | Borderline Pass | One sample a little low | Slightly high cluster density; Small amount of smearing towards the ends of reads
|  | M02641 | 2015-09-17 | Pass | One sample low | Small amount of smearing towards the ends of reads
|  | M00766 | 2015-09-18 | Pass
|  | M00766 | 2015-09-23 | Pass | One sample very high; One sample very low
|  | M02641 | 2015-09-24 | Pass | One sample a little high | Cluster density slightly on the low side
| 150925_M00766_0135_000000000-AFL75 | M00766 | 2015-09-25 | Borderline Pass | Some samples high; One sample very low | Low cluster density
|  | M02641 | 2015-09-29 | Pass | Some samples high; Some samples very low
|  | M00766 | 2015-09-30 | Borderline Pass | One sample high | Some cycles smearing in heatmap (read 1)
|  | M02641 | 2015-10-05 | Borderline Pass | One sample a little high | Cluster density a little high; Some smearing in heatmap towards end of reads (small) amount
|  | M02641 | 2015-10-08 | Pass | Some samples very high; Most samples low
|  | M02641 | 2015-10-13 | Pass
|  | M02641 | 2015-10-14 | Pass | One sample very low
|  | M00766 | 2015-10-15 | Pass | One sample a little high
|  | M02641 | 2015-10-21 | Pass | One sample high
| 151022_M00766_0146_000000000-AFN3B | M00766 | 2015-10-22 | Borderline Pass |  | Cluster density a little high; Reasonable difference between cluster density passing filter and cluster density
| 151023_M00766_0147_000000000-AFN37 | M00766 | 2015-10-23 | Borderline Pass | One sample high; One sample low | Cluster density a little high
|  | M02641 | 2015-10-28 | Pass
|  | M00766 | 2015-10-30 | Pass | Some samples low; Some samples high
|  | M02641 | 2015-10-30 | Pass | One sample high; Some samples low
|  | M02641 | 2015-11-03 | Pass | One sample high
|  | M02641 | 2015-11-05 | Pass | One sample very high; All others low
|  | M02641 | 2015-11-16 | Pass |  | Cluster density slightly low
|  | M00766 | 2015-11-17 | Borderline Pass | Some samples high; Some samples low | Some smearing towards the ends of reads (small amount)
|  | M02641 | 2015-11-17 | Pass | Some samples high; Three samples very low | Low cluster density
151125_M02641_0053_000000000-AJJ3G |  | M02641 | 2015-11-25 | Borderline Fail | One sample very high
|  | M00766 | 2015-11-26 | Pass
|  | M02641 | 2015-11-26 | Pass | One sample a little high | Cluster density slightly low
|  | M00766 | 2015-11-27 | Pass | One sample very low; Some samples high | Cluster density slightly low
|  | M02641 | 2015-12-01 | Pass | One sample very high
|  | M00766 | 2015-12-03 | Pass | Some samples low; Some samples high
151208_M00766_0167_000000000-AJ5JR |  | M00766 | 2015-12-08 | Borderline Fail | Some samples low; Some samples high | Low cluster density
|  | M02641 | 2015-12-10 | Pass
|  | M00766 | 2015-12-16 | Pass | One sample high; One sample very low
|  | M02641 | 2015-12-17 | Fail | Range | Cluster density low; Large difference between cluster density passing filter and cluster density
|  | M00766 | 2015-12-23 | Pass | One sample high | Cluster density slightly low
|  | M02641 | 2015-12-23 | Pass | One sample low; One sample high
|  | M02641 | 2015-12-31 | Pass | One sample high
|  | M00766 | 2016-01-07 | Pass | One sample high
|  | M00766 | 2016-01-08 | Fail | One sample low; One sample high | Low >Q30; Smearing in heatmap
|  | M00766 | 2016-01-14 | Pass | One sample low; Some samples high | Slightly bigger than expected difference between cluster density passing filter and cluster density
|  | M02641 | 2016-01-14 | Pass | Two samples high
|  | M02641 | 2016-01-28 | Pass | One sample high | Cluster density slightly low
|  | M00766 | 2016-01-29 | Pass | One sample high
|  | M02641 | 2016-01-29 | Pass | One sample low
|  | M00766 | 2016-02-06 | Borderline Pass | Two samples high | Cluster density slightly high; Slightly bigger than expected difference between cluster density passing filter and cluster density
|  | M02641 | 2016-02-06 | Pass | Some samples very high; Some samples very low
160209_M00766_0007_000000000-AMERJ |  | M00766 | 2016-02-09 | Borderline Fail | Two samples high | Cluster density high; >Q30 lower than expected; Bigger than expected difference between cluster density passing filter and cluster density
|  | M00766 | 2016-02-16 | Pass | One sample very high; One sample a little low
|  | M00766 | 2016-02-19 | Fail | No info | Cluster density aa little low; Low >Q30; Large difference in cluster density passing filter and cluster density; Smearing in heatmap
|  | M02641 | 2016-02-20 | Pass | One sample very low | Bit of smearing towards the ends of reads (small)
|  | M00766 | 2016-02-23 | Pass | One sample very high; Most samples low
123

## CompareManual2AutoNewCD
Pass Method 2 (Automated) | Passes | same/diff | run name for lookup | borderline? | auto wrong (to pass)? | false +ves | false -ves | true +ves | true -ves | Fail Method 2 (Automated) | Fails | same/diff | run name for lookup | borderline? | auto wrong? (to fail) | what test(s) triggered?
0 | 0 | same |  |  |  | no | no | yes | no | 121009_M00766_0002_000000000-A1U6P | 121009_M00766_0002_000000000-A1U6P | same
0 | 0 | same |  |  |  | no | no | yes | no | 130206_M00766_0002_000000000-A23JM | 130206_M00766_0002_000000000-A23JM | same
0 | 0 | same |  |  |  | no | no | yes | no | 130618_M00766_0019_000000000-A4FEU | 130618_M00766_0019_000000000-A4FEU | same
130624_M00766_0021_000000000-A53PT | 130624_M00766_0021_000000000-A53PT | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 130712_M00766_0024_000000000-A5AVV | diff | 130712_M00766_0024_000000000-A5AVV | yes | no | yes | no | no | no | 130712_M00766_0024_000000000-A5AVV | 0 | diff | 130712_M00766_0024_000000000-A5AVV | no | yes
130817_M00766_0037_000000000-A4EDD | 130817_M00766_0037_000000000-A4EDD | same |  |  |  | no | no | no | yes | 0 | 0 | same
130913_M00766_0049_000000000-A5B0E | 130913_M00766_0049_000000000-A5B0E | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 130916_M00766_0050_000000000-A5M1U | diff | 130916_M00766_0050_000000000-A5M1U |  | no | yes | no | no | no | 130916_M00766_0050_000000000-A5M1U | 0 | diff | 130916_M00766_0050_000000000-A5M1U |  | yes
0 | 0 | same |  |  |  | no | no | yes | no | 130924_M00766_0052_000000000-A5BE6 | 130924_M00766_0052_000000000-A5BE6 | same
0 | 131007_M00766_0055_000000000-A59JT | diff | 131007_M00766_0055_000000000-A59JT |  | no | yes | no | no | no | 131007_M00766_0055_000000000-A59JT | 0 | diff | 131007_M00766_0055_000000000-A59JT |  | yes
131025_M00766_0059_000000000-A5P9A | 131025_M00766_0059_000000000-A5P9A | same |  |  |  | no | no | no | yes | 0 | 0 | same
140207_M00766_0022_000000000-A7PH8 | 140207_M00766_0022_000000000-A7PH8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 140424_M00766_0033_000000000-A7BNA | 140424_M00766_0033_000000000-A7BNA | same
0 | 0 | same |  |  |  | no | no | yes | no | 140514_M00766_0036_000000000-A7BRK | 140514_M00766_0036_000000000-A7BRK | same
0 | 0 | same |  |  |  | no | no | yes | no | 140612_M00766_0039_000000000-A7BNL | 140612_M00766_0039_000000000-A7BNL | same
140616_M00766_0040_000000000-A78V9 | 140616_M00766_0040_000000000-A78V9 | same |  |  |  | no | no | no | yes | 0 | 0 | same
140813_M00766_0047_000000000-A8PJL | 140813_M00766_0047_000000000-A8PJL | same |  |  |  | no | no | no | yes | 0 | 0 | same
140901_M00766_0048_000000000-AA63M | 140901_M00766_0048_000000000-AA63M | same |  |  |  | no | no | no | yes | 0 | 0 | same
141014_M00766_0053_000000000-A8R6M | 141014_M00766_0053_000000000-A8R6M | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 141017_M02641_0022_000000000-AA66H | 141017_M02641_0022_000000000-AA66H | same
0 | 0 | same |  |  |  | no | no | yes | no | 141024_M00766_0056_000000000-A8PC5 | 141024_M00766_0056_000000000-A8PC5 | same
141031_M00766_0058_000000000-AA3GN | 141031_M00766_0058_000000000-AA3GN | same |  |  |  | no | no | no | yes | 0 | 0 | same
141107_M00766_0060_000000000-AA8PM | 141107_M00766_0060_000000000-AA8PM | same |  |  |  | no | no | no | yes | 0 | 0 | same
141118_M00766_0061_000000000-A8P8J | 141118_M00766_0061_000000000-A8P8J | same |  |  |  | no | no | no | yes | 0 | 0 | same
141208_M00766_0063_000000000-A8P7C | 141208_M00766_0063_000000000-A8P7C | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 141208_M02641_0026_000000000-A8R55 | diff | 141208_M02641_0026_000000000-A8R55 |  | no | yes | no | no | no | 141208_M02641_0026_000000000-A8R55 | 0 | diff | 141208_M02641_0026_000000000-A8R55 |  | yes
0 | 0 | same |  |  |  | no | no | yes | no | 141212_M00766_0064_000000000-ACCEB | 141212_M00766_0064_000000000-ACCEB | same
0 | 0 | same |  |  |  | no | no | yes | no | 141216_M00766_0065_000000000-ACCDT | 141216_M00766_0065_000000000-ACCDT | same
0 | 0 | same |  |  |  | no | no | yes | no | 141219_M00766_0066_000000000-ACCB1 | 141219_M00766_0066_000000000-ACCB1 | same
150120_M00766_0071_000000000-AA63K | 150120_M00766_0071_000000000-AA63K | same |  |  |  | no | no | no | yes | 0 | 0 | same
150127_M02641_0027_000000000-AA65J | 150127_M02641_0027_000000000-AA65J | same |  |  |  | no | no | no | yes | 0 | 0 | same
150203_M00766_0074_000000000-AAUMH | 150203_M00766_0074_000000000-AAUMH | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150203_M02641_0028_000000000-ACBYG | diff | 150203_M02641_0028_000000000-ACBYG |  | no | yes | no | no | no | 150203_M02641_0028_000000000-ACBYG | 0 | diff | 150203_M02641_0028_000000000-ACBYG |  | yes
150205_M00766_0075_000000000-ACC43 | 150205_M00766_0075_000000000-ACC43 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150224_M00766_0078_000000000-ACMP5 | 150224_M00766_0078_000000000-ACMP5 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150309_M02641_0035_000000000-ACC3J | 150309_M02641_0035_000000000-ACC3J | same
150311_M00766_0081_000000000-ACN0W | 150311_M00766_0081_000000000-ACN0W | same |  |  |  | no | no | no | yes | 0 | 0 | same
150311_M02641_0036_000000000-AD8VB | 150311_M02641_0036_000000000-AD8VB | same |  |  |  | no | no | no | yes | 0 | 0 | same
150402_M00766_0087_000000000-ACCAY | 150402_M00766_0087_000000000-ACCAY | same |  |  |  | no | no | no | yes | 0 | 0 | same
150414_M00766_0090_000000000-AAU39 | 150414_M00766_0090_000000000-AAU39 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150427_M02641_0042_000000000-AD8LB | 150427_M02641_0042_000000000-AD8LB | same
150430_M00766_0092_000000000-AD77P | 150430_M00766_0092_000000000-AD77P | same |  |  |  | no | no | no | yes | 0 | 0 | same
150506_M00766_0094_000000000-AEVP8 | 150506_M00766_0094_000000000-AEVP8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150508_M00766_0096_000000000-AF9MW | 150508_M00766_0096_000000000-AF9MW | same |  |  |  | no | no | no | yes | 0 | 0 | same
150508_M02641_0047_000000000-AF7NG | 150508_M02641_0047_000000000-AF7NG | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150511_M02641_0048_000000000-AEUGT | 150511_M02641_0048_000000000-AEUGT | same
150516_M00766_0099_000000000-AF81F | 150516_M00766_0099_000000000-AF81F | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150521_M02641_0051_000000000-AF9JH | diff | 150521_M02641_0051_000000000-AF9JH |  | no | yes | no | no | no | 150521_M02641_0051_000000000-AF9JH | 0 | diff | 150521_M02641_0051_000000000-AF9JH |  | yes
150522_M00766_0103_000000000-AF9M3 | 150522_M00766_0103_000000000-AF9M3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150522_M02641_0052_000000000-AF82F | diff | 150522_M02641_0052_000000000-AF82F |  | no | yes | no | no | no | 150522_M02641_0052_000000000-AF82F | 0 | diff | 150522_M02641_0052_000000000-AF82F |  | yes
150601_M00766_0107_000000000-AF9N8 | 150601_M00766_0107_000000000-AF9N8 | same |  |  |  | no | no | no | yes | 0 | 0 | same | 150601_M00766_0107_000000000-AF9N8 |  | no
150602_M00766_0108_000000000-AFMB8 | 150602_M00766_0108_000000000-AFMB8 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150603_M02641_0056_000000000-AEY8R | diff | 150603_M02641_0056_000000000-AEY8R |  | no | yes | no | no | no | 150603_M02641_0056_000000000-AEY8R | 0 | diff | 150603_M02641_0056_000000000-AEY8R |  | yes
150604_M00766_0109_000000000-AFJ62 | 150604_M00766_0109_000000000-AFJ62 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150604_M02641_0057_000000000-AFL7N | diff | 150604_M02641_0057_000000000-AFL7N |  | no | yes | no | no | no | 150604_M02641_0057_000000000-AFL7N | 0 | diff | 150604_M02641_0057_000000000-AFL7N |  | yes
150605_M00766_0110_000000000-AF80F | 150605_M00766_0110_000000000-AF80F | same |  |  |  | no | no | no | yes | 0 | 0 | same
150610_M00766_0111_000000000-AFMWF | 150610_M00766_0111_000000000-AFMWF | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 150611_M02641_0060_000000000-AFN2Y | diff | 150611_M02641_0060_000000000-AFN2Y |  | no | yes | no | no | no | 150611_M02641_0060_000000000-AFN2Y | 0 | diff | 150611_M02641_0060_000000000-AFN2Y |  | yes
0 | 150615_M02641_0062_000000000-AFH7C | diff | 150615_M02641_0062_000000000-AFH7C |  | no | yes | no | no | no | 150615_M02641_0062_000000000-AFH7C | 0 | diff | 150615_M02641_0062_000000000-AFH7C |  | yes
150619_M02641_0063_000000000-AFN2W | 150619_M02641_0063_000000000-AFN2W | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150625_M02641_0065_000000000-AEY2A | 150625_M02641_0065_000000000-AEY2A | same
150703_M00766_0117_000000000-AFN2F | 150703_M00766_0117_000000000-AFN2F | same |  |  |  | no | no | no | yes | 0 | 0 | same
150703_M02641_0001_000000000-AFMW2 | 150703_M02641_0001_000000000-AFMW2 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150709_M02641_0004_000000000-AFMPC | 150709_M02641_0004_000000000-AFMPC | same |  |  |  | no | no | no | yes | 0 | 0 | same
150727_M00766_0119_000000000-AF3TM | 150727_M00766_0119_000000000-AF3TM | same |  |  |  | no | no | no | yes | 0 | 0 | same
150727_M02641_0007_000000000-AGEW7 | 0 | diff | 150727_M02641_0007_000000000-AGEW7 |  | yes | no | yes | no | no | 0 | 150727_M02641_0007_000000000-AGEW7 | diff | 150727_M02641_0007_000000000-AGEW7 |  | no
150807_M00766_0122_000000000-AFMW7 | 150807_M00766_0122_000000000-AFMW7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150813_M00766_0124_000000000-AFYEH | 150813_M00766_0124_000000000-AFYEH | same |  |  |  | no | no | no | yes | 0 | 0 | same
150818_M00766_0125_000000000-AE8BP | 150818_M00766_0125_000000000-AE8BP | same |  |  |  | no | no | no | yes | 0 | 0 | same
150826_M00766_0127_000000000-AGKLT | 150826_M00766_0127_000000000-AGKLT | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 150904_M00766_0129_000000000-AFNA2 | 150904_M00766_0129_000000000-AFNA2 | same
150910_M00766_0130_000000000-AGJGG | 150910_M00766_0130_000000000-AGJGG | same |  |  |  | no | no | no | yes | 0 | 0 | same
150916_M02641_0025_000000000-AFN33 | 150916_M02641_0025_000000000-AFN33 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150917_M02641_0026_000000000-AFN0A | 150917_M02641_0026_000000000-AFN0A | same |  |  |  | no | no | no | yes | 0 | 0 | same
150918_M00766_0133_000000000-AH985 | 150918_M00766_0133_000000000-AH985 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150923_M00766_0134_000000000-AFN32 | 150923_M00766_0134_000000000-AFN32 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150924_M02641_0027_000000000-AGJGM | 150924_M02641_0027_000000000-AGJGM | same |  |  |  | no | no | no | yes | 0 | 0 | same
150925_M00766_0135_000000000-AFL75 | 150925_M00766_0135_000000000-AFL75 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150929_M02641_0030_000000000-AFWJ3 | 150929_M02641_0030_000000000-AFWJ3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
150930_M00766_0138_000000000-AH4BR | 150930_M00766_0138_000000000-AH4BR | same |  |  |  | no | no | no | yes | 0 | 0 | same
151005_M02641_0034_000000000-AH9ML | 151005_M02641_0034_000000000-AH9ML | same |  |  |  | no | no | no | yes | 0 | 0 | same
151008_M02641_0035_000000000-AFL73 | 151008_M02641_0035_000000000-AFL73 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151013_M02641_0038_000000000-AFLW0 | 151013_M02641_0038_000000000-AFLW0 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151014_M02641_0039_000000000-AFLFV | 151014_M02641_0039_000000000-AFLFV | same |  |  |  | no | no | no | yes | 0 | 0 | same
151015_M00766_0142_000000000-AJD8B | 151015_M00766_0142_000000000-AJD8B | same |  |  |  | no | no | no | yes | 0 | 0 | same
151021_M02641_0041_000000000-AGHAA | 151021_M02641_0041_000000000-AGHAA | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 151022_M00766_0146_000000000-AFN3B | diff | 151022_M00766_0146_000000000-AFN3B |  | no | yes | no | no | no | 151022_M00766_0146_000000000-AFN3B | 0 | diff | 151022_M00766_0146_000000000-AFN3B |  | yes
0 | 151023_M00766_0147_000000000-AFN37 | diff | 151023_M00766_0147_000000000-AFN37 |  | no | yes | no | no | no | 151023_M00766_0147_000000000-AFN37 | 0 | diff | 151023_M00766_0147_000000000-AFN37 |  | yes
151028_M02641_0043_000000000-AFL74 | 151028_M02641_0043_000000000-AFL74 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151030_M00766_0150_000000000-AFLDT | 151030_M00766_0150_000000000-AFLDT | same |  |  |  | no | no | no | yes | 0 | 0 | same
151030_M02641_0044_000000000-AJHLJ | 151030_M02641_0044_000000000-AJHLJ | same |  |  |  | no | no | no | yes | 0 | 0 | same
151103_M02641_0046_000000000-AJ5Y7 | 151103_M02641_0046_000000000-AJ5Y7 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151105_M02641_0047_000000000-AJF4E | 151105_M02641_0047_000000000-AJF4E | same |  |  |  | no | no | no | yes | 0 | 0 | same
151116_M02641_0050_000000000-AJ6L3 | 151116_M02641_0050_000000000-AJ6L3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151117_M00766_0155_000000000-AJ6LL | 151117_M00766_0155_000000000-AJ6LL | same |  |  |  | no | no | no | yes | 0 | 0 | same
151117_M02641_0051_000000000-AJF4A | 151117_M02641_0051_000000000-AJF4A | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 151125_M02641_0053_000000000-AJJ3G | 151125_M02641_0053_000000000-AJJ3G | same
151126_M00766_0159_000000000-AJD7L | 151126_M00766_0159_000000000-AJD7L | same |  |  |  | no | no | no | yes | 0 | 0 | same
151126_M02641_0054_000000000-AJD7T | 151126_M02641_0054_000000000-AJD7T | same |  |  |  | no | no | no | yes | 0 | 0 | same
151127_M00766_0161_000000000-AJERP | 151127_M00766_0161_000000000-AJERP | same |  |  |  | no | no | no | yes | 0 | 0 | same
151201_M02641_0055_000000000-AJ6B2 | 151201_M02641_0055_000000000-AJ6B2 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151203_M00766_0165_000000000-AJFAA | 151203_M00766_0165_000000000-AJFAA | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 151208_M00766_0167_000000000-AJ5JR | 151208_M00766_0167_000000000-AJ5JR | same
151210_M02641_0058_000000000-AJ5YJ | 151210_M02641_0058_000000000-AJ5YJ | same |  |  |  | no | no | no | yes | 0 | 0 | same
151216_M00766_0170_000000000-ALEAF | 151216_M00766_0170_000000000-ALEAF | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 151217_M02641_0059_000000000-AJFAD | 151217_M02641_0059_000000000-AJFAD | same
151223_M00766_0172_000000000-AJRMN | 151223_M00766_0172_000000000-AJRMN | same |  |  |  | no | no | no | yes | 0 | 0 | same
151223_M02641_0060_000000000-AGHV2 | 151223_M02641_0060_000000000-AGHV2 | same |  |  |  | no | no | no | yes | 0 | 0 | same
151231_M02641_0061_000000000-AJD8L | 151231_M02641_0061_000000000-AJD8L | same |  |  |  | no | no | no | yes | 0 | 0 | same
160107_M00766_0175_000000000-AH6C3 | 160107_M00766_0175_000000000-AH6C3 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 160108_M00766_0176_000000000-AJDAJ | 160108_M00766_0176_000000000-AJDAJ | same
160114_M00766_0177_000000000-ALYTH | 160114_M00766_0177_000000000-ALYTH | same |  |  |  | no | no | no | yes | 0 | 0 | same
160114_M02641_0065_000000000-ALRU0 | 160114_M02641_0065_000000000-ALRU0 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160128_M02641_0070_000000000-ALAGE | 160128_M02641_0070_000000000-ALAGE | same |  |  |  | no | no | no | yes | 0 | 0 | same
160129_M00766_0003_000000000-ALAH0 | 160129_M00766_0003_000000000-ALAH0 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160129_M02641_0071_000000000-AMC8N | 160129_M02641_0071_000000000-AMC8N | same |  |  |  | no | no | no | yes | 0 | 0 | same
160206_M00766_0005_000000000-ALYU1 | 160206_M00766_0005_000000000-ALYU1 | same |  |  |  | no | no | no | yes | 0 | 0 | same
160206_M02641_0074_000000000-AL604 | 160206_M02641_0074_000000000-AL604 | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 160209_M00766_0007_000000000-AMERJ | 160209_M00766_0007_000000000-AMERJ | same
160216_M00766_0009_000000000-AL73W | 160216_M00766_0009_000000000-AL73W | same |  |  |  | no | no | no | yes | 0 | 0 | same
0 | 0 | same |  |  |  | no | no | yes | no | 160219_M00766_0011_000000000-AL6GM | 160219_M00766_0011_000000000-AL6GM | same
160220_M02641_0081_000000000-AMETF | 160220_M02641_0081_000000000-AMETF | same |  |  |  | no | no | no | yes | 0 | 0 | same
160223_M00766_0012_000000000-AL5YC | 160223_M00766_0012_000000000-AL5YC | same |  |  |  | no | no | no | yes | 0 | 0 | same
[6 empty] | 13 | 1 | 24 | 85
| same pass | 109 | [8 empty] | same fail | 109
| diff pass | 14 | [8 empty] | diff fail | 14
| false passes by automated | 1 | [8 empty] | false fails by automated | 13
| missed passes by automated | 13 | [8 empty] | missed fails by automated | 2
[10 empty] | concordant | 109
[10 empty] | not concordant | 14
[9 empty] | Failing | false +ves | 13
[10 empty] | false -ves | 1
[10 empty] | true +ves | 24
[10 empty] | true -ves | 85
[10 empty] | total | 123
95% CI
[10 empty] | sensitivity (true +ve rate) | 96 | 79.65 | 99.9
[10 empty] | specificity (true -ve rate) | 86.7347 | 78.38 | 92.74
[6 empty] | Removed false manual fail |  |  | Failing | false +ves | 13
[10 empty] | false -ves | 0
[10 empty] | true +ves | 24
[10 empty] | true -ves | 86
[10 empty] | total | 123
[10 empty] | sensitivity (true +ve rate) | 100 | 85.75 | 100
[10 empty] | specificity (true -ve rate) | 86.8687 | 78.59 | 92.82